## Sheet1
Lot Number | Product Code | Date | Machine Next PM (days away) | Model Predictions
21C45963 | 8985 | 2021-09-20 | 5 | 0
21C48863X | 8881907003 | 2021-09-15 | 7 | 0.71429
21C48963X | 8881907003 | 2021-09-15 | 7 | 0.71429
21C49063X | 8881907003 | 2021-09-15 | 7 | 0.71429
21D27763X | 8881907003 | 2021-09-15 | 7 | 0.71429
21E12163 | 8537Y | 2021-09-17 | 6 | 0
21E12663 | 8930SA | 2021-10-25 | 29 | 0.033764
21E17863 | 8985PG2 | 2021-09-22 | 24 | 0
21F01963 | 8617RC | 2021-09-16 | 4 | 0.033764
21F03763 | 31323358 | 2021-10-09 | 26 | 0.14729
21F03963 | 31314886 | 2021-09-20 | 7 | 0.25
21F07163 | 31317483 | 2021-09-27 | 21 | 0.14729
21F07363 | 31353611 | 2021-10-02 | 20 | 0.033764
21F08263 | 8533SA | 2021-09-21 | 14 | 0.14729
21F08563 | 8537MW | 2021-09-22 | 29 | 0.033764
21F09163 | 8608U | 2021-10-18 | 21 | 0.033764
21F14863 | 31353603 | 2021-09-02 | 14 | 0.30134
21F15663X | 8881112059 | 2021-09-15 | 19 | 0.32529
21F15763 | 8938 | 2021-09-20 | 0 | 0.14729
21F17763 | 31353603 | 2021-09-16 | 13 | 0.30134
21F18263 | 8515 | 2021-09-22 | 7 | 0.033764
21F18463 | 8536SA | 2021-09-07 | 11 | 0.22917
21F19563 | 8990SA | 2021-10-04 | 24 | 0.30134
21F22263 | 8820 | 2021-09-17 | 21 | 0.033764
21F24163 | 8534SA | 2021-09-20 | 14 | 0.14729
21F24663 | 8607RC | 2021-09-27 | 4 | 0.033764
21F26463 | 8870 | 2021-10-18 | 4 | 0.25
21F26963 | 8997 | 2021-09-27 | 3 | 0.033764
21F27463 | 31142222 | 2021-10-11 | 1 | 0.033764
21F28763X | 8881112059 | 2021-09-15 | 6 | 0.32529
21F29163X | 8881112059 | 2021-09-15 | 3 | 0.32529
21F29363 | 8605RC | 2021-10-18 | 18 | 0.35879
21F29563 | 31144010 | 2021-10-01 | 22 | 0.14729
21F30863 | 8938 | 2021-10-04 | 22 | 0.033764
21F32263 | 8507MW | 2021-09-16 | 18 | 0.051887
21G00963 | 1522SA | 2021-09-10 | 18 | 0.32353
21G01063 | 31157111 | 2021-10-11 | 28 | 0.033764
21G01163 | 31323333 | 2021-09-22 | 8 | 0.033764
21G01363 | 8522Y | 2021-10-04 | 16 | 0.033764
21G01763 | 8605U | 2021-10-11 | 1 | 0.051887
21G01863 | 8851 | 2021-10-18 | 9 | 0.25
21G02263 | 8967Y | 2021-10-04 | 9 | 0.033764
21G02963 | 31143699 | 2021-10-04 | 1 | 0.30134
21G03863 | 89651 | 2021-10-11 | 23 | 0.051887
21G03963 | 8970 | 2021-09-20 | 13 | 0.033764
21G04063 | 8974 | 2021-09-20 | 26 | 0.033764
21G05063 | 8935Y | 2021-09-23 | 8 | 0.033764
21G07863 | 1523SA | 2021-09-27 | 10 | 0.033764
21G07963 | 31139747 | 2021-09-22 | 0 | 0.033764
21G08163 | 8860 | 2021-10-04 | 28 | 0.033764
21G08263 | 8900MW | 2021-09-27 | 1 | 0.033764
21G08563 | 8985 | 2021-10-04 | 8 | 0
21G09163 | 8918GSTC | 2021-09-16 | 7 | 0.033764
21G09263 | 8612RC | 2021-10-04 | 22 | 0.033764
21G09363 | 8985S | 2021-10-11 | 15 | 0.033764
21G10263X | 8881512852 | 2021-09-15 | 4 | 0.065915
21G10363X | 8881512852 | 2021-09-15 | 22 | 0.065915
21G10463X | 8881512852 | 2021-09-15 | 8 | 0.065915
21G10563X | 8881512852 | 2021-09-15 | 14 | 0.065915
21G11763X | 8881907102 | 2021-09-15 | 18 | 0.32529
21G11863X | 8881907102 | 2021-09-15 | 5 | 0.32529
21G13763 | 8506Y | 2021-09-16 | 24 | 0.033764
21G14463 | 8507SA | 2021-09-16 | 11 | 0.033764
21G14563 | 31314886 | 2021-09-28 | 30 | 0.25
21G15963 | 89671 | 2021-09-20 | 23 | 0.35879
21G16063 | 8836SA | 2021-10-11 | 8 | 0.14729
21G16363 | 85221R | 2021-09-20 | 18 | 0.35312
21G17063 | 8507SA | 2021-09-10 | 8 | 0.14729
21G17163 | 1522SA | 2021-09-30 | 6 | 0.32353
21G17563 | 8998PG2 | 2021-10-11 | 13 | 0
21G17663 | 8939 | 2021-09-21 | 22 | 0
21G17963 | 8936SA | 2021-09-27 | 13 | 0
21G18163 | 8935 | 2021-10-04 | 16 | 0.033764
21G18263 | 8934PG2 | 2021-10-04 | 6 | 0
21G18363 | 8934 | 2021-09-18 | 7 | 0
21G18463 | 8931PG2 | 2021-10-04 | 5 | 0
21G18763 | 65 | 2021-10-04 | 3 | 0.033764
21G18863 | 8980 | 2021-09-27 | 29 | 0.28571
21G19163 | 8985SPG2 | 2021-10-18 | 4 | 0
21G19263 | 8997PG2 | 2021-10-04 | 4 | 0
21G19663 | 85121 | 2021-09-27 | 5 | 0.40375
21G19763 | 85131 | 2021-09-27 | 25 | 0.051887
21G19863 | 8950SA | 2021-09-30 | 18 | 0.40375
21G20063 | 8871 | 2021-09-22 | 13 | 0.25
21G20163 | 8979MW | 2021-09-27 | 12 | 0.033764
21G20263 | 8617U | 2021-09-27 | 24 | 0.033764
21G20363 | 8997S | 2021-10-18 | 4 | 0
21G20463 | 8607U | 2021-09-27 | 9 | 0.033764
21G20563 | 8618U | 2021-09-20 | 18 | 0.033764
21G20663 | 47 | 2021-11-22 | 16 | 0.14729
21G20763 | 31300878 | 2021-11-16 | 4 | 0.14729
21G21063 | 8990Y | 2021-09-20 | 25 | 0.033764
21G21563 | 8608RC | 2021-10-18 | 14 | 0.033764
21G21763 | 8932PG2 | 2021-09-27 | 29 | 0
21H00163 | 31156916 | 2021-10-25 | 17 | 0.051887
21H00863 | 31143665 | 2021-10-04 | 12 | 0.35879
21H00963 | 8985CRH | 2021-09-17 | 0 | 0
21H01063 | 8556H | 2021-10-02 | 27 | 0.033764
21H01663 | 8550B | 2021-09-22 | 10 | 0.051887
21H01963 | 8989PG2 | 2021-10-25 | 26 | 0.11111
21H02163 | 8934CRS | 2021-09-23 | 29 | 0
21H02363 | 8985 | 2021-10-11 | 24 | 0
21H04463 | 8967Y | 2021-10-11 | 23 | 0.033764
21H04563 | 8970 | 2021-09-27 | 6 | 0.033764
21H04963 | 8525Y | 2021-09-27 | 15 | 0.42857
21H05063X | 412SE | 2021-09-15 | 29 | 0
21H05463 | 8933 | 2021-10-11 | 13 | 0.033764
21H05963 | 8541SA | 2021-09-27 | 5 | 0.051887
21H06363X | 412SE | 2021-09-15 | 11 | 0
21H06563 | 8989 | 2021-09-27 | 2 | 0
21H07863X | 8881512852 | 2021-09-15 | 2 | 0.065915
21H07963X | 8881512852 | 2021-09-15 | 5 | 0.065915
21H08063X | 8881512852 | 2021-09-15 | 1 | 0.065915
21H08163X | 8881512852 | 2021-09-15 | 1 | 0.065915
21H08563 | 8939CRS | 2021-09-18 | 18 | 0
21H08663X | 8881412012 | 2021-09-14 | 13 | 0.26
21H08763X | 8881412012 | 2021-09-14 | 7 | 0.26
21H09063X | 8881112059 | 2021-09-16 | 23 | 0.32529
21H09163X | 8881112059 | 2021-09-20 | 21 | 0.32529
21H09263X | 8881112059 | 2021-09-20 | 2 | 0.32529
21H09363X | 8881112059 | 2021-09-21 | 28 | 0.32529
21H09463X | 8881112059 | 2021-09-22 | 8 | 0.32529
21H09563X | 8881112059 | 2021-09-23 | 24 | 0.32529
21H10263X | 8881412012 | 2021-09-19 | 8 | 0.26
21H10363X | 8881412012 | 2021-09-20 | 11 | 0.26
21H10463X | 8881412012 | 2021-09-22 | 28 | 0.26
21H10763X | 8881112599 | 2021-09-14 | 16 | 0.27477
21H10863X | 8881112599 | 2021-09-13 | 26 | 0.27477
21H11263X | 8881512852 | 2021-11-02 | 17 | 0.065915
21H11363X | 8881512852 | 2021-09-15 | 23 | 0.065915
21H11963 | 8535Y | 2021-09-22 | 21 | 0.033764
21H12263 | 8939 | 2021-10-08 | 5 | 0
21H12463X | 412SE | 2021-09-20 | 10 | 0
21H12563X | 412SE | 2021-09-21 | 0 | 0
21J00163 | 8506SA | 2021-09-24 | 2 | 0.14729
21J00263 | 8507SA | 2021-09-10 | 26 | 0.14729
21J00563 | 8970Y | 2021-09-27 | 0 | 0
21J00663 | 8900SA | 2021-09-22 | 2 | 0.051887
21J01263X | 8881907102 | 2021-10-11 | 14 | 0.32529
21J01663 | 8906 | 2021-09-27 | 7 | 0.051887
21J01763 | 8506SA | 2021-09-21 | 21 | 0.14729
21J01963X | 8881112083 | 2021-09-15 | 3 | 0.32258
21J02463X | 8881112083 | 2021-09-16 | 11 | 0.32258
21J02663 | 85221R | 2021-09-27 | 22 | 0.35312
21J02763 | 8537SA | 2021-09-17 | 11 | 0.40375
21J02863 | 8506SA | 2021-09-16 | 29 | 0.14729
21J03163 | 1522SA | 2021-09-23 | 12 | 0.32353
21J03263 | 1522SA | 2021-10-07 | 7 | 0.32353
21J03363 | 8939 | 2021-10-29 | 24 | 0
21J03563 | 8618RC | 2021-11-08 | 25 | 0.033764
21J03663 | 8991 | 2021-09-10 | 29 | 0.033764
21J04363 | 65 | 2021-10-18 | 22 | 0.033764
21J04663 | 8935 | 2021-10-11 | 5 | 0.033764
21J04763 | 31142222 | 2021-09-27 | 6 | 0.033764
21J04863 | 89671 | 2021-10-04 | 5 | 0.35879
21J04963 | 31143699 | 2021-10-11 | 22 | 0.30134
21J05163 | 31353603 | 2021-10-15 | 21 | 0.30134
21J05263 | 8982 | 2021-11-01 | 29 | 0.34783
21J05463 | 31323333 | 2021-10-13 | 22 | 0.033764
21J05663 | 31353611 | 2021-10-09 | 29 | 0.033764
21J05763 | 8507SA | 2021-09-13 | 24 | 0.14729
21J05863 | 8507SA | 2021-09-20 | 22 | 0.14729
21J05963 | 85131 | 2021-10-11 | 29 | 0.051887
21J06063 | 8529 | 2021-10-11 | 25 | 0.033764
21J06363 | 8906 | 2021-10-18 | 13 | 0.051887
21J06463 | 89671 | 2021-10-11 | 28 | 0.35879
21J06563 | 8980 | 2021-10-11 | 25 | 0.28571
21J06663 | 8982CRH | 2021-09-19 | 28 | 0.14729
21J06763 | 8985CRS | 2021-10-16 | 25 | 0
21J06963 | 8997SPG2 | 2021-10-18 | 28 | 0
21J07163 | 8998SPG2 | 2021-10-17 | 0 | 0
21J07263 | 8536SA | 2021-10-18 | 21 | 0.22917
21J07463X | 8881412012 | 2021-09-26 | 8 | 0.26
21J07563X | 412SE | 2021-09-22 | 25 | 0
21J07663 | R7642B | 2021-09-16 | 28 | 0.14729
21J07763R | 5551761207 | 2021-09-15 | 30 | 0.033764
21J07863R | 5551761207 | 2021-10-06 | 12 | 0.033764
21J07963 | 80-8705 | 2021-10-13 | 1 | 0.033764
21J08063 | 8303SA | 2021-10-04 | 17 | 0.033764
21J08163 | 8980S | 2021-10-04 | 18 | 0.14729
21J08263 | 8985CRS | 2021-10-08 | 18 | 0
21J08363 | 8537SA | 2021-09-20 | 7 | 0.40375
21J08463 | 85221R | 2021-10-11 | 1 | 0.35312
21J08563 | 85221R | 2021-10-25 | 30 | 0.35312
21J08763 | 8506SA | 2021-09-28 | 0 | 0.14729
21J08963 | 8870 | 2021-10-04 | 30 | 0.25
21J09163 | 8618RC | 2021-10-18 | 4 | 0.033764
21J09463 | 8900SA | 2021-10-04 | 23 | 0.051887
21J09563 | 8980Y | 2021-10-11 | 29 | 0.033764
21J09863 | 1522SA | 2021-10-14 | 3 | 0.32353
21J10063 | 8541SA | 2021-10-01 | 0 | 0.051887
21J10163 | 31314886 | 2021-11-11 | 29 | 0.25
21J10263 | 8970 | 2021-10-11 | 8 | 0.033764
21J10363 | 8970 | 2021-10-18 | 8 | 0.033764
21J10463 | 89651 | 2021-10-18 | 23 | 0.033764
21J10563 | 31157111 | 2021-10-14 | 15 | 0.033764
21J10663 | 8850 | 2021-09-27 | 7 | 0
21J10763 | 8602RC | 2021-10-16 | 16 | 0.34783
21J10863 | R7642B | 2021-09-23 | 0 | 0.14729
21J10963 | R7642B | 2021-09-30 | 5 | 0.14729
21J11063X | 412SE | 2021-10-11 | 23 | 0
21J11163 | 8881112015 | 2021-10-11 | 13 | 0.46317
21J11263 | 8881112015 | 2021-10-12 | 20 | 0.46317
21J11363 | 8527R | 2021-10-15 | 26 | 0.14729
21J11563 | 31143665 | 2021-11-01 | 9 | 0.30134
21J11863 | 31314886 | 2021-09-22 | 26 | 0.25
21J11963X | 8881112083 | 2021-09-20 | 21 | 0.32258
21J12063X | 8881112083 | 2021-09-21 | 24 | 0.32258
21J12163X | 8881112083 | 2021-09-22 | 14 | 0.32258
21J12263X | 8881112083 | 2021-09-23 | 29 | 0.32258
21J12363 | 8602U | 2021-10-18 | 0 | 0.033764
21J12463 | 8938 | 2021-10-11 | 24 | 0.14729
21J12563 | 8989 | 2021-09-22 | 5 | 0
21J12663 | 8989CRH | 2021-10-15 | 21 | 0
21J12763 | 8506Y | 2021-09-28 | 10 | 0.033764
21J12863 | 8982CRH | 2021-10-13 | 9 | 0.033764
21J198634 | PT00104294 | 2021-09-15 | 0 | 0.051887

## nc_full
NC Number | Created Date | Created By | NC Source | NC Type | Discovery/Plant Area | Discovery/Plant Area Name | Item Type | Other Item | Product | Product Name | Initial Failure Mode | Description | Lot Number | Quantity Affected | Immediate Actions | Report Date | Occurrence Date | Failure Mode | Failure Mode Description | Root Cause | Root Cause Description | Disposition | Dispositioned Quantity | Disposition Cost | Phase | Status | Closure Due Date | Closed Date | TAT (Days) | Product Family | Value Stream
NC-IL081-07336 | 08-Jan-2019 | JACLYN.GRISKO | General | Internal |  |  | Equipment ID | 58002-IL081 | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-M-D-Dimensional | PR IL081 Crystal Lake Smoke Test | 123456789 | 10.0 | PR IL081 Crystal Lake Smoke Test | 08-Jan-2019 | 08-Jan-2019 | 1000 | WORKMANSHIP/WRONG ASSEMBLY | Adhesive Dispense Time Too High | Adhesive Dispense Time Too High | Rework | 10.0 | 0.0 | Closure | VOID | 07-Feb-2019 |  |  | 3 | Syringe
NC-IL081-07398 | 09-Jan-2019 | JACLYN.GRISKO | General | Supplier | 102 | Press 14-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | PR-Planned Deviation | PR Site Code Smoke Test | 123456789 | 10.0 | PR Site Code Smoke Test | 09-Jan-2019 | 09-Jan-2019 |  | smoke test | Air in the Lines | Air in the Lines |  |  | 0.0 | Initiation Verification | VOID | 08-Feb-2019 | 10-Jan-2019 | 1. | 3 | Syringe
NC-IL081-07711 | 17-Jan-2019 | BEATRIZ.GUILLEN | General | Internal | S7P | Packer-IL081 |  |  | 8881907003 | 10ML ORAL MED SYRINGE (AMBER) | SFF-SPIC-V-Labeling/Print 1 | Incorrect expiration date on inner carton labels. Labels read 211230 and it should read 240116. | 19A22163X | 85.0 | End of run was being performed when defect was found. End of run was completed with affected material being quarantined. - Donna Berceda 01/17/2019 | 17-Jan-2019 | 17-Jan-2019 | SFF-SPIC-V-Labeling/Print 1 | Lot Number and Expiration Date (where applicable) is incomplete or missing. Expiration date was entered incorrectly. | Human Error | Human Error | Rework | 85.0 | 0.0 | Closure | CLOSED | 29-Mar-2019 | 25-Mar-2019 | 66. | 7 | Syringe
NC-IL081-07740 | 18-Jan-2019 | GRACE.AUSTIN | General | Internal | WHS | WAREHOUSE-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | PR-Other | Missing skid 4578168 from product code 8980 Lot# 18G03863 of NCR 18100019. | 18G03863 | 180.0 | Escalated to QA Supervisor. DPFA initiated. | 18-Jan-2019 | 25-Oct-2018 | PR-Other | Other-MEDMFGPR | Process/Method | IT/Computer Systems - SAP transition issues led to larger than expected quantities of materials being in the warehouse. Existing QA hold area could not accommodate increased amount of hold product contained. No procedures in place to describe containment in place outside of hold area. | Other/NDF | 180.0 | 0.0 | Closure | CLOSED | 31-Mar-2020 | 27-Mar-2020 | 433. | G8 | Container
NC-IL081-07782 | 20-Jan-2019 | LAKEYSHA.SIMPSON | General | Internal | C19 | Press 19-IL081 |  |  | 8506SA | CONTAINER BIOHAZARD 5 QT | CFF-V-Flash (Functional Area) | Flash in tab lock area greater than 0.060" (WP8506) See GS100000913. | 18L18563 | 1440.0 | Turned on the water supply to the manifolds allowing the tool to cool properly. Michael Buck 02/05/2019 | 24-Jan-2019 | 20-Jan-2019 | CFF-V-Flash (Functional Area) | Sharp Flash; Hinge Pins and Holes; All other areas: Less than or equal to 0.060" and no more than 3 locations per part. | High Mold Temp | High Mold Temp - caused by the water supply to the manifolds being off. | Sort | 1440.0 | 0.0 | Closure | CLOSED | 12-Apr-2019 | 25-Mar-2019 | 64. | SS | Container
NC-IL081-07833 | 22-Jan-2019 | MARQUES.CALIMEE | Inspection | Internal | C70 | Press 70-IL081 |  |  | 8906 | 1QT N/S CONTAINER YELLOW 100 | CFF-D-Molding 1 | A low wall thickness of 0.064" was documented on an inspection for cavity 3 in the gate area. The minimum wall thickness is 0.065". | 18K14563 | 5400.0 | Issue was found during documentation review. No line start up process corrections were made. Process was verified to be within specification at time of documentation identification. | 21-Jan-2019 | 21-Jan-2019 | CFF-D-Molding 1 |  | Human Error | Human Error | UseAsIs | 5400.0 | 0.0 | Closure | CLOSED | 22-Apr-2019 | 25-Mar-2019 | 62. | G1 | Container
NC-IL081-07837 | 22-Jan-2019 | LAKEYSHA.SIMPSON | General | Internal | C46 | Press 46-IL081 |  |  | V8536 | CONTAINER 3 GALLON | CFF-D-Molding 1 | Low wall thickness cavity #2. Min spec. is 0.055" measured at 0.054" | 19A07163 | 860.0 | Added shims to increase wall thickness. | 22-Jan-2019 | 22-Jan-2019 | CFF-D-Molding 1 |  | Tooling/ fixture wear |  | Sort | 860.0 | 0.0 | Closure | CLOSED | 28-Jun-2019 | 24-Apr-2019 | 92. | SS | Container
NC-IL081-08021 | 25-Jan-2019 | MARK.VENTURA | General | Internal |  |  |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Grease/Lint/Particulate 1 | Foreign matter on rubber tip (1 defect). White and colorless. (particulate in fluid path). Guerbet Scar 18-027 | 18J17063X | 335000.0 | N/A - Post production NCR | 25-Jan-2019 | 25-Jan-2019 | SFF-M-V-Grease/Lint/Particulate 1 | Grease/Lint/Particulate Type 1: Particulate in fluid pathway | Damaged material | Damaged component causing potential miss-assemblies | UseAsIs | 335000.0 | 0.0 | Closure | CLOSED | 23-May-2019 | 15-Apr-2019 | 80. | Sub | Syringe
NC-IL081-08056 | 27-Jan-2019 | ANGELICA.AGUILAR | General | Internal | [6 empty] | CFF-V-Incorrect Items | Incorrect IFU packed into the Finished Product. | 18K19063 | 160.0 |  | 27-Jan-2019 | [7 empty] | 0.0 | Initiation Verification | VOID
NC-IL081-08057 | 27-Jan-2019 | ANGELICA.AGUILAR | General | Internal | C80 | Press 80-IL081 |  |  | 47PSS | PSS SELECT 2 GAL , RED | CFF-V-Incorrect Items | Incorrect IFU packed into the Finished Product. | 18K19063 | 160.0 | "Quarantined all affected product. Removed all incorrect IFU's before restarting press." | 27-Jan-2019 | 27-Jan-2019 | CFF-V-Incorrect Items | Incorrect IFU/Absorbent pads/Other Kit items/Other Labels in the Overshipper Box | Human Error | Root Cause: Human Error\nThe Manufacturing Associate failed to follow Standard Work FS10026366 and failed to place the required amount of instruction sheets in each finished good case. | Rework | 160.0 | 0.0 | Closure | CLOSED | 27-Apr-2019 | 02-Apr-2019 | 65.
NC-IL081-08114 | 29-Jan-2019 | LAKEYSHA.SIMPSON | General | Internal | C42 | Press 42-IL081 |  |  | PT00016490 | 12/18 GAL. HINGED CHEMO LID | CFF-D-Molding 5 | Failed Hinge thickness inspection. Hinge Thickness specification should be between 0.016"-0.028" measured at 0.029" | 186682 | 2250.0 |  | 28-Jan-2019 | 28-Jan-2019 | CFF-D-Molding 5 | Hinge Thickness | Corrective maintenance not been addressed per Workorder | R7515A mold was identified with damaged insert and misaligned shim on previous run when an NCR#18120041 was issued for hinge thickness. On 01/06/19, a workorder#CL3052050 was created in the Maximo to address the issue. On 01/28/19, the mold was placed back in production without addressing recommended corrections to run at risk. | Scrap | 2250.0 | 0.0 | Closure | CLOSED | 30-Apr-2019 | 29-Apr-2019 | 90. | G12 | Container
NC-IL081-08178 | 30-Jan-2019 | FIDELIA.CANTOR | General | Internal |  |  |  |  | R7449 | 8GAL CHEMO LID-SLIDETOP(X9895) | PR-Planned Deviation | NON-standard shop order/IQ of press 61(MP-1138) with tool R7449B and RM093 | 186837 | 100.0 | No immediate actions required-planned deviation. | 30-Jan-2019 | 30-Jan-2019 | [6 empty] | 0.0 | Initiation Verification | VOID
NC-IL081-08194 | 31-Jan-2019 | ROSEMARY.TEKSTAR | General | Internal | C76B | Press 76B-IL081 |  |  | R7450 | Chemotherapy Hood, 8/12/18 Gallon | PR-Planned Deviation | Non Standard Shop Order - Transfer test run 76B MP-1106 with tool R7450A and RM171. CDP-0015874 | 186835 | 200.0 | No immediate action required, planned deviation. | 30-Jan-2019 | 30-Jan-2019 | [6 empty] | 0.0 | Initiation Verification | VOID
NC-IL081-08222 | 31-Jan-2019 | GRACE.AUSTIN | General | Internal |  |  |  |  | R7505 | 2/3 GALLON LID | PR-Planned Deviation | Non Standard Shop Order- IQ for R7505 on MP-1137 (6D) | 186836 | 500.0 | No Immediate Action required, Non Standard Shop Order | 31-Jan-2019 | [7 empty] | 0.0 | Initiation Verification | VOID
NC-IL081-08230 | 31-Jan-2019 | GRACE.AUSTIN | General | Internal | C74 | Press 74-IL081 |  |  | 8851 | 8GA PHARMASAFETY WASTE SLIDE | CFF-F-Printing | Regulated Medical Waste - UN #3291; was not printed on Biohazard Label. per marking requirement on Formula Sheet | 18L09463' '18M17263' '18L19563' '18I15763' '18K13963' '18J09263 | 10670.0 3780.0 4800.0 3880.0 2570.0 8960.0 | Place all affected product on Hold and reach out to regulatory. | 31-Jan-2019 | 31-Jan-2019 | CFF-F-Printing | Biohazard Label Verification | Procedures: UnClear | Procedures: UnClear | Other/NDF | 3880.0 8960.0 10670.0 4800.0 3780.0 2570.0 | 0.0 | Closure | CLOSED | 06-Apr-2019 | 29-Mar-2019 | 56. | G8 | Container
NC-IL081-08231 | 31-Jan-2019 | ROSEMARY.TEKSTAR | General | Internal | C89 | Press 89-IL081 |  |  | 31353603 | 5 QT GATORGUARD JR - TRANS RED | DHRDOC | Missing PCR document from 01/14/2019. | 18H11763 | 2016.0 | PCR readings from the shifts prior to the 14th and after the 14th were reviewed and documented to be within validated parameters. | 31-Jan-2019 | 14-Jan-2019 | [6 empty] | 0.0 | Investigation | VOID |  |  |  | GG | Container
NC-IL081-08244 | 31-Jan-2019 | BERTHA.HERNANDEZ | General | Internal | S6P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | PR- Documentation Incomplete | Start up checklist missing PTL verification signature. | 19A06663X | 335000.0 | Reviewed Lot 19A06663X inspection results and PCR verifications for accuracy. - D. Rose 01/31/19 | 31-Jan-2019 | 30-Jan-2019 | Procedure Not Followed |  | Procedure Not Followed | The Production Team Lead completed all the assigned tasks and verified the start-up correctly, but did not sign the start-up checklist because they forgot to. See Human Error Form for further details. | Other/NDF | 335000.0 | 0.0 | Closure | CLOSED | 25-Feb-2019 | 25-Feb-2019 | 24. | Sub | Syringe
NC-IL081-08283 | 01-Feb-2019 | GRACE.AUSTIN | General | Internal | C82 | Press 82-IL081 |  |  | 8982 | 2 GA HINGED CHEMO - YELLOW | CFF-V-Contamination (Embedded) | Contamination (Embedded) un-melted Resin on the containers | 18M08163 | 17140.0 | Maintenance changed the heater bands. Ken Wiltse. | 01-Feb-2019 | 01-Feb-2019 | CFF-V-Contamination (Embedded) | Metal, cardboard, paper, etc | Human Error | Human Error | Other/NDF | 17140.0 | 0.0 | Closure | CLOSED | 11-May-2019 | 09-May-2019 | 97. | G2 | Container
NC-IL081-08401 | 04-Feb-2019 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 7777129656 | M906-7 TIP CAP SO | Visual- Shipping Damage | Shipping damaged (found 6 cases damaged). | 1827404764 | 150000.0 | Defective product contained at Incoming. | 04-Feb-2019 | 01-Nov-2018 | Visual- Shipping Damage | Visual- Shipping Damage | PR-Packaging | Packaging-MEDMFGPR. The product passed KenMex\'s release criteria and was released for shipment. The damage appears to have occurred in transit by the Carrier, Old Dominion, due to double stacking pallets. The Carrier may have double stacked the pallets because the pallets did not have "do not double stack" labels on them. | Scrap | 150000.0 | 0.0 | Closure | CLOSED | 06-Mar-2019 | 21-Feb-2019 | 16. | IQA | Incoming
NC-IL081-08403 | 04-Feb-2019 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 7777129656 | M906-7 TIP CAP SO | Broken / Damaged upon receipt | Shipping damaged (1 case damaged) | 1828800864 | 25000.0 | Damaged product contained at Incoming. | 04-Feb-2019 | 09-Nov-2018 | Visual- Shipping Damage | Visual- Shipping Damage | PR-Packaging | Packaging-MEDMFGPR. The product passed KenMex\'s release criteria and was released for shipment. The damage appears to have occurred in transit by the Carrier, Old Dominion, due to double stacking pallets. The Carrier may have double stacked the pallets because the pallets did not have "do not double stack" labels on them. | Scrap | 25000.0 | 0.0 | Closure | CLOSED | 06-Mar-2019 | 21-Feb-2019 | 16. | IQA | Incoming
NC-IL081-08442 | 05-Feb-2019 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | PT00012138 | (COV) Printed Insert, CHEMOSAFETY Waste Disposal Container | Incorrect lot number | PT00012138 was received with a lot number (42476) on the product label that does not match the lot number on the CoC (42472). | 42476 | 10000.0 | Placed product on hold. | 05-Feb-2019 | 05-Feb-2019 | Label information incorrect | Label information incorrect. Received shipment with lot number (42476) on the product label that does not match the lot number on the CoC (42472). | Process/Method | Process/Method. Supplier did not have a check in place to verify label information. Supplier stated in their investigation "After expanding our procedures to include extra checking on the Covidien paperwork, we still experience human error in a typo on the label itself." | Return to Vendor | 10000.0 | 0.0 | Closure | CLOSED | 22-Mar-2019 | 27-Feb-2019 | 21.
NC-IL081-08500 | 06-Feb-2019 | BERTHA.HERNANDEZ | General | Internal | S8F | Form, Fill Sealer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SPBP-F-UDI Barcode | Lidding cavities 7 & 8 failed barcode scan, found during scheduled audit by process inspector. | 19A32163X | 154560.0 | Front paper station taken out of service for tension issues. Parts on order. - P. Myerson 02/09/19 | 06-Feb-2019 | 04-Feb-2019 | SFF-SPBP-F-UDI Barcode | UDI Barcode | Incorrect Setup | Incorrect Setup | Sort | 154560.0 | 0.0 | Closure | CLOSED | 29-Mar-2019 | 22-Mar-2019 | 43. | 8 | Syringe
NC-IL081-08501 | 07-Feb-2019 | BERTHA.HERNANDEZ | General | Internal | S8K | Kahle-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SA-V-Particulate 2 | Spraymation exceeded upper specification limits. | 19A32163X | 154560.0 | Spraymation head replaced with reading above upper specification limit. | 06-Feb-2019 | 05-Feb-2019 | SFF-SA-V-Particulate 2 | Particulate Type 2: Extraneous matter outside of syringe; Excessive lubricant in or on the syringe | Incorrect Pressure | Incorrect Pressure | UseAsIs | 154560.0 | 0.0 | Closure | CLOSED | 12-Apr-2019 | 21-Mar-2019 | 42. | 8 | Syringe
NC-IL081-08511 | 07-Feb-2019 | JESSE.SOBEL | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-F-Spraymation | Spraymation values exceed upper control limit for Line 6. | 19A18963X' '19A18863X' '19A18763X | 232000.0 335000.0 | Spraymation tested multiple times until values were acceptable. Lot was cut short. - Berenice Villatoro 02/07/2019 | 07-Feb-2019 | 06-Feb-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | Sub | Syringe
NC-IL081-08515 | 07-Feb-2019 | ALEJANDRA.CASTANEDA | Inspection | Internal | C46 | Press 46-IL081 |  |  | 8536SA | 3 GA SHARPSTAR - CLEAR | CFF-D-Molding 1 | Low wall thickness on Cavity # 1- 0.053, Should be 0.055 | 18L12263 | 5770.0 | Water line were checked for clogs and the mold was shimmed to correct wall thickness. | 07-Feb-2019 | 06-Feb-2019 | CFF-D-Molding 1 | Wall Thickness | Equipment variation requiring maintenance adjustments | The take-out robot found to be faulty with bad “up & down” cylinder and was intermittently failing to remove containers from core. Mold closed on a stack container and double shot resulted in a core shift causing the low wall. | Sort | 5770.0 | 0.0 | Closure | CLOSED | 30-May-2019 | 07-May-2019 | 89. | SS | Container
NC-IL081-08633 | 09-Feb-2019 | ANGELICA.AGUILAR | General | Internal | C81 | Press 81-IL081 |  |  | 89671 | 2 GA HORIZONTAL DROP - T RED | CFF-V-Label Verification and Barcode Verification 6 | Failed barcode verification (UDI) on both sides of container | 18L06963 | 2160.0 | Replaced cartridges and adjusted print head positioning to restore acceptable UDI barcode print quality. S. Hahn 02/10/2019 | 09-Feb-2019 | 09-Feb-2019 | CFF-V-Label Verification and Barcode Verification 6 | Print - failed UDI barcode inspection | Printer Out of Alignment | Printer Out of Alignment | Scrap | 2160.0 | 0.0 | Closure | CLOSED | 31-May-2019 | 15-May-2019 | 95. | G2 | Container
NC-IL081-08644 | 09-Feb-2019 | BERTHA.HERNANDEZ | General | Internal | S8K | Kahle-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SA-V-Particulate 2 | Spraymation pressure was found out of parameters spec is 50-60, setting found at 64. The resulting potential product impact would most likely result in excessive lubricant in/on the barrel. Until product can be confirmed to be/not be impacted, we will utilize this potential impact as the failure mode. | 19A32263X | 37920.0 | Line stopped, all affected product was placed on hold. Spraymation setting placed back to validated parameter. | 09-Feb-2019 | 08-Feb-2019 | SFF-SA-V-Particulate 2 | Particulate Type 2: Extraneous matter outside of syringe; Excessive lubricant in or on the syringe | People/Human Error/ | People/Human Error/ | Scrap | 37920.0 | 0.0 | Closure | CLOSED | 30-Apr-2019 | 15-May-2019 | 95. | 8 | Syringe
NC-IL081-08646 | 10-Feb-2019 | JESSE.SOBEL | General | Internal | S7P | Packer-IL081 | Room ID | PF 7-IL081 | 8881907465 | 10ML ORAL DOSE SYR BULK NS | PR-Process Step Missed | Start up checklist not performed on plunger mold for line 7 (MP-1114). PCR was signed at 1930. All product from 1900 and forward was contained. | 19B04563X | 16100.0 | Samples were pulled from case 2-45 and up to performed startup checklist and first article inspection for MP-1114 plunger mold. Berenice Villatoro 02/11/2019 | 09-Feb-2019 | 09-Feb-2019 | Startup checklist |  | Human Error | Human Error- Cell tech failed to identify startup checklist was not completed for plunger press prior to starting the machine. See Human Error Form for further details. | Other/NDF | 16100.0 | 0.0 | Closure | CLOSED | 30-Apr-2019 | 17-Apr-2019 | 66. | 7 | Syringe
NC-IL081-08649 | 10-Feb-2019 | JESSE.SOBEL | General | Internal |  |  |  |  | 8881907465 | 10ML ORAL DOSE SYR BULK NS | SFF-P-V-Numbering 2 | Graduation Lines touching numerals. | 19B04663X | 16100.0 | Lot was cut short. Removed excessive thinner from printer and added ink to bring the viscosity level up. | 10-Feb-2019 | 10-Feb-2019 | SFF-P-V-Numbering 2 | Type 1: Numeral(s) are touching graduation lines | Print Pad Issue | Print Pad Issue | Scrap | 16100.0 | 0.0 | Closure | CLOSED | 20-Jun-2019 | 19-Jun-2019 | 129. | 7 | Syringe
NC-IL081-08765 | 12-Feb-2019 | GRACE.AUSTIN | General | Internal | C56 | Press 56-IL081 |  |  | 8997S | 8 GA GASKETED SLIDE LID - RED | CFF-F-Molding 3 | Lid Fit Inspection Failed. | 18L16863' '185959 | 2650.0 3911.0 | Change jobs over to non-gasketed lids. M. Sailor 04/24/19. | 12-Feb-2019 | 12-Feb-2019 | CFF-F-Molding 3 | Lid Fit | Gasket Overlap under a lifter | The lids that were produced were observed with overlap location of gasket under a lifter. | Scrap | 2650.0 3911.0 | 0.0 | Closure | CLOSED | 02-Jul-2019 | 19-Jun-2019 | 126. | G8 | Container
NC-IL081-08780 | 13-Feb-2019 | MARIA.ROSALES | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SPIC-V-Labeling/Print 2 | Smeared and illegible print found on inner carton an unscheduled inspection. | 19A28563X | 828.0 | Print block straightened and backed off. - Kim Kaecker 02/13/19 | 13-Feb-2019 | 13-Feb-2019 | SFF-SPIC-V-Labeling/Print 2 |  | incorrect tooling/ fixture | Use of incorrect tooling/ fixture | Sort | 828.0 | 0.0 | Closure | CLOSED | 14-Apr-2019 | 27-Mar-2019 | 42. | 3 | Syringe
NC-IL081-08816 | 13-Feb-2019 | GRACE.AUSTIN | General | Internal | C74 | Press 74-IL081 |  |  | 8850 | 8 GA PHARMASAFETY CONT -HINGED | CFF-V-Incorrect Items | Miss-Assembly, Packing incorrect component. wrong Absorbent Pads | 18L19463 | 900.0 | Contained product, removed incorrect pads, had material handler bring correct pads. - Steve Schirmer 04/02/19 | 13-Feb-2019 | 13-Feb-2019 | CFF-V-Incorrect Items |  | Human Error | Human Error | Rework | 900.0 | 0.0 | Closure | CLOSED | 14-Apr-2019 | 10-Apr-2019 | 56. | G8 | Container
NC-IL081-08830 | 13-Feb-2019 | ROSEMARY.TEKSTAR | General | Internal | C31 | Press 31-IL081 |  |  | 8982CRH | 2 GALLON CHEMO RED | CFF-V-Quantity per Case | Package missing one absorbent pad package per SW. | 18K20163 | 560.0 | Operator was instructed to follow standard work which required 1 pack of absorbent pads inside the box. Product was reworked by the press and press was restarted. O. Ratinskiy 03/22/2019 | 13-Feb-2019 | 13-Feb-2019 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | The Manufacturing Associate failed to follow Standard Work DC00008108 and failed to place the required absorbent pad pack in each finished good case. | Rework | 560.0 | 0.0 | Closure | CLOSED | 14-Apr-2019 | 20-May-2019 | 95. | G2 | Container
NC-IL081-08853 | 14-Feb-2019 | GRACE.AUSTIN | General | Internal | C77 | Press 77-IL081 |  |  | 31317483 | SAG4836 TY TRANSLUCENT YELLOW | CFF-V-Blisters, Bubbles and Voids | Bubbles on Cavity 2 bottom area out of specification per GS10000913 | 18K22863 | 2592.0 | Re-hung R7597C tool, shimmed R7597C tool 3 separate times. Performed AQL with passing results. Steve Shirmer. | 14-Feb-2019 | 14-Feb-2019 | CFF-V-Blisters, Bubbles and Voids |  | Venting |  | Sort | 2592.0 | 0.0 | Closure | CLOSED | 15-May-2019 | 08-May-2019 | 83. | GG | Container
NC-IL081-08854 | 14-Feb-2019 | BERTHA.HERNANDEZ | General | Internal | S1R | RAM-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-M-V-Broken/Damaged 2 | Damaged component (plunger) functional but less than desired. | 19A25463X | 11200.0 | "Technician found the low level sensor on side 2 wasn't working and 1A the high sensor not working. It wasn't seeing the part. Both sensors adjusted. Product reworked. - P. Myerson 02/14/2019" | 14-Feb-2019 | 14-Feb-2019 | SFF-M-V-Broken/Damaged 2 |  | Sensor not aligned | Sensor not aligned | Sort | 11200.0 | 0.0 | Closure | CLOSED | 15-Apr-2019 | 27-Mar-2019 | 41. | 1 | Syringe
NC-IL081-08865 | 14-Feb-2019 | ROSEMARY.TEKSTAR | General | Internal | S3R | RAM-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | PR-Planned Deviation | Non-standard shop order - Air piping and heat stake retrofit test run. CL-3052084 and CL-3044616 | TSTA021419 | 28800.0 | N/A, planned deviation | 14-Feb-2019 | 14-Feb-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | 3 | Syringe
NC-IL081-08871 | 14-Feb-2019 | ROSEMARY.TEKSTAR | General | Internal | C25 | Press 25-IL081 |  |  | 8507SA | CONTAINER SHARPS 5QT RED | Documentation Incomplete/Missing | Product was packed into cases before the start-up checklist was completed and signed. | 18H18763 | 160.0 | Start-Up checklist was reviewed and signed. - S. Schirmer 04/02/19 | 14-Feb-2019 | 14-Feb-2019 | Documentation Incomplete/Missing | Documentation Incomplete/Missing | Human Error | The Manufacturing Associate failed to follow GS10000122 Rev 206 Section 5.9 which requires the start -up form to be signed off prior to any product being packed. | No Issue / No Impact | 160.0 | 0.0 | Closure | CLOSED | 30-Jun-2019 | 12-Jun-2019 | 117. | SS | Container
NC-IL081-08963 | 16-Feb-2019 | ROSEMARY.TEKSTAR | General | Internal | C75 | Press 75-IL081 |  |  | R7568 | Tortuous Door | Documentation Missing | DHR missing inspection table from 02/14/19. | 186322 | 7959.0 | All product produced has been placed on hold. | 15-Feb-2019 | 15-Feb-2019 | [6 empty] | 0.0 | Initiation Verification | VOID
NC-IL081-08976 | 16-Feb-2019 | FIDELIA.CANTOR | General | Internal | C43C | Press 43C-IL081 |  |  | 8920SA | 0.5 GA SHARPS CONTAINER | CFF-V-Packaging 1 | Missing one container in the box packed only 59 correct QTY is 60 containers. Not per standard work. | 18K24363 | 420.0 | Removed affected product to QA hold. Instructed associate to read and follow quality procedure available at the job station. Minda Sailor. | 16-Feb-2019 | 16-Feb-2019 | CFF-V-Packaging 1 |  | Human Error | Human Error | Rework | 420.0 | 0.0 | Closure | CLOSED | 17-Apr-2019 | 27-Mar-2019 | 39. | MS | Container
NC-IL081-09038 | 19-Feb-2019 | GRACE.AUSTIN | General | Internal | AS1 | Assembly 1-IL081 |  |  | 8974 | HOLDER BEIGE PLASTIC W ADHESIV | CFF-V-Quantity per Case | Incorrect Quantity in the Box, it should 5 ea tray holder per case, packed only 4. | 19A10163 | 270.0 | 100% screened all product at the press. Verified correct standard work was present also. - S. Hahn 02/18/19 | 19-Feb-2019 | 19-Feb-2019 | CFF-V-Quantity per Case |  | Human Error | Human Error | Sort | 270.0 | 0.0 | Closure | CLOSED | 20-May-2019 | 07-May-2019 | 77. | AS | Container
NC-IL081-09084 | 20-Feb-2019 | JESSE.SOBEL | General | Internal | S8F | Form, Fill Sealer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-PPSP-V-Labeling/Print 2 | During rework activities, the soft pack lidding print was found to be smeared and not legible. An AQL was performed and failed. | 19A32163X | 19200.0 | N/A - Lot was completed and the line was currently not running. J. Burle 03/20/2019 | 20-Feb-2019 | 20-Feb-2019 | SFF-PPSP-V-Labeling/Print 2 | Excessive ink/print greater than 0.080" | Printer Die Issue | Printer Die Issue | Sort | 19200.0 | 0.0 | Closure | CLOSED | 20-Jun-2019 | 13-Jun-2019 | 113. | 8 | Syringe
NC-IL081-09101 | 20-Feb-2019 | BERTHA.HERNANDEZ | General | Internal | S8B | 8B-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | Outside Validated Parameters | Charge recovery time out of parameters for MP-1133/58119, found at 3.75 spec is 3.8-4.2. | 19A32763X | 123840.0 | Parameters brought back into acceptable range, press restarted, and monitoring. Product placed on hold. - P. Myerson 02/20/2019 | 20-Feb-2019 | 19-Feb-2019 | Outside Validated Parameters |  | Outside Validated Parameters |  | Re-Inspect' 'Other/NDF | 200.0 123640.0 | 0.0 | Closure | CLOSED | 21-Apr-2019 | 01-Apr-2019 | 40. | 8 | Syringe
NC-IL081-09147 | 21-Feb-2019 | ALEJANDRA.CASTANEDA | General | Internal | C82 | Press 82-IL081 |  |  | 31300919 | SAG4838-TY 5 QT TURTUOUS PATH | PR-Process Step Missed | Start up checklist not perform during In process start up, Press was Down over 24 hours Start up Checklist required per GS10000133 | 19B09363 | 1290.0 | Performed an inprocess start-up and inspection. Ken Wiltse 02/21/2019 | 21-Feb-2019 | 21-Feb-2019 | PR-Process Step Missed | Failed to follow procedure GS10000122. | Human Error | Production lead failed to perform startup checklist at restart per GS10000122, sec.5.8.3). The production lead got distracted with another task and forgot to initiate a startup checklist prior to resume production. | Other/NDF | 1290.0 | 0.0 | Closure | CLOSED | 22-May-2019 | 17-Jun-2019 | 116. | G5 | Container
NC-IL081-09232 | 23-Feb-2019 | LAKEYSHA.SIMPSON | General | Internal | C31 | Press 31-IL081 |  |  | 8979C | 2 GA ROTOR LID SHARPS - T RED | CFF-V-Packaging 1 | Incorrect packaging-missing inserts PT00044732/PT000447281 and PT00044733. | 19B26263 | 5300.0 | Missing components (PT00044728, PT00044732 and PT00044733) were brought to the press. Operator was instructed to follow packaging standard work, press restarted. - O. Ratinskiy 02/22/2019 | 22-Feb-2019 | 22-Feb-2019 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | Per standard work DC00047470 operator failed to place dividers and inserts into finished goods box | Rework | 5300.0 | 0.0 | Closure | CLOSED | 22-Jun-2019 | 19-Jun-2019 | 116. | G2 | Container
NC-IL081-09279 | 25-Feb-2019 | BERTHA.HERNANDEZ | General | Internal | S4A | 412A-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-M-V-Bubbles 1 | Bubble in fluid pathway. | 19B11363X | 12500.0 | Slowed down the cycle and increased mold heats. 100% case 120-144. - P. Myerson 02/25/19 | 25-Feb-2019 | 25-Feb-2019 | SFF-M-V-Bubbles 1 |  | High Mold Temp | High Mold Temp | Sort | 12500.0 | 0.0 | Closure | CLOSED | 26-Apr-2019 | 27-Mar-2019 | 30. | 412 | Syringe
NC-IL081-09291 | 26-Feb-2019 | ROSEMARY.TEKSTAR | General | Internal | AS1 | Assembly 1-IL081 |  |  | 8556H- | "5QT ENCLOSURE, GB, SHARPSTAR" | Documentation Incomplete/Missing | Start-Up checklist not completed after being down over 24 hrs. Operator was packing cases. BPCS hold # 19020054. | 18B22363 | 9.0 | Completed start-up checklist. - J. Pixius 05/31/19 | 25-Feb-2019 | 25-Feb-2019 | Documentation Incomplete/Missing | Documentation Incomplete/Missing | HUMAN ERROR | The Manufacturing Associate failed to follow GS10000122 Rev 206 Section 5.9 which requires the start -up form to be signed off prior to any product being packed. | Other/NDF | 9.0 | 0.0 | Closure | CLOSED | 30-Jun-2019 | 12-Jun-2019 | 106.
NC-IL081-09356 | 26-Feb-2019 | JACLYN.GRISKO | General | Internal | QA | QA-IL081 |  |  |  |  | PR-Procedure Not Followed | On 2/25/19, as related to CAPA-IL081-00533, a new version of Bartender print software was installed to printing work stations on the Crystal Lake manufacturing floor and a new version of the governing procedure PS20000402 was issued. On 2/27/19, it was identified that the new version of software was causing unanticipated physical print complication (multiple labels exiting printer at once). The software needed to be rolled back to the still-valid former version, but the governing procedure has already been issued. This NCR documents this fact and reviews the gap analysis for the two procedure versions and its impact. |  |  | 1. Rollback of Bartender software version from version 11 to version 10. | 26-Feb-2019 | 26-Feb-2019 | PR-Procedure Not Followed | Procedure Not Followed-MEDMFGPR | Other | Other |  |  | 0.0 | Closure | CLOSED | 31-Mar-2020 | 27-Mar-2020 | 394.
NC-IL081-09481 | 01-Mar-2019 | BERTHA.HERNANDEZ | General | Internal | S7R | RAM-IL081 |  |  | 8881112059 | 12ML SYRINGE REGULAR NS | SFF-SA-V-Mixed Product 2 | Incomplete assembly of product, plunger not engaged to rubber tip. | 19B51463X | 12000.0 | The plunger push down rods were lowered on the ram. 100% cases 130-153. - P. Myerson 03/01/2019 | 01-Mar-2019 | 01-Mar-2019 | SFF-SA-V-Mixed Product 2 |  | Jam |  | Sort | 12000.0 | 0.0 | Closure | CLOSED | 30-May-2019 | 07-May-2019 | 67. | 7 | Syringe
NC-IL081-09530 | 03-Mar-2019 | MARK.VENTURA | General | Internal | S6P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SPBP-V-Labeling/Print 1 | Inputted correct information and reprinted on to DHR time and case when corrected. K. Kaecker 03/03/2019 | 19B14663X | 122000.0 | corrected the printer to include the letter B. | 03-Mar-2019 | 03-Mar-2019 | SFF-SPBP-V-Labeling/Print 1 | Label/Print, Incomplete label control number on over-shipper. The letter B is missing from 1000BX. | Lot number, expiry, UDI; missing/incorrect/not legible | Lot number, expiry, UDI; missing/incorrect/not legible | Rework | 122000.0 | 0.0 | Closure | CLOSED | 12-Apr-2019 | 22-Mar-2019 | 19. | Sub | Syringe
NC-IL081-09543 | 04-Mar-2019 | LAKEYSHA.SIMPSON | General | Internal | C63 | Press 63-IL081 |  |  | WP8998 | 8931 LID | CFF-D-Molding 3 | Failed height gasket measurements and leak test inspection. | 19B187294 | 834.0 | Cell Tech slowed the robot speed. Fran Henry. | 04-Mar-2019 | 03-Mar-2019 | CFF-D-Molding 3 | Gasket Height | Forming head malfunction | Forming head malfunction | Scrap | 834.0 | 0.0 | Closure | CLOSED | 03-Apr-2019 | 02-Apr-2019 | 29. | G18 | Container
NC-IL081-09547 | 04-Mar-2019 | LAKEYSHA.SIMPSON | General | Internal | S8P | Packer-IL081 |  |  | 8881540122 | 12ML SMARTIP MED-PREP CANNULA | SFF-PPSP-V-Seal Integrity 1 | Unacceptable seal integrity on cavity 1 and 2. | 19B38563X | 13920.0 | Sealing gasket and plate cleaned for stuck paper. Per KK. | 04-Mar-2019 | 04-Mar-2019 | SFF-PPSP-V-Seal Integrity 1 | Assembly Defect - Damage -Jammed Parts | Part Out of Pocket | Part Out of Pocket | Sort | 13920.0 | 0.0 | Closure | CLOSED | 03-May-2019 | 01-Apr-2019 | 28. | 8 | Syringe
NC-IL081-09563 | 04-Mar-2019 | GRACE.AUSTIN | General | Internal | C67 | Press 67-IL081 |  |  | 8934CRS | 12 GAL CHEMO RED SLIDE LID | CFF-V-Label Verification and Barcode Verification 16 | Wrinkled Biohazard Labels | 19BB25963 | 130.0 | Press was down for break. Removed affected products and put to QA hold. Re-started and address associate to smooth label before stacking them. - M. Sailor 03/04/2019 | 04-Mar-2019 | 04-Mar-2019 | CFF-V-Label Verification and Barcode Verification 16 | BH - Wrinkled labels | CFF-V-Label Verification and Barcode Verification 16 |  | Rework | 130.0 | 0.0 | Closure | CLOSED | 03-May-2019 | 27-Mar-2019 | 23. | G12 | Container
NC-IL081-09603 | 05-Mar-2019 | MARIA.ROSALES | General | Internal | S7N | Pad Printer-IL081 |  |  | 8881112059 | 12ML SYRINGE REGULAR NS | SFF-P-V-Graduations 3 | Graduation lines do not have same thickness | 19B38963X | 9000.0 | Discarded/scrapped affected parts. Technicians changed out the pad for the printer and verified print was acceptable. - L. Reed 03/05/2019 | 05-Mar-2019 | 05-Mar-2019 | SFF-P-V-Graduations 3 | Type 3: Inconsistent graduation lines. Graduation lines do not have same thickness | Print Pad Issue | Print Pad Issue | Scrap | 9000.0 | 0.0 | Closure | CLOSED | 03-Jun-2019 | 15-May-2019 | 71. | 7 | Syringe
NC-IL081-09645 | 05-Mar-2019 | ROSEMARY.TEKSTAR | General | Internal | C54 | Press 54-IL081 | Equipment ID | R7536A-IL081 | 8930SA | "30GA RED W/ HINGED, GASKET LID" | CFF-V-Flash (Non-Functional Area), Gate | Sharp flash on gate area. Operator did not trim flash as instructed per FM20011813. | 19B28463 | 231.0 | Standard work for trimming flash was at press, operator was instructed to trim the gates of each container. | 05-Mar-2019 | 05-Mar-2019 | CFF-V-Flash (Non-Functional Area), Gate | Sharp Flash; Surfaces A & B: greater than or equal to 0.080”; Surface C: greater than or equal to 0.125”; All Surfaces: more than 3 locations in total | High Mold Temp | High Mold Temp. The thermolator temperature of valves side was higher than temperature of the thermolator of the core side. Both temperatures were within the limits. | Rework | 111.0 120.0 | 0.0 | Closure | CLOSED | 26-Mar-2019 | 10-Apr-2019 | 35. | MS | Container
NC-IL081-09668 | 06-Mar-2019 | GRACE.AUSTIN | General | Internal | C54 | Press 54-IL081 |  |  | 8930SA | "30GA RED W/ HINGED, GASKET LID" | CFF-V-Label Verification and Barcode Verification 13 | "Wrong product code on the over shipper label, it's labeled 851201, Product code should be 8930SA" | 19B28463 | 210.0 | Stop production and verified that all labels currently at the line contained the correct information. - Justin Pixius 03/06/19 | 06-Mar-2019 | 05-Mar-2019 | CFF-V-Label Verification and Barcode Verification 13 | OS - Wrong label | Human Error | Human Error | Rework | 210.0 | 0.0 | Closure | CLOSED | 05-Apr-2019 | 05-Apr-2019 | 30. | MS | Container
NC-IL081-09669 | 06-Mar-2019 | GRACE.AUSTIN | General | Internal | C54 | Press 54-IL081 |  |  | 8930SA | "30GA RED W/ HINGED, GASKET LID" | CFF-V-Flash (Non-Functional Area), Gate | Sharp Flash at the gate area | 19B28463 | 111.0 | Place all affected Product on hold for rework. | 06-Mar-2019 | 05-Mar-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | MS | Container
NC-IL081-09785 | 08-Mar-2019 | MARQUES.CALIMEE | General | Internal | C87 | Press 87-IL081 |  |  | 31353603 | 5 QT GATORGUARD JR - TRANS RED | PCR out of Specification | PCR was documented out of parameters, 3.23 sec. Spec is 3.25-3.75 sec. in a critical area, during the previous shift at press 87. The lids went directly to press 89 used into FG 31353603 | 19B187341 | 1512.0 | Cell Tech adjusted the hold time back within the specified tolerances of the key process variables and documented same on the process control record. Francis Henry. | 07-Mar-2019 | 07-Mar-2019 | PCR out of Specification |  | Human Error | Human Error | Scrap | 1512.0 | 0.0 | Closure | CLOSED | 07-May-2019 | 19-Apr-2019 | 42. | GG | Container
NC-IL081-09829 | 09-Mar-2019 | LAKEYSHA.SIMPSON | General | Internal |  |  |  |  | WP8998 | 8931 LID | Failed Leak Test | Failed Leak Test | 19B187294 | 1.0 |  | 09-Mar-2019 | [7 empty] | 0.0 | Initiation Verification | VOID |  |  |  | G18 | Container
NC-IL081-09830 | 09-Mar-2019 | LAKEYSHA.SIMPSON | General | Internal | C63 | Press 63-IL081 |  |  | 8931 | 12 GA HINGED CHEMO - YELLOW | Failed Leak Test | Failed Leak Test Position C | 19B187294 | 1200.0 | Replaced the gearbox hat supplies ISO to the head. Mike Buck 04/04/19. | 09-Mar-2019 | 09-Mar-2019 | Failed Leak Test |  | Machine Failure | Machine Failure | Scrap | 1200.0 | 0.0 | Closure | CLOSED | 08-Apr-2019 | 15-Apr-2019 | 37. | G12 | Container
NC-IL081-09848 | 11-Mar-2019 | ROSEMARY.TEKSTAR | General | Internal | C89 | Press 89-IL081 | Equipment ID | R7599B-IL081 | 31353603 | 5 QT GATORGUARD JR - TRANS RED | CFF-V-Short | Shorts on handle. | 18M05163 | 98.0 | Appropriate adjustments were made to the process to correct the issue and all WIP parts were scrapped. - M. Islam 04/05/2019 | 11-Mar-2019 | 11-Mar-2019 | CFF-V-Short | Part is not completely formed; Missing plastic | human error | Human Error - A technical operator performed process inspection and documented actual value without noticing it was out of specification. MP-1041 is a hydraulic injection molding machine and takes time to fully stabilize the process. Machine was down for 12 hours during previous shift. Tech performed their process inspection immediately upon restarting the machine and did not allow time to stabilize the process. | Sort | 98.0 | 0.0 | Closure | CLOSED | 30-Aug-2019 | 12-Aug-2019 | 154. | GG | Container
NC-IL081-09905 | 12-Mar-2019 | GRACE.AUSTIN | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | PR-Planned Deviation | Run at Risk awaiting for CDP-0011281 | 19C08863X | 230400.0 | Product will be placed on hold until the CDP-0011281 has reached or surpassed the implementation phase then product may be released. | 12-Mar-2019 | 12-Mar-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | 8 | Syringe
NC-IL081-09906 | 12-Mar-2019 | GRACE.AUSTIN | General | Internal |  |  |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | PR-Planned Deviation | Run at Risk awaiting for CDP-0011281 | 19C08963X | 230400.0 | Product will be placed on Hold until the CDP-0011281 has reached or surpassed the implementation phase and then product may be released. | 12-Mar-2019 | 12-Mar-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | 8 | Syringe
NC-IL081-09907 | 12-Mar-2019 | GRACE.AUSTIN | General | Internal |  |  |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | PR-Planned Deviation | Run at Risk awaiting CDP-0011281 | 19C09063 | 230400.0 | Product will be placed on hold until the CDP-11281 reached or surpassed the implementation phase and then product ;may be released. | 12-Mar-2019 | 12-Mar-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | 8 | Syringe
NC-IL081-09928 | 12-Mar-2019 | GRACE.AUSTIN | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | CFF-V-Packaging 1 | Missing lids, was packed wrong quantity. | 19B17363 | 300.0 | Quarantined all affected product. Instructed associate to read and follow work instructions. Minda Sailor. | 12-Mar-2019 | 12-Mar-2019 | CFF-V-Packaging 1 |  | Human Error | Human Error | Rework | 300.0 | 0.0 | Closure | CLOSED | 11-May-2019 | 27-Mar-2019 | 15. | G2 | Container
NC-IL081-09975 | 13-Mar-2019 | MARK.VENTURA | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SA-V-Mixed Product 1 | Improper line clearance. Soft-pack from the previous lot were found in the packing area of Line 8 during production of the current lot. | 1181200777 | 61920.0 | Completed line clearance. J. Pixius 04/18/19 | 13-Mar-2019 | 13-Mar-2019 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Human Error | Human Error | Sort | 61920.0 | 0.0 | Closure | CLOSED | 12-May-2019 | 19-Apr-2019 | 37. | 8 | Syringe
NC-IL081-10057 | 15-Mar-2019 | JACLYN.GRISKO | General | Internal | S7C | 7C-IL081 |  |  |  |  | Calibration Out of Tolerance | Out of tolerance/to be repaired. "Bad Load Cell" Error messages displayed during full calibration. |  |  | Production is not at risk. \n1. Reached out to vendor for replacement load cell. \n2. Production has switched to jobs that do not require the load cell until this issue has been resolved. | 14-Mar-2019 | 23-Feb-2019 | Calibration Out of Tolerance | Calibration Out of Tolerance | maintenance | Technician installed the incorrect load cell |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 520.
NC-IL081-10087 | 16-Mar-2019 | MARK.VENTURA | General | Internal | S6K | Kahle-IL081 | Room ID | PF Sub-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-F-Spraymation | Spraymation exceeded upper specification limits | 19C07663X | 335000.0 | Changed head 5. Justin Pixius | 16-Mar-2019 | 15-Mar-2019 | SFF-SA-F-Spraymation |  | PM Frequency |  | Other/NDF | 335000.0 | 0.0 | Closure | CLOSED | 15-May-2019 | 24-Apr-2019 | 39. | Sub | Syringe
NC-IL081-10114 | 18-Mar-2019 | GRACE.AUSTIN | General | Internal | C32 | Press 32-IL081 |  |  | 851301 | 5QT ALWAYS OPEN INROOM -T RED | CFF-V-Packaging 1 | Missing IFU in one Finish Good case. | 19A26963 | 500.0 | Did rework at press, placed IFU in 1 case that was missing it. - Steve Schirmer 03/19/19 | 18-Mar-2019 | 18-Mar-2019 | CFF-V-Packaging 1 |  | Human Error | Human Error | Rework | 500.0 | 0.0 | Closure | CLOSED | 17-May-2019 | 10-Apr-2019 | 23. | G5 | Container
NC-IL081-10120 | 18-Mar-2019 | BERTHA.HERNANDEZ | General | Internal | C89 | Press 89-IL081 |  |  | 8605U | 5 QT UNIVERSAL WASTE IN ROOM | CFF-V-Quantity per Case | Quantity per case, missing instruction sheet in F/G. | 19B22863 | 504.0 | Press stopped. | 18-Mar-2019 | 18-Mar-2019 | CFF-V-Quantity per Case |  | Human Error | Human Error | Rework | 504.0 | 0.0 | Closure | CLOSED | 17-May-2019 | 29-Apr-2019 | 41. | G5 | Container
NC-IL081-10122 | 18-Mar-2019 | FIDELIA.CANTOR | Receipt | Supplier | Incoming Inspection | Multivac 5-NE017 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Loose Particulate | During Incoming inspection of 4001204, lot 19C15, skids 1-8 on 3/18/19, white particles were found on 1 rubber tip and "reflective" particles were found on 7 rubber tips. Eight defects were identified in total, failing the inspection, AQL of Normal Level II, 0.25 AQL, A=7, R=8 for cleanliness of rubber tip- loose foreign matter, particles. | 19C15-1-8 | 2880000.0 | Product contained and put on hold. Supplier Quality Notification 0028 issued to supplier. | 18-Mar-2019 | 18-Mar-2019 | Visual- Loose Particulate | Visual- Loose Particulate | Procedure/Method | Procedure/Method. Refer to SCAR CAPA-MA-029-00916 | Return to Vendor | 2880000.0 | 0.0 | Closure | CLOSED | 17-Apr-2019 | 15-Apr-2019 | 27. | Tips | Syringe
NC-IL081-10124 | 18-Mar-2019 | DANIEL.CALLAHAN01 | General | Internal | S1P | Packer-IL081 |  |  | 8881112518 | 12ML KK LLK SMOOTH TIP BULK NS | PR-Other | Lot size of 8881112518 code is at 504,000 and sampling plans are defined up to 500,000. | 19C08653X | 504000.0 | Assessment of # of samples taken during production of lot and verifying against ANSI Z1.4 standard for appropriate AQL. See attached memo. | 18-Mar-2019 | 16-Mar-2019 | INSAMP | Incorrect Sampling | PR-Personnel | Personnel-MEDMFGPR | UseAsIs | 504000.0 | 0.0 | Closure | CLOSED | 05-Apr-2019 | 01-Apr-2019 | 13. | 1 | Syringe
NC-IL081-10155 | 19-Mar-2019 | FIDELIA.CANTOR | Receipt | Supplier | Incoming Inspection | Multivac 5-NE017 |  |  | RM177 | Resin, Polypropylene, 8 IZOD Mixed Color Reprocessed | PR-Labeling Incorrect | The supplier labeled the gay lords with the incorrect PO # (820643) on the label. Correct PO # is 820970. See attachments. | 341130/350230 | 41694.0 |  | 19-Mar-2019 | 19-Mar-2019 | PR-Labeling Incorrect | Labeling Incorrect-MEDMFGPR | Process/Method | Process/Method. Salesman cloned and old order and did not enter a new order. This carried the old PO# forward to the new order. | Return to Vendor | 41694.0 | 0.0 | Closure | CLOSED | 30-Apr-2019 | 28-Mar-2019 | 8.
NC-IL081-10173 | 20-Mar-2019 | JESSE.SOBEL | General | Internal |  |  | Room ID | PF Sub-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | Out of Validated Parameter - Parameter not to Specification | PCR parameter for stationary Thermo was out of range. | 19A14463X | 33000.0 | Clear out the line. The lot was cut short. Process Tech lowered the set temperature to 90°F to prevent the temperature to go above the validated parameters. Berenice Villatoro 03/20/2019 | 20-Mar-2019 | 20-Mar-2019 | PR-Process Settings | Process Settings-MEDMFGPR | Human Error | Human Error - Set temperature for the stationary thermolater was manually changed by unknown personnel but the value was not verified against PCR range nor documented. | Other/NDF | 33000.0 | 0.0 | Closure | CLOSED | 15-May-2019 | 02-May-2019 | 43. | Sub | Syringe
NC-IL081-10237 | 21-Mar-2019 | GRACE.AUSTIN | General | Internal | C67 | Press 67-IL081 |  |  | 8611U | 12 GA UNIVERSAL WASTE HINGED | CFF-V-Quantity per Case | Quantity per case, missing lids in one FG case. | 19B23763 | 80.0 | Performed Checked Back. Place affected product on Hold. | 21-Mar-2019 | 21-Mar-2019 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | Human Error | Rework | 80.0 | 0.0 | Closure | CLOSED | 18-Aug-2019 | 22-Jul-2019 | 123. | G12 | Container
NC-IL081-10248 | 21-Mar-2019 | ROSEMARY.TEKSTAR | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | CFF-V-Packaging 1 | Missing four lids from finished case. Acceptable failure amount: 0 per/DC00032963. | 19B17463 | 960.0 | Quarantine product. Reworked at press. Instructed associate to read and follow quality procedure available at the press. Minda Sailor | 21-Mar-2019 | 21-Mar-2019 | CFF-V-Packaging 1 |  | Human Error | Human Error | Rework |  | 0.0 | Closure | CLOSED | 20-May-2019 | 27-Mar-2019 | 6. | G2 | Container
NC-IL081-10316 | 23-Mar-2019 | LAKEYSHA.SIMPSON | General | Internal |  |  |  |  | WP8506 | 8507SA LID | CFF-V-SHORT | Shorts found on R7502E (door) in a non-functional area on assembled lid WP8506 | 19B51563 | 480.0 |  | 23-Mar-2019 | [7 empty] | 0.0 | Initiation Verification | VOID |  |  |  | SS | Container
NC-IL081-10317 | 23-Mar-2019 | LAKEYSHA.SIMPSON | General | Internal | C18 | Press 18-IL081 | Equipment ID | R7502E-IL081 | WP8506 | 8507SA LID | CFF-V-SHORT | Shorts found on R7502E (door) in a non functional area on assembled lid WP8506. | 18B51563' '19B51563 | 420.0 2400.0 | Technician adjusted the cut off. F.Henry 03/23/19 | 23-Mar-2019 | 23-Mar-2019 | CFF-V-Short | Part is not completely formed; Missing plastic | Materials | Materials | Sort | 420.0 2400.0 | 0.0 | Closure | CLOSED | 30-Aug-2019 | 21-Aug-2019 | 151. | SS | Container
NC-IL081-10319 | 23-Mar-2019 | BERTHA.HERNANDEZ | General | Internal | S3P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-PPSP-V-Labeling/Print 4 | Failed UDI inspection. Inspector had 2 inner cartons with "F" grade during scheduled inspection. | 19C04063X | 144.0 | Changed ink cartridge. | 23-Mar-2019 | 23-Mar-2019 | SFF-PPSP-V-Labeling/Print 4 |  | Bracket |  | Rework | 144.0 | 0.0 | Closure | CLOSED | 22-May-2019 | 15-Apr-2019 | 23. | 3 | Syringe
NC-IL081-10323 | 23-Mar-2019 | BERTHA.HERNANDEZ | General | Internal | S8B | 8B-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 2 | Flash (string) in functional area exceeds 1/32" on barrel tip, press MP1133/58119. | 19A14863X' '19A14763X | 255000.0 264000.0 9600.0 | Cavities were blocked. Line cleared of all barrels. | 23-Mar-2019 | 22-Mar-2019 | SFF-M-V-Flash/Gate 2 |  | Core Shift |  | Scrap' 'Sort | 9600.0 255000.0 264000.0 | 0.0 | Closure | CLOSED | 22-May-2019 | 02-May-2019 | 39. | Sub | Syringe
NC-IL081-10346 | 25-Mar-2019 | BERTHA.HERNANDEZ | General | Internal | S8B | 8B-IL081 | Equipment ID | 58119-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Bubbles 3 | Bubble in non-fluid pathway (thread area). While reworking product for flash a bubble was discovered on the thread area, AQL was performed and failed. | 19A14863X' '19A14963X' '19A14763X | 335000.0 147000.0 | Pulled tool for repair under work order number CL3072740. - P. Myerson 03/25/2019 | 24-Mar-2019 | 22-Mar-2019 | SFF-M-V-Bubbles 3 |  | Communication error | Communication Error | UseAsIs | 335000.0 147000.0 | 0.0 | Closure | CLOSED | 24-May-2019 | 29-Apr-2019 | 35. | Sub | Syringe
NC-IL081-10397 | 26-Mar-2019 | GRACE.AUSTIN | General | Internal | C52 | Press 52-IL081 |  |  | 31378089 | 19 GA SHARPS CONTAINER | CFF-V-Packaging 1 | Incorrect Quantity per case. Missing 3 ea. of lids on one Finish Good Case. | 19B20863 | 120.0 | Quarantined affected product and rework completed at the press. Instructed associate to read and follow quality procedure available at the press. - M. Sailor 03/26/19 | 26-Mar-2019 | 26-Mar-2019 | CFF-V-Packaging 1 |  | Human Error | Human Error | Rework | 120.0 | 0.0 | Closure | CLOSED | 25-May-2019 | 29-Apr-2019 | 34. | G18 | Container
NC-IL081-10414 | 26-Mar-2019 | FIDELIA.CANTOR | Receipt | Supplier | Incoming Inspection | Multivac 5-NE017 |  |  | RM001 | RESIN,HIPS,ATOFINA #825 | PR-Labeling Incorrect | Incorrect PO # (820887) on the product label and BOL for RM001. Correct PO # is 820928. | CL50026856 | 144000.0 | Placed product on hold under BPCS hold #19030055 | 26-Mar-2019 | 26-Mar-2019 | Label information incorrect | Label information incorrect | Human Error | Human Error. Account Coordinator entered incorrect customer PO on the order. | Return to Vendor | 144000.0 | 0.0 | Closure | CLOSED | 31-Jul-2019 | 04-Sep-2019 | 161.
NC-IL081-10442 | 27-Mar-2019 | MARIA.ROSALES | General | Internal | S8P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-F-Spraymation | Spraymation exceeded upper control limits, excessive lubricant in the syringe. Performed AQL for visible excessive lubricant in the syringe and failed. | 19C16763X | 335000.0 | Line stopped. Kahle was cleaned. Work order CL-3072835 completed to replace silicone gun head number 11. - L. Reed 03/27/2019 | 27-Mar-2019 | 27-Mar-2019 | Assembly Defect - Silicone - Excessive Application | Assembly Defect - Silicone - Excessive Application | Machine Failure | Machine Failure | Scrap | 335000.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 19-Jun-2019 | 84. | Sub | Syringe
NC-IL081-10594 | 29-Mar-2019 | ROSEMARY.TEKSTAR | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Mixed Product 1 | Mixed product sent under one lot number. | 19B40063X | 335000.0 | All material contained at internal customers warehouse. | 29-Mar-2019 | 01-Mar-2019 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Handling/Storage | Handling/Storage | Sort' 'Scrap' 'UseAsIs | 96000.0 50836.0 188164.0 | 0.0 | Closure | CLOSED | 10-May-2019 | 03-Jul-2019 | 95. | Sub | Syringe
NC-IL081-10676 | 02-Apr-2019 | MARIA.ROSALES | General | Internal | S2P | Packer-IL081 | Room ID | PF 2 or 3-IL081 | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-PPHP-V-Lot Number 3 | Incorrect lot # on cap, 15H08263X should be 19C19063X | 19C19063X | 34560.0 | We reprogrammed cap side A, and verified the correct info on it before restarting the packer. D. Breceda 04/02/2019 | 02-Apr-2019 | 01-Apr-2019 | SFF-PPHP-V-Lot Number 3 |  | Human Error | Human Error | Sort | 34560.0 | 0.0 | Closure | CLOSED | 02-May-2019 | 30-Apr-2019 | 28. | 3 | Syringe
NC-IL081-10690 | 02-Apr-2019 | JACLYN.GRISKO | General | Internal | C54 | Press 54-IL081 |  |  | 8536SA | 3 GA SHARPSTAR - CLEAR | CFF-D-Molding 1 | blah | TST011519 | 100000.0 | Training tool | 02-Apr-2019 | [7 empty] | 0.0 | Initiation Verification | VOID |  |  |  | SS | Container
NC-IL081-10738 | 03-Apr-2019 | LAKEYSHA.SIMPSON | General | Internal | C40 | Press 40-IL081 |  |  | WP8964 | LID AND ROTOR ASSEMBLY | Documentation Missing | Technician Crew 3 failed to verify PCR Parameter and document readings once per shift per GS100004646 section 6.3.1 | 19C187658 | 720.0 | PCR reading were done and all reading were in parameter. All affected product contain. | 02-Apr-2019 | 02-Apr-2019 | Documentation Incomplete/Missing | Documentation Incomplete/Missing | Human Error | Human error- Process tech failed to perform process inspection per GS10004646, sec#6.3. Human error investigation revealed, a startup inspection was performed late in the night shift at 4:58 am and day shift tech assumed that covers day crew inspection requirements. His misunderstanding of different procedural requirements between product and process inspections led to skipping inspection. | Other/NDF | 720.0 | 0.0 | Closure | CLOSED | 30-Aug-2019 | 21-Aug-2019 | 140. | G2 | Container
NC-IL081-10766 | 03-Apr-2019 | ROSEMARY.TEKSTAR | General | Internal | C31 | Press 31-IL081 | Equipment ID | R7028M-IL081 | 8970 | 2 GA ROTOR LID CONTAINER - RED | CFF-V-Warping, Deformation and Pulling | Deformed rim and bowed Bottom of containers. The containers rock. | 19B33163 | 4440.0 | Line stopped, process technician changed parameters to eliminate bowed bottom on Cavity 2. Once parts were acceptable, the press was restarted. | 03-Apr-2019 | 03-Apr-2019 | CFF-V-Warping, Deformation and Pulling |  | Poppet |  | Scrap | 4440.0 | 0.0 | Closure | CLOSED | 02-Jul-2019 | 20-Jun-2019 | 77. | G2 | Container
NC-IL081-10790 | 04-Apr-2019 | ALEJANDRA.CASTANEDA | General | Internal | 171 | Cell 171-IL081 |  |  | WP8536 | SharpStar II Lid, Door and Cowl | CFF-V-Short | Short on R7460C Cavity# 2, WP8536 lids used to Finish Good product 8536SA | 19A16963 | 1650.0 10800.0 | Process adjustments made were unsuccessful, so we pulled the R7460 tool and set the R7460B tool. Justin Pixius 04/04/19 | 04-Apr-2019 | 04-Apr-2019 | CFF-V-Short | Part is not completely formed; Missing plastic | Tooling/ fixture wear | Wear on tooling causes variation-\nSome of the vents inside mold surface washed out. A lack of air escapement was causing the trapped air pocket in cavity-2, which hindered the thermoforming and not allowing mold closure to be completed properly. Also, the process parameters have not been validated and the specified limit seems to have fast fill time that might be a contributing factor. | Sort | 10800.0 1650.0 | 0.0 | Closure | CLOSED | 30-May-2019 | 24-Jun-2019 | 81. | SS | Container
NC-IL081-10849 | 05-Apr-2019 | MARK.VENTURA | General | Internal |  |  |  |  | 8881512852 | Syringe Reg Tip 12ml Monoject 80/bx 480/cs | SFF-PPHP-V-Lot Number 3 | Incorrect print on CAP | 19C23263X | 3840.0 | Reprogrammed the correct information on the caps and verified per lot control. - Kim Kaecker 04/07/19 | 05-Apr-2019 | 05-Apr-2019 | SFF-PPHP-V-Lot Number 3 |  | Human Error | Human Error | Scrap | 3840.0 | 0.0 | Closure | CLOSED | 04-Jun-2019 | 15-Apr-2019 | 10. | 3 | Syringe
NC-IL081-10858 | 05-Apr-2019 | FIDELIA.CANTOR | Receipt | Supplier | Incoming Inspection | Multivac 5-NE017 |  |  | EXP6029 | Thermal Transfer Ribbon, General Wax, 4.33 Inch | Procedural- Procedure not updated | Ink Ribbon is inappropriately categorized at its inception as (Foil- Syringe), which resulted in an inappropriate level of incoming inspection., | B229 | 144.0 | Product was put on hold until further research could be completed. | 05-Apr-2019 | 05-Apr-2019 | Procedural- Procedure not updated | Procedural- Procedure not updated | PR-Documentation | This event was the result of incomplete information received from a plant that was closing, and therefore led to miscategorization. | UseAsIs | 144.0 | 0.0 | Closure | CLOSED | 05-May-2019 | 29-Apr-2019 | 24.
NC-IL081-10873 | 07-Apr-2019 | ANGELICA.AGUILAR | General | Internal | C52 | Press 52-IL081 |  |  | V8870 | NOVAPLUS 18 GA HINGED PHARMA | CFF-V-Label Verification and Barcode Verification 2 | One side printed Incorrect lot number on the biohazard label. | 19B50363 | 535.0 | Turn off labeler and reprogram. Place all affected product on Hold. | 07-Apr-2019 | 07-Apr-2019 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Process/Method | Process/Method-The method used to verify the formatting of the information was not present at the time of this printer programming. At the time of this defect the labeling system was still on the Medtronic server / network and the information was able to be manually entered . The Medtronic server and network did not provide a verification system that would have caught the extra digit being programmed into the system. Upgrades to the label printing software were scheduled to be implemented after the Cardinal transition. Since then the labeling system has been converted over to the Cardinal server and network. During that changeover changes were made to the lot# field on Bartender web site which would have prevented the lot number defect noted on this NCR. The changes verify the alpha numeric format and number of digits allowed in the lot number field.) | Rework | 535.0 | 0.0 | Closure | CLOSED | 27-Sep-2019 | 09-Sep-2019 | 155. | G18 | Container
NC-IL081-10883 | 08-Apr-2019 | ALEJANDRA.CASTANEDA | General | Internal | C33 | Press 33-IL081 |  |  | WP7445 | 8 GA HEAT STK LID RAILS DOORS | Assembly Defect - Incorrect Assembly | Assembled wrong. Sliding door with the Notch put on Backwards. | 19B187524 | 760.0 | Ensured doors were being welded correctly. Jr Cruz 04/08/19 | 08-Apr-2019 | 08-Apr-2019 | Assembly Defect - Incorrect Assembly | Assembly Defect - Incorrect Assembly | Human Error | Human Error - Associate assembled lids incorrectly and did not realize during their rotation period. Defect was identified by following associate during a non-scheduled audit. Associate admitted, he failed to read stand work instruction prior to start of task. | Sort | 720.0 40.0 | 0.0 | Closure | CLOSED | 30-Aug-2019 | 26-Aug-2019 | 140. | G8 | Container
NC-IL081-10924 | 09-Apr-2019 | JESSE.SOBEL | General | Internal | S1B | 1B-IL081 | Room ID' 'Equipment ID | PF 1-IL081' '58140-IL081 | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-M-V-Bubbles 1 | Gas bubble in fluid pathway. May be interpreted as bubble in content | 19D04463X | 92000.0 | Cavity 42 was blocked CL-3077952. Removed all affected parts from line. Initiated a lot change. Berenice Villatoro 04/09/19. | 09-Apr-2019 | 08-Apr-2019 | Bubbles | Bubbles - Cavity specific #42 | Blocked Vents | A clogged/non-vented core pin on cavity 42 led to excessive decompression resulted in cavity specific bubbles. | Sort | 92000.0 | 0.0 | Closure | CLOSED | 03-May-2019 | 13-May-2019 | 34. | 1 | Syringe
NC-IL081-10930 | 09-Apr-2019 | ALEJANDRA.CASTANEDA | General | Internal | C66 | Press 66-IL081 |  |  | 31156550 | 7 GA SHARPS CONTAINER | CFF-V-Broken, Cracking, and Crazing | Broken locking tap on WP4810, lids were packed into F/G 31156550 | 19C11663 | 300.0 | Polished and deburred tool. G. Cruz 04/09/19 | 09-Apr-2019 | 09-Apr-2019 | CFF-V-Broken, Cracking, and Crazing | Broken, Cracking, and Crazing all surfaces | Mold Wear or Damage | Mold Wear or Damage - R7579C mold was found to have burr inside of the pocket which caused the plastic to remain and stay in the front half. When the mold closed on the next cycle the plastic was harden and cause the lug to not insert into the pocket and there for snapped the lug off | Scrap | 300.0 | 0.0 | Closure | CLOSED | 31-Aug-2019 | 28-Aug-2019 | 141. | G7 | Container
NC-IL081-11021 | 11-Apr-2019 | MARIA.ROSALES | General | Internal | S8K | Kahle-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | Out of Validated Parameter | Incorrect setting used for Spraymation ( 47 psi ) spec is 50-60, and Temp.( 160 ) spec. 185-225 | 19D10163X | 29280.0 | Stopped line. Printed out new PCR for Kahle and set all areas to within parameters. Retested spraymation and had passing results. - L. Reed 04/11/2019 | 10-Apr-2019 | 10-Apr-2019 | Out of Validated Parameter |  | Human Error | Human Error | Other/NDF | 29280.0 | 0.0 | Closure | CLOSED | 30-Jun-2019 | 24-Jun-2019 | 74. | 8 | Syringe
NC-IL081-11061 | 12-Apr-2019 | JESSE.SOBEL | General | Internal |  |  |  |  | 8881112059 | 12ML SYRINGE REGULAR NS | PR- Mixed product | Mixed Product | 18K12963X | 1.0 |  | 11-Apr-2019 | [7 empty] | 0.0 | Initiation Verification | VOID |  |  |  | 7 | Syringe
NC-IL081-11092 | 12-Apr-2019 | ROSEMARY.TEKSTAR | General | Internal | C63 | Press 63-IL081 |  |  | WP7417 | Hot Air Staker Chemotherapy Lid, 12/18 Gallon | CFF-F-Assembly 2 | Failed leak test in positions A & B per requirement LS10000076 Rev 12. | 187832 | 3225.0 | Start/Stop position was moved to corner. Ryan Syafrin 04/12/19 | 12-Apr-2019 | 12-Apr-2019 | CFF-F-Assembly 2 | Leak Test Failure | Blockage | Blockage of Iso flow due to kinks and twists on the Iso hose that prevents flow | Scrap | 3225.0 | 0.0 | Closure | CLOSED | 28-Jun-2019 | 24-Jun-2019 | 72. | G12 | Container
NC-IL081-11116 | 13-Apr-2019 | BERTHA.HERNANDEZ | General | Internal | C80 | Press 80-IL081 |  |  | 85321R | 2GA ALWAYS OPEN INROOM -T RED | CFF-V-Packaging 2 | Overshipper is not packaged as per packaging standard, found during QA F/G audit. Step #2 of FM20007222 was not followed. | 19A26063 | 2800.0 | Shut down press ,contained affected product and retrained associates per FM20007222 Rev 0. Minda I. Sailor 04/13/19. | 13-Apr-2019 | 12-Apr-2019 | CFF-V-Packaging 2 |  | Human Error | Human Error | Rework | 2800.0 | 0.0 | Closure | CLOSED | 12-Jun-2019 | 29-Apr-2019 | 16. | G2 | Container
NC-IL081-11162 | 16-Apr-2019 | MARIA.ROSALES | Inspection | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SSF-PPSP-VLabeling/Print4 | Failed LVS Barcode verification on over shipper labels. | 19D10363X | 84960.0 | Line was stopped. New overs hipper labels were printed and verified to be passing. Labeler changed out in label cage. Work Order CL-3078498. - L. Reed 04/16/2019 | 15-Apr-2019 | 15-Apr-2019 | SSF-PPSP-VLabeling/Print4 |  | Printer Malfunction | Printer Malfunction | Rework | 84960.0 | 0.0 | Closure | CLOSED | 24-May-2019 | 17-May-2019 | 31. | 8 | Syringe
NC-IL081-11184 | 16-Apr-2019 | BERTHA.HERNANDEZ | General | Supplier | STT | Tip Tumbling-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Water in package/on tip | Extraneous matter, water is present in rubber tip bag | 19D12-1 | 90000.0 | Placed product on hold. | 16-Apr-2019 | 16-Apr-2019 | Visual- Water in package/on tip | Visual- Water in package/on tip | Documentation | Documentation. The Lot card form used by the operators was insufficient (lacking detailed information that can affect part quality) | Return to Vendor | 90000.0 | 0.0 | Closure | CLOSED | 05-Jul-2019 | 02-Jul-2019 | 76. | Tips | Syringe
NC-IL081-11208 | 16-Apr-2019 | BERTHA.HERNANDEZ | General | Internal | S8F | Form, Fill Sealer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-PPSP-V-Seal Integrity 1 | Seal integrity, void allows air into sealed package on cavities 14 & 15. | 19D10363X | 7200.0 | Replaced the gasket. - P. Myerson 04/16/2019 | 16-Apr-2019 | 16-Apr-2019 | SFF-PPSP-V-Seal Integrity 1 | Void allows air into sealed package | Damaged gasket | Damaged Gasket | Sort | 7200.0 | 0.0 | Closure | CLOSED | 15-Jun-2019 | 13-Jun-2019 | 57. | 8 | Syringe
NC-IL081-11246 | 17-Apr-2019 | BERTHA.HERNANDEZ | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SPBP-V-Labeling/Print 4 | Labeling/print- incorrect expiration date on overshipper labels. | 19D10363X | 480.0 | Defect found during end of run. Line was down for product change. Reprint labels with the correct expiration date. - M. Buck 04/24/2019 | 17-Apr-2019 | 16-Apr-2019 | SFF-SPBP-V-Labeling/Print 4 | Expiration Date is incomplete or missing (where applicable) | Human Error | Human Error | Rework | 480.0 | 0.0 | Closure | CLOSED | 17-May-2019 | 17-May-2019 | 30. | 8 | Syringe
NC-IL081-11330 | 21-Apr-2019 | MARIA.ROSALES | Inspection | Internal | C56 | Press 56-IL081 |  |  | 8997 | 8 GA GASKETED HINGED LID -RED | CFF-V-Label Verification and Barcode Verification 3 | Missing lot # on over shipper label. | 19B49863 | 320.0 | Removed and discard all labels with incorrect Print. New set of labels printed and verified to ensure accuracy by reprogramming labeler. | 20-Apr-2019 | 20-Apr-2019 | CFF-V-Label Verification and Barcode Verification 3 | OS - Bad/Missing label print | IT/Computer Systems | Associate used a Wi-Fi enabled labeling computer (ELS_WS07) to print over shipper label. Using Wi-Fi instead of hardwire data port connection was causing poor signal strength. As a result, transmitting data was interrupted and the printer was not receiving full data to process the label print. Partially loss of input data caused the missing GTIN print. | Rework | 320.0 | 0.0 | Closure | CLOSED | 19-Aug-2019 | 19-Jul-2019 | 89. | G8 | Container
NC-IL081-11446 | 24-Apr-2019 | GRACE.AUSTIN | General | Internal | C67 | Press 67-IL081 |  |  | 8931PG2 | SHARPS CONTAINER 12GAL YELLOW HINGED LID | CFF-V-Packaging 1 | Missing one package of absorbent Pad, required 2 package per BOM. | 19B30563 | 160.0 |  | 24-Apr-2019 | 24-Apr-2019 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | Human Error | Rework | 160.0 | 0.0 | Closure | CLOSED | 23-Jul-2019 | 22-Jul-2019 | 89. | G12 | Container
NC-IL081-11479 | 25-Apr-2019 | GRACE.AUSTIN | General | Internal | C81 | Press 81-IL081 |  |  | 89671 | 2 GA HORIZONTAL DROP - T RED | CFF-V-Label Verification and Barcode Verification 4 | Missing Pad Print, Missing Barcode, Product Code, Lot number and UDI on Biohazard Label. | 19B32863 | 40.0 | Loaded Program and changed Ink cartridges. | 25-Apr-2019 | 25-Apr-2019 | CFF-V-Label Verification and Barcode Verification 4 |  | Power loss |  | Sort | 40.0 | 0.0 | Closure | CLOSED | 24-Jul-2019 | 05-Jun-2019 | 41. | G2 | Container
NC-IL081-11516 | 26-Apr-2019 | BENJAMIN.KELLY | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-F-Spraymation | Spraymation exceed upper control limit. | 19C13863X | 335000.0 | Cleaned spraymation heads. - J. Pixius 06/24/2019 | 26-Apr-2019 | 26-Apr-2019 | SFF-SA-F-Spraymation | Spraymation | Cause Unknown | Cause Unknown | Scrap | 335000.0 | 0.0 | Closure | CLOSED | 25-Jul-2019 | 19-Jul-2019 | 84. | Sub | Syringe
NC-IL081-11526 | 26-Apr-2019 | GRACE.AUSTIN | General | Internal | C56 | Press 56-IL081 |  |  | 8997S | 8 GA GASKETED SLIDE LID - RED | CFF-V-Quantity per Case | Missing absorbent pad and IFU. | 19B50063 | 240.0 | Checkback was performed by QAT, AQL was done and passed. Went over standard work with associate. - S. Schirmer 05/08/2019 | 26-Apr-2019 | 26-Apr-2019 | CFF-V-Quantity per Case |  | Human Error | Human Error | Rework | 240.0 | 0.0 | Closure | CLOSED | 25-Jun-2019 | 05-Jun-2019 | 39. | G8 | Container
NC-IL081-11558 | 27-Apr-2019 | BEATRIZ.GUILLEN | General | Internal | C85 | Press 85-IL081 |  |  | 85031 | 2 QT MAILBOX INROOM - T RED | CFF-V-Label Verification and Barcode Verification 10 | Missing biohazard label on one side of container | 19B09963 | 640.0 | Contained affected product, re-trained associate on FM20008303 Rev 2, section 2.B. - Minda Sailor 04/27/2019 | 27-Apr-2019 | 27-Apr-2019 | CFF-V-Label Verification and Barcode Verification 10 |  | Human Error | Human Error | Rework | 640.0 | 0.0 | Closure | CLOSED | 26-Jun-2019 | 29-Apr-2019 | 2. | HG | Container
NC-IL081-11731 | 01-May-2019 | GRACE.AUSTIN | General | Internal | C67 | Press 67-IL081 |  |  | 8935Y | 12 GA NEEDLE DISP CONT YEL 10 | CFF-V-Packaging 1 | Missing 5 ea. lids on one FG box | 19C17763 | 10.0 | Rework at the Press, add 5 ea. lid on affected FG box. | 01-May-2019 | 01-May-2019 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | Human Error - The Manufacturing Associate did not follow all machine setup and operating procedures. The formula sheet for 8935Y called for the use of a scale to verify the quantities packed in the finished goods boxes at press 67 per DC00032963. The operator packed finished goods boxes without using the scale verification. It was verified that the formula sheet for 8935Y does call out for the setup and use of a scale through FM20013285. | Rework | 10.0 | 0.0 | Closure | CLOSED | 20-Sep-2019 | 26-Sep-2019 | 148. | G12 | Container
NC-IL081-11749 | 02-May-2019 | ALEJANDRA.CASTANEDA | General | Internal | C70 | Press 70-IL081 |  |  | 8906 | 1QT N/S CONTAINER YELLOW 100 | PR-Parameters out of specification | Parameters out of specification on the PCR ( hold time and hold pressure) | 19C25563 | 3500.0 | Hold Time and Hold Pressure was changed to bring back the process within validated limit. | 02-May-2019 | 01-May-2019 | PR-Parameters out of specification | Parameters out of specification on the PCR (hold time and hold pressure) | Human Error | Human error- Process tech failed to change/adjust settings following the PQ run and prior to startup on regular production. He was rushing to complete process inspection and missed the task due to not verifying actual values against current validated PCR (DC#150821-2878). | Other/NDF | 3500.0 | 0.0 | Closure | CLOSED | 31-Jul-2019 | 17-Jul-2019 | 76. | G1 | Container
NC-IL081-11876 | 04-May-2019 | BENJAMIN.KELLY | General | Internal | S6A | 6A-IL081 | Equipment ID | 58110-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Shorts 1 | Out Of Spec. Visual - Shorts in fluid pathway, Cav #16 | 19D24863X | 13000.0 | Cleared out the line. Blocked cavity 16. Scrapped cases 20-32. CL-3084865 - B. Villatoro 05/04/2019 | 04-May-2019 | 04-May-2019 | SFF-M-V-Shorts 1 |  | Mold Wear or Damage | Mold#58110 found to have a bad cavity block with clogged gate pin. It was failing to transfer hit during molding process resulted in short shot on cavity#16. | Scrap | 13000.0 | 0.0 | Closure | CLOSED | 30-Jun-2019 | 24-Jun-2019 | 50. | Sub | Syringe
NC-IL081-11877 | 05-May-2019 | BERTHA.HERNANDEZ | General | Internal | S6C | 6C-IL081 | Equipment ID | 58114-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 1 | Drag mark on barrel that can be felt with bare finger on cavity 21B. | 19D24863X' '19D24763X' '19D24663X | 200000.0 335000.0 | Cavity 21B was blocked. All work in process parts from 6C were scrapped. - M. Islam 05/05/2019 | 05-May-2019 | 02-May-2019 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 1 | Material removed - Depth is measureable. Scratch inside Fluid Pathway - No Defects Allowed (Scratches/Scuffs on the inside of the barrel are not acceptable; Drag Marks that can be felt with a bare finger are not acceptable; Drag marks extended along the entire barrel length that can be felt with a bare finger) | Equipment | Equipment | UseAsIs | 200000.0 335000.0 | 0.0 | Closure | CLOSED | 21-Jun-2019 | 26-Sep-2019 | 144. | Sub | Syringe
NC-IL081-11941 | 07-May-2019 | JACLYN.GRISKO | General | Internal | S7P | Packer-IL081 | Equipment ID | 58002-IL081 | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SA-F-Spraymation | Group average exceeding upper specification limits on head A. | 123456789 | 1.0 | Enter the line start up process correction here. - Name and date of who wrote it. | 07-May-2019 | 03-May-2019 |  |  |  |  | Scrap |  | 0.0 | Investigation | VOID |  |  |  | 3 | Syringe
NC-IL081-12016 | 09-May-2019 | GRACE.AUSTIN | General | Internal | C31 | Press 31-IL081 |  |  | 8602RC | 2 GA RCRA HAZARDOUS WASTE CONT | CFF-V-Label Verification and Barcode Verification 3 | Missing Print GTIN and Lot number on Overshipper Label. | 19B22563 | 2240.0 | Reprogrammed the Overshipper Labeler. | 09-May-2019 | 09-May-2019 | CFF-V-Label Verification and Barcode Verification 3 | OS - Bad/Missing label print | IT/Computer Systems | IT/Computer Systems: Associate used a Wi-Fi enabled labeling computer (ELS_WS07) to print over shipper label. Using Wi-Fi instead of hardwire data port connection was causing poor signal strength. As a result, transmitting data was interrupted and the printer was not receiving full data to process the label print. Partially loss of input data caused the missing GTIN print. | Rework | 2240.0 | 0.0 | Closure | CLOSED | 06-Sep-2019 | 19-Jul-2019 | 71. | G2 | Container
NC-IL081-12031 | 09-May-2019 | BERTHA.HERNANDEZ | General | Internal | S8P | Packer-IL081 |  |  | 8881540122 | 12ML SMARTIP MED-PREP CANNULA | Sterilization Process Failure | Incorrect loading pattern used for totes at sterilizer. | 19B38563X' '19A32763X | 172320.0 225520.0 | Sterilizer notified us of error and are performing dose value certification. - J. Grisko 05/10/2019 | 09-May-2019 | 23-Apr-2019 | Sterilization Parameters Out of Specification | Sterilization Parameters Out of Specification | People/Human Error/ | People/Human Error/ | No Issue / No Impact' 'Saleable | 225520.0 172320.0 | 0.0 | Closure | CLOSED | 31-May-2019 | 31-May-2019 | 22. | 8 | Syringe
NC-IL081-12078 | 10-May-2019 | ROSEMARY.TEKSTAR | General | Internal | C52 | Press 52-IL081 |  |  | 8939PG2 | SHARPS CONTAINER 18GAL SLIDE- | CFF-V-Flash (Non-Functional Area), Gate | Sharp flash in gate area. | 19B32063 | 630.0 | Assigned additional labor/personnel to clip the gate. M. Islam 05/14/19. | 10-May-2019 | 10-May-2019 | CFF-V-Flash (Non-Functional Area), Gate | Sharp Flash; Surfaces A & B: greater than or equal to 0.080”; Surface C: greater than or equal to 0.125”; All Surfaces: more than 3 locations in total | Tooling/ fixture damaged | Valve gate was not shifting consistently during injection molding process and was causing the gate flash as result of protrusion. A worn-out piston cup was identified as a root cause by tool room during their mold evaluation under WO #CL-3078031. | Rework | 630.0 | 0.0 | Closure | CLOSED | 31-Aug-2019 | 03-Sep-2019 | 115. | G18 | Container
NC-IL081-12088 | 11-May-2019 | MARIA.ROSALES | General | Internal | S8B | 8B-IL081 | Equipment ID | 58119-IL081 | 3150403 | 12 mL Printed MAGELLAN Barrel - P5M6K-080 | Out of Validated Parameter | Recovery time out parameters 3.76 spec is 3.8-4.2 | 19D18963X | 10000.0 | Performed new start-up on press 8B with new PCR inspection, ensuring all readings were within required parameters. - J. Williams 05/11/2019 | 11-May-2019 | 11-May-2019 | Out of Validated Parameter | Out of Validated Parameter | Human Error | The settings for Mold#58119 in 8B were not adjusted to the current MSG/PCR following a commodity change. At startup the process value was out of limit as a result. Technical operator did not properly observe the high and low limits for the charge recovery time on the process control record and failed to identify actual value is out. | Scrap | 10000.0 | 0.0 | Closure | CLOSED | 10-Jul-2019 | 26-Jun-2019 | 46.
NC-IL081-12114 | 13-May-2019 | GRACE.AUSTIN | General | Internal | C31 | Press 31-IL081 |  |  | 8982CRH | 2 GALLON CHEMO RED | CFF-V-Label Verification and Barcode Verification 3 | Missing Print GTIN and Lot number on the Overshipper Labels | 19B35063 | 1120.0 | Reprogramed the labeler. | 13-May-2019 | 13-May-2019 | CFF-V-Label Verification and Barcode Verification 3 | OS - Bad/Missing label print | IT/Computer Systems | IT/Computer Systems: Associate used a Wi-Fi enabled labeling computer (ELS_WS07) to print over shipper label. Using Wi-Fi instead of hardwire data port connection was causing poor signal strength. As a result, transmitting data was interrupted and the printer was not receiving full data to process the label print. Partially loss of input data caused the missing GTIN print. | Rework | 1120.0 | 0.0 | Closure | CLOSED | 10-Sep-2019 | 19-Jul-2019 | 67. | G2 | Container
NC-IL081-12217 | 14-May-2019 | ROSEMARY.TEKSTAR | General | Internal | C37 | Press 37-IL081 | Equipment ID | R7461D-IL081 | 8537SA | 3 GA SHARPSTAR - TRANS RED | PR-- Incorrect Raw Material | Wrong raw material being used at the press. RM-129 was being used. Shop order called for RM-093. This resulted in the wrong PCR being used as well. | 19D22863 | 49340.0 | Performed an end of run line clearance on 8537SA. A new startup was completed on a different lot after changing to the correct material. M. Islam | 14-May-2019 | 14-May-2019 | PR-- Incorrect Raw Material | Incorrect Raw Material-MEDMFGPR | Human Error | Human Error - Employees failed to verify raw material against shop order per startup/line clearance procedure GS10000122. | Other/NDF | 49340.0 | 0.0 | Closure | CLOSED | 31-Aug-2019 | 28-Aug-2019 | 105. | SS | Container
NC-IL081-12219 | 14-May-2019 | ROSEMARY.TEKSTAR | General | Internal | C46 | Press 46-IL081 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | Documentation Incomplete/Missing | In-process inspection completed after 4:00pm. | 19D22863 | 0.0 | Documentation issue only - no product impact. | 14-May-2019 | 14-May-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | SS | Container
NC-IL081-12365 | 18-May-2019 | GRACE.AUSTIN | General | Internal | C31 | Press 31-IL081 |  |  | 8967Y | 2 GA H-DROP TRNS YEL W/UNWND20 | CFF-V-Packaging 1 | Product not packed as per Standard Work Packaging, FM20007220. The Insert Divider, the half of the Flap was not inserted down to divide layers of overshipper. | 19C12063 | 340.0 | Shut Down the press and repacked the Product per Standard Work Packaging. B. Dominguez 05/18/19 | 18-May-2019 | 18-May-2019 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | Human Error - Associate failed to follow packaging instruction per applicable standard work document FM20007220. H-insert was not placed in a proper method during finished goods packaging. Employee skipped reading standard work and was working on “autopilot” due to familiar to the task. If the employee had read the work instruction and not assumed the method for packing the box, the H insert would have been inserted correctly) | Rework | 340.0 | 0.0 | Closure | CLOSED | 21-Aug-2019 | 21-Aug-2019 | 95. | G2 | Container
NC-IL081-12371 | 18-May-2019 | BERTHA.HERNANDEZ | General | Internal | C42 | Press 42-IL081 |  |  | 8527R | 3 GA ROTOR HINGED LID - T RED | Documentation Incomplete/Missing | Failed to performed a start up checklist after press been down greater than 24 hours per GS10000122 Rev#206 | 19B13063 | 360.0 | Start checklist was completed no product impact. | 18-May-2019 | 18-May-2019 | Documentation Incomplete/Missing | Documentation Incomplete/Missing | Human Error | Human Error - Backup lead failed to perform required startup with checklist per GS1000122, Sec#5.8). A miscommunication during shift change hand off led to misunderstanding of procedural requirement. | Other/NDF | 360.0 | 0.0 | Closure | CLOSED | 15-Oct-2019 | 08-Oct-2019 | 143. | G3 | Container
NC-IL081-12372 | 18-May-2019 | BERTHA.HERNANDEZ | General | Internal | C43 | Press 43-IL081 |  |  | WP8980-C | 8 Gallon Locking Rotor Lid | Documentation Incomplete/Missing | Failed to perform start-up checklist when tool was pulled per GS10000122 Rev#206 | 19C187586 | 360.0 | Start up checklist was completed. No product impact. | 18-May-2019 | 18-May-2019 | Documentation Incomplete/Missing | Documentation Incomplete/Missing | human error | Human Error - Backup lead failed to perform required startup with checklist per GS1000122, Sec#5.8). A miscommunication during shift change hand off led to misunderstanding of procedural requirement. Mold was pulled and set back in after repair which required a startup checklist, but backup lead failed to perform due to lack of knowledge. | Other/NDF | 360.0 | 0.0 | Closure | CLOSED | 15-Oct-2019 | 08-Oct-2019 | 143. | G8 | Container
NC-IL081-12406 | 20-May-2019 | BENJAMIN.KELLY | General | Internal | S7N | Pad Printer-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-P-V-Alpha Numeric 1 | Post production NCR - Missing Print SCAR # CAPA-MEX15-01101 | 18B14963X' '18B14863X' '17L23063X | 1.0 |  | 20-May-2019 | 19-May-2019 | SFF-P-V-Alpha Numeric 1 | Type 1: Illegible print or missing one unit of measure | Equipment | Equipment | Not Applicable | 1.0 | 0.0 | Closure | CLOSED | 14-Aug-2019 | 30-Jul-2019 | 70. | 7 | Syringe
NC-IL081-12456 | 21-May-2019 | GRACE.AUSTIN | General | Internal | C52 | Press 52-IL081 |  |  | 8870- | 18GA PHARMSAFETY WASTE HINGED | Incorrect Assembly - Wrong Color/Material | Colorant should be .5% and was running below. | 19C10963' '19B45963' '19B45863' '19B50363 | 4735.0 1.0 535.0 | End of Run and Lot Change to New Job with the correct color percentage. - Ken Wiltse 05/21/2019 | 21-May-2019 | 21-May-2019 | Class2 - COLOR N/SPEC | Wrong color percentage used in the formulation of the component. | Human Error | Human Error | Other/NDF | 1.0 4735.0 535.0 | 0.0 | Closure | CLOSED | 31-Jul-2019 | 26-Jul-2019 | 65. | G18 | Container
NC-IL081-12551 | 23-May-2019 | FIDELIA.CANTOR | Receipt | Supplier | Incoming Inspection | Multivac 5-NE017 |  |  | RM018 | YELLOW, 25/1,HIPS,#YHC-9358-A | Documentation Incorrect/Unclear | The quantity on the product label did not match the quantity on the packing list or on the COA. The quantity on the product label states 414lb and the quantity on the packing list and COA states 434lb. | USPN052904 | 414.0 |  | 23-May-2019 | 23-May-2019 | Documentation Incorrect/Unclear | Documentation Incorrect/Unclear | Human Error | Human Error. Clariant employee did not verify the correct weight from the container before entering it on the finished good label. | Other/NDF | 414.0 | 0.0 | Closure | CLOSED | 05-Jul-2019 | 21-Jun-2019 | 29.
NC-IL081-12563 | 23-May-2019 | BENJAMIN.KELLY | General | Internal | S7A | 7A-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-M-V-Grease/Lint/Particulate 2 | Post Production NCR. Red/Brown material found on top of plunger | 18D23763X' '18G03963X | 307200.0 | Post Production. Red/Brown material found on top of plunger SCAR # CASE-2019-00058038\nBPCS Hold # 19050044 | 23-May-2019 | 14-May-2019 | SFF-M-V-Grease/Lint/Particulate 2 | Grease/Lint/Particulate Type 2: Particulate outside of syringe | Equipment | Equipment | Other/NDF | 307200.0 | 0.0 | Closure | CLOSED | 31-Jul-2019 | 31-Jul-2019 | 68. | 7 | Syringe
NC-IL081-12622 | 25-May-2019 | ALEJANDRA.CASTANEDA | General | Internal | C56 | Press 56-IL081 |  |  | 8980S | 8 GA RENEWABLES SLIDE LID | CFF-V-Label Verification and Barcode Verification 3 | Missing Print, Missing lot number and GTIN code on the overshipper labels | 19B48063 | 540.0 | Reprogrammed labeler and ensured lot number was visible and correct. - G. Cruz 05/25/2019 | 24-May-2019 | 24-May-2019 | CFF-V-Label Verification and Barcode Verification 3 | OS - Bad/Missing label print | IT/Computer Systems | Associate used a Wi-Fi enabled labeling computer (ELS_WS07) to print over shipper label. Using Wi-Fi instead of hardwire data port connection was causing poor signal strength. As a result, transmitting data was interrupted and the printer was not receiving full data to process the label print. Partially loss of input data caused the missing GTIN print. | Rework | 540.0 | 0.0 | Closure | CLOSED | 22-Aug-2019 | 19-Jul-2019 | 55. | G8 | Container
NC-IL081-12697 | 29-May-2019 | BERTHA.HERNANDEZ | General | Internal | S7P | Packer-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-SPBP-V-Labeling/Print 6 | Failed LVS barcode verification grade, spec is grade A, B or C. Scanner read grade F. | 19D17463X | 87.0 | Line stopped, checkback was performed and affected product was placed on hold. Extra labels were destroyed and new labels printed. | 29-May-2019 | 28-May-2019 | SFF-SPBP-V-Labeling/Print 6 | Fails UDI Grade Test | Misaligned Printer | Misaligned Printer | Rework | 87.0 | 0.0 | Closure | CLOSED | 27-Aug-2019 | 21-Aug-2019 | 84. | 7 | Syringe
NC-IL081-12721 | 29-May-2019 | BERTHA.HERNANDEZ | General | Internal | S4P | Packer-IL081 | Equipment ID | PF 412-IL081 | 8881412012 | 412 CURVED TIP SYRINGE | SFF-SPBP-F-UDI Barcode | Failed LVS barcode verification for overshipper labels, scanner read grade F, spec is A, B or C. | 19E15263X | 240.0 | Completed the lot and stared a new lot with a set of new labels. - B. Villatoro 05/29/2019 | 29-May-2019 | 27-May-2019 | SFF-SPBP-F-UDI Barcode | UDI Barcode | Insufficent white space | Not enough clearance within label template for the scanner to read the barcode | Rework | 240.0 | 0.0 | Closure | CLOSED | 28-Jun-2019 | 21-Jun-2019 | 22. | 412 | Syringe
NC-IL081-12725 | 30-May-2019 | JESSE.SOBEL | General | Internal | S2P | Packer-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-SPBP-F-UDI Barcode | Overshipper labels scan under C rating on LVS barcode inspection. | 19D07563X | 111.0 | Lot was cut short and a new lot with a new set of labels was started. B. Villatoro 05/29/19 | 29-May-2019 | 29-May-2019 | SFF-SPBP-F-UDI Barcode | UDI Barcode | Insufficient white space | Not enough clearance within label template for the scanner to read the barcode. | Rework | 111.0 | 0.0 | Closure | CLOSED | 28-Jun-2019 | 25-Jun-2019 | 26. | 7 | Syringe
NC-IL081-12764 | 30-May-2019 | MARK.VENTURA | General | Internal | S2N | Hot Stamp Printer-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 1 | Scratch on barrel that can be felt with bare finger | 19E15063X | 77600.0 | Line stopped and cleared all printed barrels . Lower dial needed to be deburred. - K. Kaecker 05/30/2019 | 30-May-2019 | 30-May-2019 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 1 |  | Misalignment |  | Sort | 77600.0 | 0.0 | Closure | CLOSED | 29-Jul-2019 | 19-Jun-2019 | 20. | 7 | Syringe
NC-IL081-12834 | 02-Jun-2019 | BENJAMIN.KELLY | General | Internal | S7A | 7A-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-M-V-Grease/Lint/Particulate 2 | Case-2019-00057614-1, particulate outside of syringe. BPCS Hold #19050040 assigned, no product will be moved to the hold area. Post production NCR | 18E12763X' '18F10463X' '18D23663X | 1.0 |  | 02-Jun-2019 | 20-May-2019 | SFF-M-V-Grease/Lint/Particulate 2 | Grease/Lint/Particulate Type 2: Particulate outside of syringe | Equipment | Equipment | Not Applicable | 1.0 | 0.0 | Closure | CLOSED | 31-Jul-2019 | 30-Jul-2019 | 57. | 7 | Syringe
NC-IL081-12835 | 02-Jun-2019 | BENJAMIN.KELLY | General | Internal | S7P | Packer-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-M-V-Contamination/Specks 2 | Case-2019-00053566-1, particulate outside of syringe. BPCS Hold #19050041 assigned for tracking and trending only, no product will be moved to the hold area. Post production NCR | 18E12763X' '18F10463X' '18D23663X | 1.0 | N/A - Post Production | 02-Jun-2019 | 20-May-2019 | SFF-M-V-Contamination/Specks 2 | Contamination/Specks Type 2/3: Inclusions in the non-fluid pathway. Length: Greater than 2 millimeters (0.078 inch), Area: Greater than 3.00 squared mm, or Groups or clusters: Sum of all bubble areas accumulated to a total of 4.00 squared mm shall be considered defective. Use Tappi chart. | Cause Unknown | Cause Unknown | Other/NDF |  | 0.0 | Closure | CLOSED | 31-Jul-2019 | 31-Jul-2019 | 59. | 7 | Syringe
NC-IL081-12836 | 02-Jun-2019 | ANGELICA.AGUILAR | General | Internal | C81 | Press 81-IL081 |  |  | 89671 | 2 GA HORIZONTAL DROP - T RED | CFF-V-Label Verification and Barcode Verification 4 | Bad UDI print on biohazard label for both sides of container. | 19B32963 | 1300.0 | Stopped the press and change the cartridges from both sides. - Beronica Domiquez 06/02/2019 | 02-Jun-2019 | 02-Jun-2019 | CFF-V-Label Verification and Barcode Verification 4 | BH - Bad/Missing pad print | Machine Failure | Print cartridge ran out of ink. | Scrap | 1300.0 | 0.0 | Closure | CLOSED | 06-Sep-2019 | 05-Sep-2019 | 94. | G2 | Container
NC-IL081-12851 | 03-Jun-2019 | GRACE.AUSTIN | General | Internal | C42 | Press 42-IL081 | Equipment ID | R7515A-IL081 | PT00016490 | 12/18 GAL. HINGED CHEMO LID | CFF-V-Broken, Cracking, and Crazing | Cracked hinge, due to lids was not flexed within one minute from injection from the Press | 19E188781' '19D188092 | 4740.0 3395.0 | Standard Work for 12/18 Gallon Hinge Assembly-Orange Dot DC00069083 brought to the press. Step 1A states lids are required to be flexed. K. Wiltse 6/3/19 | 03-Jun-2019 | 03-Jun-2019 | CFF-V-Broken, Cracking, and Crazing | Broken, Cracking, and Crazing all surfaces | Human Error | Human Error | Sort | 4740.0 3395.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 23-Sep-2019 | 112. | G12 | Container
NC-IL081-12979 | 06-Jun-2019 | BENJAMIN.KELLY | General | Internal | S7P | Packer-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-M-V-Grease/Lint/Particulate 2 | CAPA-MEX15-01109, Particulate inside syringe BPCS Hold #19060008 | 18G09363X' '18F10463X' '18F10563X | 1.0 |  | 06-Jun-2019 | 21-May-2019 | SFF-M-V-Grease/Lint/Particulate 2 | Grease/Lint/Particulate Type 2: Particulate outside of syringe | PR-Unknown | Unknown-MEDMFGPR | Not Applicable | 1.0 | 0.0 | Closure | CLOSED | 31-Jul-2019 | 31-Jul-2019 | 55. | 7 | Syringe
NC-IL081-12980 | 06-Jun-2019 | BENJAMIN.KELLY | General | Internal | S7R | RAM-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-M-V-Contamination/Specks 2 | CAPA-MEX15-01117, Black Foreign Material in tray. | 18F10563X' '18G09363X | 1.0 | N/A - Post production | 06-Jun-2019 | 22-May-2019 | SFF-M-V-Grease/Lint/Particulate 2 | Grease/Lint/Particulate Type 2: Particulate outside of syringe | PR-Unknown | Unknown-MEDMFGPR | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 31-Jul-2019 | 31-Jul-2019 | 55. | 7 | Syringe
NC-IL081-13000 | 07-Jun-2019 | ROSEMARY.TEKSTAR | General | Internal | C82 | Press 82-IL081 | Equipment ID | R7103A-IL081 | 31157111 | FISHER 1 GA SHARPS CONT - RED | CFF-V-Flash (Functional Area) | Flash exceeding 0.060" in functional area - rim. | 19C19863 | 2944.0 | Tech adjusted parameters by changing settings, including lowering clamp tonnage, slowing injection speed and hold speed for proper gas escapement. M. Islam 06/07/19 | 06-Jun-2019 | 06-Jun-2019 | CFF-V-Flash (Functional Area) | Sharp Flash; Hinge Pins and Holes; All other areas: Less than or equal to 0.060" and no more than 3 locations per part. | Mold Wear or Damage | Mold Wear or Damage | Rework | 2944.0 | 0.0 | Closure | CLOSED | 06-Jul-2019 | 24-Jun-2019 | 17. | G1 | Container
NC-IL081-13018 | 07-Jun-2019 | GRACE.AUSTIN | General | Internal | C64 | Press 64-IL081 |  |  | WP7449 | Hotair Staked Chemotherapy Lid, 8 Gallon | Documentation Missing | Aux. PCR verification was not performed during in process run. per GS10004646. | 19D188385 | 0.0 |  | 07-Jun-2019 | 07-Jun-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | G8 | Container
NC-IL081-13115 | 11-Jun-2019 | GRACE.AUSTIN | General | Internal | C52 | Press 52-IL081 |  |  | 8870- | 18GA PHARMSAFETY WASTE HINGED | CFF-V-Label Verification and Barcode Verification 3 | Incorrect product code number printed on the overshipper labels. Print 8989 instead of 8870 | 19E11863 | 490.0 | Cleared the printer queue and reprogrammed the overshipper with correct product code. F. Henry 06/10/19 | 11-Jun-2019 | 10-Jun-2019 | CFF-V-Label Verification and Barcode Verification 3 | OS - Bad/Missing label print | Human Error | Human Error | Rework | 490.0 | 0.0 | Closure | CLOSED | 30-Aug-2019 | 14-Aug-2019 | 64. | G18 | Container
NC-IL081-13117 | 11-Jun-2019 | GRACE.AUSTIN | General | Internal | C64 | Press 64-IL081 |  |  | WP8982 | Foam in Place Gasketed Lid, Hinged, 2 Gallon, White | Out of Validated Parameter | PCR out of validated parameter, Fill Time it was 0.79 sec. Specification 0.50-0.60 sec. | 19C187839 | 6191.0 | Stop the press and performed the reading to put it back into Fill Time parameters. - Beronica Domingez 06/11/19 | 11-Jun-2019 | 11-Jun-2019 | Out of Validated Parameter | Out of Validated Parameter - When the cell tech took his PCR reading he recorded the fill time to be .79 seconds when the acceptable range is .50 to .60 seconds. The cell tech failed to recognize that the fill time was out of parameters and continued to run | Human Error | Human Error - The Manufacturing Associate failed to follow GS10004646 Rev 208 Section 6.3 which requires that the PCR readings are recorded and verified once per shift. When the cell tech took his PCR reading he recorded the fill time to be .79 seconds when the acceptable range is .50 to .60 seconds. The cell tech failed to recognize that the fill time was out of parameters and continued to run. The cell tech took the fill time reading from the computer screen on the press and not from the RJG (EDART) Monitor as required. (Attached) | Other/NDF | 6191.0 | 0.0 | Closure | CLOSED | 27-Sep-2019 | 04-Sep-2019 | 85. | G2 | Container
NC-IL081-13212 | 13-Jun-2019 | ROSEMARY.TEKSTAR | General | Internal | C67 | Press 67-IL081 |  |  | 8935 | CONTAINER BIOHAZARD 12GL SHARP | CFF-V-Label Verification and Barcode Verification 3 | Biohazard label missing print. | 19C11063 | 270.0 | Printer roller was bad, roller. Changed out biohazard labels and restarted. Ken Wiltse 06/13/19 | 13-Jun-2019 | 13-Jun-2019 | CFF-V-Label Verification and Barcode Verification 3 |  | Broken Bearing |  | Rework | 270.0 | 0.0 | Closure | CLOSED | 12-Aug-2019 | 22-Jul-2019 | 39. | G12 | Container
NC-IL081-13307 | 18-Jun-2019 | ROSEMARY.TEKSTAR | General | Internal | C15 | Press 15-IL081 | Equipment ID | R7502D-IL081 | 8507SA | CONTAINER SHARPS 5QT RED | Documentation Incomplete/Missing | Process Technician failed to verify validated parameters once during 12 hour shift per GS10004646 (Rev 208) Section 6.3.1. | 19D11463 | 3840.0 | Process tech filled out PCR within parameters. - Jr. Cruz 06/17/19 | 17-Jun-2019 | 17-Jun-2019 | Documentation Incomplete/Missing | Documentation Incomplete/Missing | Human Error | Human Error | Rework | 3840.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 03-Oct-2019 | 107. | SS | Container
NC-IL081-13347 | 18-Jun-2019 | GRACE.AUSTIN | General | Internal | C56 | Press 56-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | CFF-V-Label Verification and Barcode Verification 3 | Missing Pint on Overshipper Labels. Missing Lot# and UDI | 19B47963 | 530.0 | Power down printer and reprogram. | 18-Jun-2019 | 18-Jun-2019 | CFF-V-Label Verification and Barcode Verification 3 |  | Print Queue Buildup |  | Rework | 530.0 | 0.0 | Closure | CLOSED | 17-Sep-2019 | 12-Sep-2019 | 85. | G8 | Container
NC-IL081-13441 | 20-Jun-2019 | GRACE.AUSTIN | General | Internal | C20 | Press 20-IL081 |  |  | 8303SA | 1 QT TRANSPORTABLE CONT | CFF-V-Molding 6 | Hole on the gate area of the Lids R7507A | 19B43663 | 440.0 | Tech turned on Air to Valve gates. | 20-Jun-2019 | 20-Jun-2019 | CFF-V-Molding 6 | Hole Location | Human Error | Human Error. Task is not a documented task with human unknown. Air to the ejector pins was turned off. The most probable cause is it occurred during an extended shutdown; technicians normally shutdown the water to eliminate sweating from the tool. It is most probable that the air valve was shutoff instead of the water. | Scrap | 440.0 | 0.0 | Closure | CLOSED | 04-Aug-2019 | 26-Jun-2019 | 6. | G1 | Container
NC-IL081-13502 | 21-Jun-2019 | ROSEMARY.TEKSTAR | General | Internal | AS1 | Assembly 1-IL081 |  |  | 1523SA | TRAY HOLDER FOR 2.2 QT CONT | CFF-V-Incorrect Items | Non-skid tape applied to the outside short end of the container. SW1523SA (Rev 8) states to apply the tape to the inside short end of container. | 19E13063 | 60.0 | Stopped production, rework was done at press. B. Villatoro | 21-Jun-2019 | 21-Jun-2019 | CFF-V-Incorrect Items | Incorrect IFU/Absorbent pads/Other Kit items/Other Labels in the Overshipper Box | Human Error | The Associate was told by a temporary employee on how to assemble the trays. Associate did not read Work Instruction Sheets and followed what the temp employee said. See Human Error Form for further details. | Rework | 60.0 | 0.0 | Closure | CLOSED | 07-Sep-2019 | 12-Aug-2019 | 51. | AS | Container
NC-IL081-13526 | 23-Jun-2019 | LAKEYSHA.SIMPSON | General | Internal | S2N | Hot Stamp Printer-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag marks on barrel, that can be felt with bare finger. | 19F12963X' '19F13063X | 96000.0 48000.0 | The infeed dial was replaced the sharp edge on the rail was polished to remove the scratches. L. Reed 07/03/19 | 23-Jun-2019 | 23-Jun-2019 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Incorrect Tools |  | UseAsIs | 96000.0 48000.0 | 0.0 | Closure | CLOSED | 21-Oct-2019 | 15-Oct-2019 | 114. | 7 | Syringe
NC-IL081-13568 | 25-Jun-2019 | LAKEYSHA.SIMPSON | General | Internal | C56 | Press 56-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | CFF-D-Molding 1 | Wall thickness out of specification @0.058". Specification is 0.060". | 19F04863 | 5220.0 | The machine was stopped. Tool room shimmed the wear plate. B. Villatoro 06/25/19 | 25-Jun-2019 | 25-Jun-2019 | CFF-D-Molding 1 | Wall Thickness | Materials | Materials | Other/NDF | 5220.0 | 0.0 | Closure | CLOSED | 13-Sep-2019 | 06-Sep-2019 | 73. | G8 | Container
NC-IL081-13601 | 25-Jun-2019 | BENJAMIN.KELLY | General | Internal | S8D | 8D-IL081 | Equipment ID | 58121-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 3 | Flash in non-functional area; string flash on cavity 14E plunger thumbpad area. | 19F10063X | 71000.0 | Cleared line of all sub-assemblies and plungers from 8D mold 58121. - P. Myerson 06/25/2019 | 25-Jun-2019 | 25-Jun-2019 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Flash | Change in cushion caused by cavities not filling resulted in extra plastic/ pressure coming through the mold. | Sort | 71000.0 | 0.0 | Closure | CLOSED | 04-Oct-2019 | 02-Oct-2019 | 98. | Sub | Syringe
NC-IL081-13707 | 27-Jun-2019 | GRACE.AUSTIN | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | CFF-V-Packaging 1 | Missing 4 lids in one FG box. during QA final case inspection. | 19D04963 | 300.0 | Retrained associate per Standard Work - Steve Schirmer 06/27/2019 | 27-Jun-2019 | 27-Jun-2019 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | Human Error | Rework | 300.0 | 0.0 | Closure | CLOSED | 26-Aug-2019 | 17-Jul-2019 | 20. | G2 | Container
NC-IL081-13820 | 02-Jul-2019 | ROSEMARY.TEKSTAR | General | Internal | C67 | Press 67-IL081 |  |  | 31143665 | 10 GA SHARPS CONTAINER | CFF-V-Incorrect Items | Incorrect IFU CL63R98334A being packed into finish good box, the correct IFU is CL63R98396A. | 19D20763 | 180.0 | Removed all incorrect IFUs from press and replaced with correct IFUs. Reworked all affected product at the press. M. Islam 07/01/19 | 01-Jul-2019 | 01-Jul-2019 | CFF-V-Incorrect Items | Incorrect IFU/Absorbent pads/Other Kit items/Other Labels in the Overshipper Box | Human Error | Human Error: Associate failed to verify component per shop order prior to inserting IFU’s into finished good. Backup lead misread the IFU number and incorrect IFU’s were mistakenly delivered to the associate. | Rework | 180.0 | 0.0 | Closure | CLOSED | 06-Sep-2019 | 05-Sep-2019 | 65. | G10 | Container
NC-IL081-13858 | 03-Jul-2019 | JESSE.SOBEL | General | Internal | C74 | Press 74-IL081 |  |  | 8985 | 8 GA HINGED LID CHEMO -YELLOW | CFF-D-Molding 3 | Gasket height does not meet minimum requirement per LS10000108, Rev 11.\nWP8985 Lids was used into Finish Good 8985. | 19B48563' '19C187846 | 1440.0 825.0 | Lowered cooling time and turned on the fans. F. Henry 07/03/19 | 03-Jul-2019 | 02-Jul-2019 | CFF-D-Molding 3 | Gasket Height | Machine |  | Sort' 'Scrap | 1440.0 825.0 | 0.0 | Closure | CLOSED | 28-Sep-2019 | 17-Sep-2019 | 76. | G8 | Container
NC-IL081-13881 | 03-Jul-2019 | GRACE.AUSTIN | General | Internal | C63 | Press 63-IL081 |  |  | WP8999 | FIP Gasketed Lid, Hinged, 12/18 Gallon | Documentation Missing | Missing Aux. PCR from the Start-up | 19C187825 | 630.0 | Performed PCR verification to ensure we were running within Parameter | 03-Jul-2019 | 03-Jul-2019 | Documentation Missing | Documentation Missing | Misplaced Paperwork |  | Other/NDF | 630.0 | 0.0 | Closure | CLOSED | 01-Sep-2019 | 18-Jul-2019 | 15. | G12 | Container
NC-IL081-13921 | 05-Jul-2019 | BENJAMIN.KELLY | General | Internal | CME | CME-IL081 |  |  |  |  | Procedure Not Followed | Project leader did not confirm revision or obtain necessary signatures as required per QS00002152 Rev 203. Issue found during review of DHR 8881412012, Lot# 19F22363X. Further review found that the following Non-Standard Shop Order were also affected: 188928, 188927, 189056, 189214, and 189123. |  |  | N/A - Non-Product NCR | 05-Jul-2019 | 25-Jun-2019 | Procedure Not Followed | Procedure Not Followed | Procedure Not Followed | Procedure not followed due to formatting of form |  |  | 0.0 | Closure | CLOSED | 03-Sep-2019 | 24-Jul-2019 | 19.
NC-IL081-13975 | 09-Jul-2019 | ROSEMARY.TEKSTAR | General | Internal | C81 | Press 81-IL081 |  |  | 89671 | 2 GA HORIZONTAL DROP - T RED | Out of Validated Parameter | Barrel Temp reading documented at 470. Validated parameter control limit is 490. | 19D16763 | 2780.0 | Adjusted barrel temp back within validated parameters. F. Henry 07/08/19 | 08-Jul-2019 | 08-Jul-2019 | Out of Validated Parameter |  | Human Error | Human Error | Other/NDF | 2780.0 | 0.0 | Closure | CLOSED | 30-Aug-2019 | 15-Aug-2019 | 37. | G2 | Container
NC-IL081-14021 | 10-Jul-2019 | GRACE.AUSTIN | General | Internal | C67 | Press 67-IL081 |  |  | 8611RC | 12 GA RCRA HINGED LID | CFF-V-Quantity per Case | Incorrect Quantity per case, missing 5 ea. of lids packed in FG box. | 19G04463 | 70.0 | Stopped Press removed scale MD-682 and Red Tagged. Brought a new scale | 10-Jul-2019 | 10-Jul-2019 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | Human Error | Rework | 70.0 | 0.0 | Closure | CLOSED | 04-Oct-2019 | 23-Sep-2019 | 75. | G12 | Container
NC-IL081-14058 | 11-Jul-2019 | JESSE.SOBEL | General | Internal | S3D | 3D-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-M-V-Flash/Gate 3 | Flash on caps (Non-functional area) exceeds 1/16". | 19F21563X' '19F21463X | 23520.0 4800.0 |  | 10-Jul-2019 | 10-Jul-2019 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Machine |  | Sort | 23520.0 4800.0 | 0.0 | Closure | CLOSED | 30-Sep-2019 | 11-Sep-2019 | 62. | 3 | Syringe
NC-IL081-14077 | 11-Jul-2019 | FIDELIA.CANTOR | Receipt | Supplier | Incoming | Incoming -MYS03 |  |  | PT00006610 | (COV) Printed Product Label, Generic, Biohazard Waste Dispos | Broken / Damaged upon receipt | Five (5) damaged boxes found upon receipt, the inside contents found damaged( 14 rolls of labels). | 1192-52 | 105000.0 | Damaged product placed on hold per NC-IL081-14077 | 11-Jul-2019 | 11-Jul-2019 | PR- Damage | Damage-MEDMFGPR | Handling/Storage | Handling/Storage. Both the Carrier and Manufacturer agreed that the way the boxes were stacked on the pallet led to the damage. More detail in attached investigations. | Scrap | 105000.0 | 0.0 | Closure | CLOSED | 06-Sep-2019 | 27-Aug-2019 | 46.
NC-IL081-14130 | 12-Jul-2019 | ROSEMARY.TEKSTAR | General | Internal | S7A | 7A-IL081 |  |  | 994501820 | In-Process Syringe 12ml Oral Purple Sub, Bulk | Tracking/Trending | Customer complaint: Incorrect or mixed product. Scar#CAPA-NE017-01201 | 19A23063X' '19F05563X' '19F05663X | 1.0 | N/A, Post-production | 12-Jul-2019 | 12-Jul-2019 | SFF-M-V-Discoloration/Marbling/Color Streaks 1 | Molded components do not meet color and clarity including marbling | Component Design | Component Design | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 11-Sep-2019 | 28-Aug-2019 | 46.
NC-IL081-14132 | 12-Jul-2019 | ROSEMARY.TEKSTAR | General | Internal | S7A | 7A-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-SA-V-Particulate 2 | Particulate outside syringe-plunger; Case-2019-00072830-1, CAPA-MEX15-01204 | 17J05363X' '17J05163X' '17H31563X | 1.0 | N/A Post Production | 12-Jul-2019 | 12-Jul-2019 | SFF-SA-V-Particulate 2 | Particulate Type 2: Extraneous matter outside of syringe; Excessive lubricant in or on the syringe | Process/Method | Process/Method | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 30-Aug-2019 | 26-Aug-2019 | 44. | 7 | Syringe
NC-IL081-14134 | 12-Jul-2019 | ROSEMARY.TEKSTAR | General | Internal | S7A | 7A-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-SA-V-Particulate 2 | Particulate outside syringe-plunger thumb area; CASE-2019-00070512-1 | 17L23063X' '18D23663X | 1.0 | N/A Post Production | 12-Jul-2019 | 12-Jul-2019 | SFF-SA-V-Particulate 2 |  | Human Error | Human Error | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 30-Aug-2019 | 26-Aug-2019 | 44. | 7 | Syringe
NC-IL081-14135 | 12-Jul-2019 | ROSEMARY.TEKSTAR | General | Internal | S7C | 7C-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-SA-V-Particulate 2 | Particulate outside syringe-Non-Fluid path | 18F10463X' '18G09363 | 1.0 | N/A Post Production | 12-Jul-2019 | 12-Jul-2019 | SFF-SA-V-Particulate 2 | Particulate Type 2: Extraneous matter outside of syringe; Excessive lubricant in or on the syringe | Machine Failure | Machine Failure | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 30-Aug-2019 | 26-Aug-2019 | 44. | 7 | Syringe
NC-IL081-14180 | 16-Jul-2019 | JESSE.SOBEL | General | Internal | S4A | 412A-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-SPBP-V-Labeling/Print 6 | Overshipper labels failed UDI barcode verification | 19F22363X | 50.0 | Destroyed remaining labels, printed new labels, de-labeled and re-labeled completed cases 1 -50. - J. Williams 07/15/19 | 15-Jul-2019 | 03-Jul-2019 | SFF-SPBP-V-Labeling/Print 6 | Fails UDI Grade Test | Machine | Print head dirty | Rework | 50.0 | 0.0 | Closure | CLOSED | 30-Sep-2019 | 11-Sep-2019 | 57. | 412 | Syringe
NC-IL081-14206 | 16-Jul-2019 | JACLYN.GRISKO | General | Internal | [6 empty] | Calibration Damage Gauge | WSV-005 Calibration Alert for Damaged Repairable |  |  | Issued Cal Alert and updated GageTrak. - Fred McNatt 12/06/2018 | 16-Jul-2019 | 26-Nov-2018 | Calibration Damage Gauge | Calibration Damage Gauge | Damaged Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 398.
NC-IL081-14253 | 17-Jul-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Over Due | MD-222 Calibration alert for overdue |  |  | Not applicable for non product. | 17-Jul-2019 | 12-Feb-2018 | Calibration Over Due | Calibration Over Due | Overdue calibration |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 396.
NC-IL081-14254 | 17-Jul-2019 | ROSEMARY.TEKSTAR | General | Internal | C74 | Press 74-IL081 |  |  | 8850 | 8 GA PHARMASAFETY CONT -HINGED | Out of Spec Visual Inspection (Visual Defect) | WP8850 lids used in FG have dents in the foam causing low gasket height. | 19C15363' '19C187851 | 800.0 2834.0 | Air rack with defective lids was removed and new rack with acceptable lids was inspected before press started. O. Ratinskii 07/17/19 | 17-Jul-2019 | 17-Jul-2019 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Cooling fans |  | Sort | 800.0 2834.0 | 0.0 | Closure | CLOSED | 15-Oct-2019 | 25-Sep-2019 | 70. | G8 | Container
NC-IL081-14333 | 20-Jul-2019 | MARK.VENTURA | General | Internal | S2R | RAM-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-M-V-Grease/Lint/Particulate 1 | Extraneous matter in fluid pathway. | 19G04063X | 71200.0 | Scrapped all printed barrels, repaired screw on in-feed dial mount. Cleaned printer and RAM. - J. Williams 07/19/19 | 20-Jul-2019 | 19-Jul-2019 | SFF-M-V-Grease/Lint/Particulate 1 | Extraneous matter in the fluid pathway | Machine Failure | Machine Failure - Screw damaging luer tip caused debris/particulate as well as damaged component. | Sort | 71200.0 | 0.0 | Closure | CLOSED | 26-Oct-2019 | 03-Oct-2019 | 74. | 7 | Syringe
NC-IL081-14412 | 23-Jul-2019 | BENJAMIN.KELLY | General | Internal | S2R | RAM-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-SA-V-Damaged Components 3 | Damaged product / component. Functional but less than desired. damage at outer luer . | 19G04063X | 192000.0 | Scrapped all printed barrels, repaired screw on in-feed dial barrel mount. - J. Williams 07/19/19 | 23-Jul-2019 | 19-Jul-2019 | SFF-M-V-Broken/Damaged 2 | Broken or damaged product but still Functional | Machine Failure | Machine Failure - Screw damaging luer tip caused debris/particulate as well as damaged component. | Sort | 192000.0 | 0.0 | Closure | CLOSED | 26-Oct-2019 | 03-Oct-2019 | 71. | 7 | Syringe
NC-IL081-14451 | 25-Jul-2019 | BENJAMIN.KELLY | General | Supplier | STT | Tip Tumbling-IL081 |  |  | 4006202 | Plunger Tip, 12mL, Smooth, Black, Latex Free | Visual- Water in package/on tip | 3 totes suspected to have water contamination after tip tumbling | 00001755160010 | 200000.0 | Cleaned the tumblers and verified no water present. | 24-Jul-2019 | 24-Jul-2019 | Visual- Water in package/on tip | Visual- Water in package/on tip | PR-Undetermined | PR-Undetermined-MEDMFGPR. Per West, a definitive root cause of the moisture observed in the bags could not be determined at the time of the investigation. The supplier stated that "it is possible that carton 1 was packaged warm immediately after the rinse/dry cycle it can cause small amounts of condensation to form inside of the unopened bags. However, the amount of water seen in these bags is not typical." | Sort' 'Scrap | 170000.0 30000.0 | 0.0 | Closure | CLOSED | 07-Sep-2019 | 19-Sep-2019 | 56. | Tips | Syringe
NC-IL081-14492 | 25-Jul-2019 | MARK.VENTURA | General | Internal | S2R | RAM-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-M-V-Grease/Lint/Particulate 1 | Extraneous matter in fluid pathway. Found during functional inspection by manufacturing. | 19G10863X' '19G10763X' '19G10663X | 192000.0 191200.0 480.0 | Cleared inside of vacuum tubes. Cleaned inside of RAM. Cleaned inside of printer. - J. Willams 07/25/2019 | 25-Jul-2019 | 25-Jul-2019 | SFF-M-V-Grease/Lint/Particulate 1 | Grease/Lint/Particulate Type 1: Particulate in fluid pathway | PR-Housekeeping | Housekeeping-MEDMFGPR | Sort' 'Scrap | 191200.0 192000.0 480.0 | 0.0 | Closure | CLOSED | 01-Nov-2019 | 25-Oct-2019 | 91. | 7 | Syringe
NC-IL081-14506 | 26-Jul-2019 | GRACE.AUSTIN | General | Internal | C64 | Press 64-IL081 |  |  | WP8985 | Foam in Place Gasketed Lid, Single Pass, 8 Gallon | Out of Validated Parameter | Used incorrect PCR R7481A, the correct PCR should be R7481A/Robot (Cool Time is out of parameter) | 19F189143 | 939.0 | Re-did start-up with correct PCR. Steve Schirmer | 26-Jul-2019 | 26-Jul-2019 | Out of Validated Parameter | Out of Validated Parameter | Running Outside Of Specified Range And Or Validated Parameters | Running Outside Of Specified Range And Or Validated Parameters | Other/NDF | 939.0 | 0.0 | Closure | CLOSED | 07-Sep-2019 | 06-Sep-2019 | 42. | G8 | Container
NC-IL081-14553 | 29-Jul-2019 | BERTHA.HERNANDEZ | General | Internal |  |  |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Validation | Validation parts mold # 8. | 19G26-7-8 | 540000.0 | Placed product on hold. | 29-Jul-2019 | 29-Jul-2019 |  | No investigation required. NC was opened for control of material for internal validation per CDP-0018674. | PR-Undetermined | PR-Undetermined-MEDMFGPR. No investigation required. NC was opened for control of material for internal validation per CDP-0018674. | Other/NDF | 30000.0 510000.0 | 0.0 | Closure | CLOSED | 30-Sep-2019 | 13-Sep-2019 | 46. | Tips | Syringe
NC-IL081-14557 | 29-Jul-2019 | BERTHA.HERNANDEZ | General | Internal | INC | Incoming-INC-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Validation | Validation PQ parts, mold # 9. | 19G26-9-10 | 540000.0 | Placed product on hold. | 29-Jul-2019 | 29-Jul-2019 |  | No investigation required. NC was opened for control of material for internal validation per CDP-0018674 | PR-Undetermined | PR-Undetermined-MEDMFGPR. No investigation required. NC was opened for control of material for internal validation per CDP-0018674 | Other/NDF | 30000.0 510000.0 | 0.0 | Closure | CLOSED | 30-Sep-2019 | 13-Sep-2019 | 46. | Tips | Syringe
NC-IL081-14588 | 30-Jul-2019 | JACLYN.GRISKO | General | Internal | Laboratory | Laboratory |  |  |  |  | Other | Melt Flow Tester (GW-002) is having inconsistent communication between the tester and computer. |  |  |  | 30-Jul-2019 | 24-Jul-2019 | PR-Other |  | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 07-Apr-2020 | 03-Apr-2020 | 248.
NC-IL081-14638 | 31-Jul-2019 | ROSEMARY.TEKSTAR | General | Internal | C85 | Press 85-IL081 | Equipment ID | R7089C-IL081 | R7089W-1 | 2 GAL. HINGED LID, WHT, NO CAP | Documentation Incorrect/Unclear | Incorrect PCR used at press. PCR should be for R7089W-1, but R7089C had been used. | 19B187353 | 1.0 | No immediate action required, post-production NC. All parameters were verified to have been running within documented specifications. | 31-Jul-2019 | 08-Apr-2019 | Documentation Incorrect/Unclear | Documentation Incorrect/Unclear | Human Error | Human Error | No Issue / No Impact | 1.0 | 0.0 | Closure | CLOSED | 29-Sep-2019 | 16-Sep-2019 | 46. | G2 | Container
NC-IL081-14643 | 01-Aug-2019 | MARIA.ROSALES | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-P-V-Alpha Numeric 4 | Smeared and illegible print on cap. | 19G22663X | 48000.0 | Stopped line and cleared all in-process printed caps. Began Check-back procedure. | 31-Jul-2019 | 31-Jul-2019 | SFF-PPHP-V-Lot Number 5 | Cap printing is illegible | Procedure/Method | Procedure/Method | Sort | 48000.0 | 0.0 | Closure | CLOSED | 14-Sep-2019 | 11-Sep-2019 | 41. | 3 | Syringe
NC-IL081-14686 | 02-Aug-2019 | JESSE.SOBEL | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 3 | Non-cavity specific string flash on gate exceeds 1/16". | 4502570 | 42000.0 | Line stopped and cleared, product placed on hold, check back performed. | 01-Aug-2019 | 01-Aug-2019 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Incorrect Temperature | Incorrect Temperature: Elevated tip temperature resulted in excess material pulling off the gate pin in the cavity, creating string flash. Tip temperature was elevated due to fluctuations in the gate tip temperature from the DME unit. DME temperature is adjustable along with process adjustments. These parameters are not critical to quality but reference only in the validated PCR. | Scrap | 42000.0 | 0.0 | Closure | CLOSED | 30-Sep-2019 | 18-Sep-2019 | 47. | Sub | Syringe
NC-IL081-14713 | 02-Aug-2019 | BENJAMIN.KELLY | General | Internal | S1A | 1A-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-M-V-Broken/Damaged 1 | Broken / Damaged luer Skirt rendering product non functional | 19G16063X | 1600.0 | Cleared out all barrels and sub-assemblies. Scrapped 2 cases. Process tech adjusted cool air pressure. - P. Myerson 08/02/19 | 02-Aug-2019 | 02-Aug-2019 | SFF-M-V-Broken/Damaged 1 | Broken or damaged product rendering it Non Functional | Lure thread not seated on the pinion shaft | During tool repair, the lure thread was not seated on the pinion shaft, causing it to not fully align with the rack component. | Scrap | 1600.0 | 0.0 | Closure | CLOSED | 04-Oct-2019 | 23-Sep-2019 | 52. | 1 | Syringe
NC-IL081-14715 | 02-Aug-2019 | GRACE.AUSTIN | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | 6" Ruler MD-0712 Lost and can\'t be locate. |  |  | Issued Cal Alert and updated Gage Trak - Fred McNatt 3/4/19 | 02-Aug-2019 | 07-Mar-2016 | Calibration Lost Gauge | Calibration Lost Gauge | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 380.
NC-IL081-14720 | 02-Aug-2019 | GRACE.AUSTIN | General | Internal | QA | QA-IL081 |  |  |  |  | Fail to Locate | 6" Ruler, MD-0645 lost, and can\'t be locate. |  |  | Issued Cal Alert and updated Gage Trak - Fred McNatt 3/19/19 | 02-Aug-2019 | 23-Mar-2016 | Calibration Lost Gauge | Calibration Lost Gauge | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 380.
NC-IL081-14721 | 02-Aug-2019 | GRACE.AUSTIN | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | Gage ID: 50; Calibration Alert for out of Tolerance |  |  | Issued Cal Alert and updated Gage Trak - Fred McNatt 4/3/19 | 02-Aug-2019 | 04-Oct-2018 | Calibration Out of Tolerance | Calibration Out of Tolerance | Normal use and age of equipment | Balances inherently go out of calibration due to normal use and wear and tear, which is why calibration is required. |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 380.
NC-IL081-14722 | 02-Aug-2019 | GRACE.AUSTIN | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | Gage ID: 54; Scale - Calibration Alert for out of Tolerance |  |  | Issued Cal Alert and updated Gage Trak - Fred McNatt 4/3/19 | 02-Aug-2019 | 04-Oct-2018 | Calibration Out of Tolerance | Calibration Out of Tolerance | Normal use and age of equipment | Balances inherently go out of calibration due to normal use and wear and tear, which is why calibration is required. |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 380.
NC-IL081-14723 | 02-Aug-2019 | GRACE.AUSTIN | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Damage Gauge | MD-007 Calibration Alert for Damaged Unrepairable |  |  | Issued Cal Alert and updated Gage Trak - Fred McNatt 4/15/19 | 02-Aug-2019 | 18-Apr-2018 | Calibration Damage Gauge | Calibration Damage Gauge | PR-Undetermined | PR-Undetermined-MEDMFGPR. The exact origin of the damage is unknown. The issue is most likely to have occurred through normal wear and tear and age of the gage. |  |  | 0.0 | Closure | CLOSED | 13-Sep-2019 | 03-Sep-2019 | 32.
NC-IL081-14724 | 02-Aug-2019 | GRACE.AUSTIN | General | Internal | [6 empty] | Calibration Out of Tolerance | LVS-010 Calibration alert for out of Tolerance |  |  | Issued Cal Alert and updated Gage Trak - Fred McNatt 6/3/19 | 02-Aug-2019 | 01-May-2019 | Calibration Out of Tolerance | Calibration Out of Tolerance | Out of Tolerance |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 380.
NC-IL081-14726 | 02-Aug-2019 | GRACE.AUSTIN | General | Internal | QA | QA-IL081 |  |  |  |  | PR-Other QS Issue | SR-304 Calibration Alert to set/modify specifications. |  |  | SR-304 Cal alert and updated Gage Trak - Fred McNatt 6/28/19 | 02-Aug-2019 | 03-Oct-2018 |  | An inconsistency with unit of measure for calibration by TCS of Torque Sensors Mark-10/MR50-100 was identified upon review of calibration certificates for Cardinal Health gages SR-304, SR-1160, and SR-1161. \n\nSR-304 was calibrated using in/oz on 10/3/18; SR-1160 was calibrated on 2/26/19 using lbf-in; and SR-1161 was calibrated on 7/22/19 using lbf-in. | 2210 MISSING/EXTRA/WRONG COMPONENT | MISSING/EXTRA/WRONG COMPONENT- The display units are plug and play for use of different torque cells, Cardinal Health has sent in display units with different range torque cells. |  |  | 0.0 | Closure | CLOSED | 13-Jan-2020 | 09-Jan-2020 | 160.
NC-IL081-14742 | 03-Aug-2019 | MARK.VENTURA | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SPBP-V-Labeling/Print 1 | Expiration date is incomplete or missing on over shipper labels | 19G22863X | 116640.0 | Lot cut short, QA placed product on hold, New lot started. - P. Myerson 08/03/2019 | 03-Aug-2019 | 03-Aug-2019 | SFF-SPBP-V-Labeling/Print 1 | Missing Lot Number /Label | Printer Malfunction | Printer Malfunction | Rework | 116640.0 | 0.0 | Closure | CLOSED | 13-Sep-2019 | 04-Sep-2019 | 32. | 3 | Syringe
NC-IL081-14828 | 06-Aug-2019 | BENJAMIN.KELLY | General | Internal | S8K | Kahle-IL081 |  |  |  |  | Procedure Not Followed | Weekly Volumetric testing not completed for line 8. | 19G23263X | 1.0 | 08/06/18 Preformed Volumetric testing with samples produced 08-02-19 with passing results. B. Kelly | 06-Aug-2019 | 02-Aug-2019 | Other Failure | Other Failure | Cause Unknown | Cause Unknown |  |  | 0.0 | Closure | CLOSED | 31-Mar-2020 | 27-Mar-2020 | 233.
NC-IL081-14859 | 07-Aug-2019 | BENJAMIN.KELLY | General | Internal | S3C | 3C-IL081 | Equipment ID | 58040-IL081 | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-M-V-Shorts 1 | Cavity Specific short that affects the sterility of the assembled part. Cap cavity 9-1 from mold 58040. | 19G23163X' '19G23063X' '19G22963X | 43200.0 172800.0 | Raised shot size to allow more material into cavity. | 07-Aug-2019 | 07-Aug-2019 | SFF-M-V-Shorts 1 | Shorts in fluid pathway or in area where Sterility is required | Tooling/ fixture wear | Defect resulted from worn nozzle probe. Tech’s failure to block the cavity 9-1 allowed material flow back into the cavity inconsistently and produce parts with short shots. | Sort | 43200.0 172800.0 | 0.0 | Closure | CLOSED | 05-Nov-2019 | 20-Dec-2019 | 135. | 3 | Syringe
NC-IL081-14938 | 09-Aug-2019 | MARIA.ROSALES | General | Internal | S1R | RAM-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-SA-F-Spraymation | Group average exceeds max limit allowed. | 19G24663X' '19G16163X' '19G16263X | 435200.0 499200.0 | Spraymation head was replaced and re-testing results are passing. Mike Buck 08/09/19 | 08-Aug-2019 | 08-Aug-2019 | SFF-SA-F-Spraymation | Spraymation | Machine Failure | Machine Failure | Other/NDF | 499200.0 435200.0 | 0.0 | Closure | CLOSED | 08-Oct-2019 | 23-Aug-2019 | 14. | 1 | Syringe
NC-IL081-14940 | 09-Aug-2019 | LAKEYSHA.SIMPSON | General | Internal | Warehouse-CRI-01 | Warehouse-CRI-01 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-V-Broken, Cracking, and Crazing | Material handler knocked over multiple skids in the warehouse. Product suspected broken, cracked or damage. | 19E06363 | 540.0 | Contain all affected product on hold. | 09-Aug-2019 | 09-Aug-2019 | CFF-V-Broken, Cracking, and Crazing | Broken, Cracking, and Crazing all surfaces | Human Error | Human Error | Rework | 540.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 25-Oct-2019 | 77. | SS | Container
NC-IL081-14956 | 09-Aug-2019 | ROSEMARY.TEKSTAR | General | Internal | C85 | Press 85-IL081 |  |  | PT00013692 | Gasketed Lid, Single Pass, 8 Gallon | CFF-F-Molding 2 | Lids found with cracked hinges upon incoming inspection from RAMPF. Lids manufactured at CAH. | 19F189050 | 17548.0 | All affected product placed on hold. | 09-Aug-2019 | 09-Aug-2019 | CFF-F-Molding 2 | Hinged Lid | Human Error | Human Error | Sort | 17548.0 | 0.0 | Closure | CLOSED | 08-Oct-2019 | 30-Sep-2019 | 52. | G8 | Container
NC-IL081-14963 | 09-Aug-2019 | MARK.VENTURA | General | Internal | CME | CME-IL081 |  |  | 8881112059 | 12ML SYRINGE REGULAR NS | PR-Environmental | Spider found on the floor of the CME processing area of line 6. | 19G25063X | 1.0 | Removed spider from floor and turned into Quality in bag. - K. Kaecker 08/13/2019 | 09-Aug-2019 | 09-Aug-2019 | PR-Environmental | Environmental-MEDMFGPR | PR-Environmental Control | Environmental Control-MEDMFGPR | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 01-Nov-2019 | 18-Oct-2019 | 69. | 7 | Syringe
NC-IL081-14964 | 10-Aug-2019 | MARIA.ROSALES | General | Internal | S2R | RAM-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | PCR out of parameter. | Spraymation on time-side B out of parameter, spec is 13-19 it was changed to 20. | 19H10163X | 9600.0 | The parameter for spraymation on time was lowered to be within parameters. Mike Buck 08/09/19 | 09-Aug-2019 | 09-Aug-2019 | Out of Validated Parameter | Out of Validated Parameter | Human Error | Human Error | Scrap | 9600.0 | 0.0 | Closure | CLOSED | 14-Sep-2019 | 11-Sep-2019 | 32. | 3 | Syringe
NC-IL081-14991 | 12-Aug-2019 | BENJAMIN.KELLY | General | Internal | S6A | 6A-IL081 |  |  | 3150450 | Barrel, 12 mL Luer Lock with Finger Flange | Failure to follow Procedure | Failure to follow Non-Standard Shop order procedure. No BPCS hold number assigned before the S.O. was ran. | TSTA080619 | 20000.0 | Documented the issue with this NCR. No product effected, all parts made scrapped per non-standard shop order instructions. BPCS hold number 19080014 created for nonstandard shop order. | 12-Aug-2019 | 07-Aug-2019 | Failure to follow Procedure |  | PR-Training | Training-MEDMFGPR | Scrap | 20000.0 | 0.0 | Closure | CLOSED | 29-Nov-2019 | 21-Oct-2019 | 69.
NC-IL081-15001 | 12-Aug-2019 | FIDELIA.CANTOR | Receipt | Supplier |  |  |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Water in package/on tip | Water/condensation was found in the bag of one case of rubber tips for 4001204, lot 19H09-1-9, during incoming inspection on 8/12/19, failing incoming inspection for presence of water in package or on/in rubber tips, Inspection Level II, AQL 0.065, (A=0/R=1) per GS20003456, Rev. 2. | 19H09-1-9 | 3000000.0 |  | 12-Aug-2019 | 12-Aug-2019 | Visual- Water in package/on tip | Visual- Water in package/on tip | PR-Undetermined | PR-Undetermined-MEDMFGPR. To be documented in SCAR CAPA-IL081-01268 | Rework | 3000000.0 | 0.0 | Closure | CLOSED | 30-Sep-2019 | 12-Sep-2019 | 31. | Tips | Syringe
NC-IL081-15044 | 13-Aug-2019 | ROSEMARY.TEKSTAR | General | Internal | C63 | Press 63-IL081 |  |  | WP7417BL | 8871 LID | Procedure/ Missing | Formula sheet not yet updated for each WP part that can be ran in FIPC-2. We are currently routing them in Agile.. This is a process NC only, no product affected. | 1234567890 | 1.0 | No immediate action required, Process NC, no product affected. | 13-Aug-2019 | 13-Aug-2019 | Procedure/ Missing | Procedure/ Missing | Incomplete update of documentation |  |  |  | 0.0 | Closure | CLOSED | 12-Sep-2019 | 04-Sep-2019 | 21. | G18 | Container
NC-IL081-15087 | 14-Aug-2019 | GRACE.AUSTIN | General | Internal | C54 | Press 54-IL081 |  |  | 8930SA | "30GA RED W/ HINGED, GASKET LID" | CFF-V-Label Verification and Barcode Verification 8 | Biohazard Label out of Specification, Incorrect Label Placement some too low/high. Per FM20008973 Rev. 2 | 19E05263 | 939.0 | Engineer made adjustment to the pick and place of the container to move the label within spec. - M. Ozimek 08/15/2019 | 14-Aug-2019 | 14-Aug-2019 | CFF-V-Label Verification and Barcode Verification 8 | BH - Incorrect label placement | Inherent Variation of Cooling Rack |  | Sort | 939.0 | 0.0 | Closure | CLOSED | 13-Oct-2019 | 06-Sep-2019 | 23. | MS | Container
NC-IL081-15208 | 17-Aug-2019 | JESSE.SOBEL | General | Internal | S2R | RAM-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-SA-V-Particulate 3 | Particulate found in fluid pathway of assembled barrel | 19H03763X | 51200.0 | Rubber tip bowls & RAM deep cleaned. - K. Kaecker 08/17/2019 | 17-Aug-2019 | 16-Aug-2019 | SFF-SA-V-Particulate 3 | Particulate Type 3: ‘Enteral Oral’ Syringes ONLY: Extraneous matter in fluid pathway. (Does not apply to Enfit Enteral Syringe with external luer threads | Incorrect guide rail setup | The Apex printer had a broken mandrel guide causing the mandrel to bent/misaligned and parts were rubbing inside creating the particulate as a result. Most probable cause for the bent mandrel is restarting the printer without clearing barrel jam on infeed dial. | Sort | 51200.0 | 0.0 | Closure | CLOSED | 16-Oct-2019 | 11-Oct-2019 | 55. | 7 | Syringe
NC-IL081-15238 | 19-Aug-2019 | GRACE.AUSTIN | General | Internal | Rework | Rework |  |  | 8930SA | "30GA RED W/ HINGED, GASKET LID" | CFF-V-Labeling - Product & Overshipper 1 | Missing Lot number and wrinkled Biohazard Label | 19E05263 | 939.0 | N/A- Post production. | 19-Aug-2019 | 19-Aug-2019 | CFF-V-Labeling - Product & Overshipper 1 | Product and Lot Number is incomplete or missing | Label Guide |  | Rework | 939.0 | 0.0 | Closure | CLOSED | 18-Oct-2019 | 06-Sep-2019 | 18. | MS | Container
NC-IL081-15244 | 19-Aug-2019 | ROSEMARY.TEKSTAR | General | Internal | C77 | Press 77-IL081 |  |  | 31322723 | FISHER 3.5 GA TP - T. RED | CFF-V-Label Verification and Barcode Verification 2 | Product label missing ink from product code and lot number. | 19F13763 | 570.0 | Ribbon on biohazard labeler was straightened and not to have wrinkles. Oleg Ratinskiy 08/19/19 | 19-Aug-2019 | 19-Aug-2019 | CFF-V-Labeling - Product & Overshipper 1 | Product and Lot Number is incomplete or missing | Printer Tape Issue | Printer Tape Issue | Rework | 570.0 | 0.0 | Closure | CLOSED | 18-Sep-2019 | 09-Sep-2019 | 20. | G3 | Container
NC-IL081-15263 | 20-Aug-2019 | LAKEYSHA.SIMPSON | General | Internal | C31 | Press 31-IL081 |  |  | 8790- | 2 GA NON-INFECTIOUS WASTE CONT | CFF-V-Blisters, Bubbles and Voids | Bubbles out of specification on container per GS10000913 . | 19F14763 | 2140.0 | Tech made PCR adjustment to eliminate bubbles. See attached PCR . Doc. No. 1180804-2496. Ken Wiltse 8/20/19 | 20-Aug-2019 | 20-Aug-2019 | CFF-V-Blisters, Bubbles and Voids | Single bubble greater than or equal to 0.080”; Up to 3 bubbles between 0.060 and 0.080”; Up to 4 bubbles less than or equal to 0.060”; Up to 8 bubbles less than 0.030”; Living Hinge = Reject All | Blocked Vents | Blocked Vents: R7028M mold was producing container with bubbles due to lack of venting and improper air escapement. The vents found to be clogged from plastic residue. Tech cleaned the mold to wipe off the residue, also process adjustment was made to ensure proper air escapements. Tech increased hold psi & hold time which helped the part pack out longer, so the air escapement pushed toward end of the fill and vented out. (see attached PCR) | Sort | 2140.0 | 0.0 | Closure | CLOSED | 19-Oct-2019 | 18-Sep-2019 | 29.
NC-IL081-15265 | 20-Aug-2019 | MARIA.ROSALES | General | Internal | S8A | 8A-IL081 | Equipment ID | 58118-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-D-Dimensional | Barrel luer exceeds no-go of go/no-go gauge (cavity # 12D from mold 58118) | 19H09963X' '19H09863X' '19H09763X' '19H09663X' '19H09563X' '19H09463X' '19H09363X | 205000.0 335000.0 | Stopped the line. Scrapped all molded parts from mold 58118 inside hoppers, conveyors, and kahle. Cut lot short. Blocked cavity 12D. - L. Reed 08/20/2019 | 20-Aug-2019 | 14-Aug-2019 | SFF-M-D-Dimensional | Dimensional Out of Spec | Non-OEM component manufactured near out of tolerance |  | UseAsIs | 205000.0 335000.0 | 0.0 | Closure | CLOSED | 19-Oct-2019 | 06-Sep-2019 | 17. | Sub | Syringe
NC-IL081-15286 | 20-Aug-2019 | FIDELIA.CANTOR | Receipt | Supplier |  |  |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Water in package/on tip | Water/condensation was found in the bag of rubber tips for 4001204, lot 19H16-1-8, during incoming inspection on 8/20/19, failing inspection for presence water in package or on/in rubber tips, inspection level II AQL 0.065 (A-0/R-1) per GS20003456, Rev. 2. Also found a rubber band inside the bag of rubber tips. | 19H16-1-8 | 2880000.0 | Placed all affected product on hold. | 20-Aug-2019 | 20-Aug-2019 | Visual- Water in package/on tip | Visual- Water in package/on tip | PR-Undetermined | PR-Undetermined-MEDMFGPR. To be documented in SCAR CAPA-IL081-01268 | Rework | 2880000.0 | 0.0 | Closure | CLOSED | 30-Sep-2019 | 12-Sep-2019 | 23. | Tips | Syringe
NC-IL081-15382 | 23-Aug-2019 | MARIA.ROSALES | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-PPHP-V-Lot Number 5 | Smeared and illegible print on cap, it was found during rework for short on cap. | 19G22963X | 11520.0 | Defect discovered during rework processes. No line startup process correction needed. M. Buck 09/11/19 | 23-Aug-2019 | 23-Aug-2019 | SFF-PPHP-V-Lot Number 5 | Printing is in the incorrect position and illegible | Printer Malfunction | Printer Malfunction | Sort | 11520.0 | 0.0 | Closure | CLOSED | 21-Oct-2019 | 31-Oct-2019 | 69. | 3 | Syringe
NC-IL081-15421 | 24-Aug-2019 | MARIA.ROSALES | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | Out of Spec Visual Inspection (Visual Defect) | Incorrect lot number found during rework in which MFG was sorting for shorts on cap. | 19G22963X | 13440.0 | Defect was discovered during rework processes, no line startup process correction needed. M. buck 09/11/19 | 24-Aug-2019 | 24-Aug-2019 | SFF-PPHP-V-Lot Number 3 | Lot number is incomplete/incorrect | Human Error | Human Error | Sort | 13440.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 31-Oct-2019 | 68. | 3 | Syringe
NC-IL081-15425 | 24-Aug-2019 | BERTHA.HERNANDEZ | General | Internal | S4R | RAM-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-SA-V-Damaged Components 3 | Damaged product functional but less than desired. Picture attached. | 19F18763X | 2000.0 | Cleared line of all sub-assemblies, 100% 4 cases for damage. Damage caused by jams, clear rail when jams. - P. Myerson 08/24/2019 | 24-Aug-2019 | 24-Aug-2019 | SFF-SA-V-Damaged Components 3 | Damaged Components Type 3: Damaged product/component (functional but less than desired); Scratches/drag marks on the barrel | Complex machine set-up | Machine set-up requires fine tuning | Sort | 2000.0 | 0.0 | Closure | CLOSED | 30-Sep-2019 | 04-Sep-2019 | 10. | 412 | Syringe
NC-IL081-15440 | 26-Aug-2019 | BENJAMIN.KELLY | General | Internal | S2P | Packer-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-M-V-Grease/Lint/Particulate 1 | Particulate in fluid pathway | 19H15763X | 13600.0 | Cleared line of all parts, scrapped 17cs affected. Cleaned RAM, Printer, and Tubes. - P. Myerson 08/26/2019 | 26-Aug-2019 | 26-Aug-2019 | SFF-M-V-Grease/Lint/Particulate1 | Particulates in fluid pathway | Particulate | Build up of particulate due to friction of the barrels going through the delivery tube system. In-sufficient intervals of cleaning of the delivery tubes. | Scrap | 13600.0 | 0.0 | Closure | CLOSED | 26-Oct-2019 | 08-Oct-2019 | 43. | 7 | Syringe
NC-IL081-15454 | 26-Aug-2019 | ROSEMARY.TEKSTAR | General | Internal | C52 | Press 52-IL081 |  |  | 8989 | 18 GA HINGED CHEMO - YELLOW | CFF-V-Labeling - Product & Overshipper 1 | Biohazard label missing print from lot number and GTIN number. | 19G04163 | 210.0 | Adjusted labeler print head to print correctly. De-labeled and relabeled containers. G. Cruz 08/26/19 | 26-Aug-2019 | 26-Aug-2019 | CFF-V-Labeling - Product & Overshipper 1 | Product and Lot Number is incomplete or missing | Human Error | Human Error- An adjustment was made to the Print head and the head thumb screw was not tightened back. | Rework | 195.0 15.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 30-Oct-2019 | 64. | G18 | Container
NC-IL081-15523 | 28-Aug-2019 | GRACE.AUSTIN | General | Internal | S6D | 6D-IL081 |  |  | X9900 | Plunger, 12 mL with Extended Flange | Documentation Not Approved | Not Following the Process of Non-Standard Shop Order. QS00002152 | TSTC082319 | 10000.0 |  | 28-Aug-2019 | 28-Aug-2019 | Failure to follow Procedure |  | PR-Training | Training-MEDMFGPR | Scrap | 10000.0 | 0.0 | Closure | CLOSED | 27-Oct-2019 | 31-Oct-2019 | 64.
NC-IL081-15591 | 30-Aug-2019 | BERTHA.HERNANDEZ | General | Internal | S4A | 412A-IL081 |  |  | 3156001 | Molded Component, 412 Curved Tip Barrel | SFF-M-V-Bubbles 1 | Bubble in fluid pathway, in the barrel tip of cavity 19B. | 19H16163X | 3600.0 5400.0 | Cleared line, scrapped affected product, cavity blocked. - P. Myerson 08/30/19 | 30-Aug-2019 | 30-Aug-2019 | SFF-M-V-Bubbles 1 | Gas Bubbles in critical area such as barrel luer, fluid pathway, ability to create leakage or affect sterility or in barrel walls that during use may be interpreted as bubbles in the content (i.e. in medicine, etc…) | clogged bubbler | The small cooling tubes (bubblers) in the affected cores were blocked / restricted with debris limiting their ability to affectively remove heat from the core tip. | Scrap | 3600.0 5400.0 | 0.0 | Closure | CLOSED | 11-Oct-2019 | 02-Oct-2019 | 33. | 412 | Syringe
NC-IL081-15603 | 01-Sep-2019 | LAKEYSHA.SIMPSON | General | Internal | 131 | Cell 131-IL081 |  |  | 8506SA | CONTAINER BIOHAZARD 5 QT | CFF-F-Welding | Failed weld strength inspection per LS10000087. | 19G01563' '19G01763 | 9980.0 360.0 | Technician cleaned welder horn, increased Amplitude from 70 to 100 and also increased Trigger Force from 15 to 20. Fran Henry 08/31/19 | 31-Aug-2019 | 31-Aug-2019 | CFF-F-Welding | Weld Strength | Incomplete, improper weld | Incomplete, improper weld | Scrap | 9980.0 360.0 | 0.0 | Closure | CLOSED | 29-Oct-2019 | 10-Oct-2019 | 39. | SS | Container
NC-IL081-15632 | 03-Sep-2019 | BENJAMIN.KELLY | General | Internal | S4A | 412A-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-M-V-Flash/Gate 3 | Excessive flash found on barrel, causing parts to get stuck in printer. | 19H18463X | 500.0 | Cleared line of all barrels and sub-assemblies. Scrapped affected product. - P. Myerson 09/03/19 | 03-Sep-2019 | 03-Sep-2019 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Excessive material | Excessive material due to shot size, clogged the stripper ring. | Scrap | 500.0 | 0.0 | Closure | CLOSED | 03-Oct-2019 | 23-Sep-2019 | 19. | 412 | Syringe
NC-IL081-15638 | 04-Sep-2019 | JESSE.SOBEL | General | Internal | S1P | Packer-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-SA-V-Damaged Components 3 | Plunger and finger flange damaged, but still functional. | 19H13863X | 19200.0 | Line stopped and affected product placed on hold. | 04-Sep-2019 | 03-Sep-2019 | SFF-SA-V-Damaged Components 3 | Damaged Components Type 3: Damaged product/component (functional but less than desired); Scratches/drag marks on the barrel | Equipment | The packer transfer tube was scratched up from normal wear and tear causing sporadic functionality of packer (root cause). | Sort | 19200.0 | 0.0 | Closure | CLOSED | 03-Nov-2019 | 03-Oct-2019 | 29. | 1 | Syringe
NC-IL081-15658 | 04-Sep-2019 | ROSEMARY.TEKSTAR | General | Internal | C31 | Press 31-IL081 |  |  | 47 | PSS SELECT 2 GAL , RED | CFF-V-Label Verification and Barcode Verification 5 | Biohazard label missing lot number and UDI print. | 19G07863 | 760.0 | Turn off the printer and reprogram the labeler. V. Marrero 09/05/19 | 04-Sep-2019 | 04-Sep-2019 | CFF-V-Label Verification and Barcode Verification 5 | OS - Bad/Missing pad print | Printer Malfunction | Printer Malfunction - Printer (Significant – The printer lost the programmed information during the run. The label on the DHR indicated that the first labels was programmed and printed correctly.) The printer connections were confirmed to be correct and we were unable to recreate the defect. | Rework | 760.0 | 0.0 | Closure | CLOSED | 08-Nov-2019 | 04-Nov-2019 | 61. | G2 | Container
NC-IL081-15665 | 05-Sep-2019 | ROSEMARY.TEKSTAR | General | Internal | C54 | Press 54-IL081 |  |  | WP7417P | 8618RC LID | CFF-F-Molding 8 | Failed gasket height requirements. | 19H189753' '19H20263' '19J02363' '19J02463' '19H19963' '19H01063' '19H02763 | 2075.0 1230.0 835.0 210.0 3995.0 4420.0 942.0 | Line stayed down until corrective actions could be completed. | 04-Sep-2019 | 04-Sep-2019 | CFF-F-Molding 8 | Visual anomaly was identified and created an excessively tall gasket height which caused a leaking condition | PR-Process Control | Process Control-MEDMFGPR | No Issue / No Impact' 'Pending Product Impact | 4420.0 3995.0 1230.0 835.0 210.0 2075.0 942.0 | 0.0 | Closure | CLOSED | 02-Nov-2019 | 21-Oct-2019 | 46. | G18 | Container
NC-IL081-15667 | 05-Sep-2019 | JESSE.SOBEL | General | Internal | S0A | S0A-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | PR-Process Settings | Missed PCR reading MP-1123/0A on 09/03/2019 for product packed into finished goods. | 19H13863X | 176000.0 | Crew 3 verified all settings were within specification after restart. Product made between press restart and next verified inspection was placed on QA Hold. - J. Williams 09/04/2019 | 04-Sep-2019 | 02-Sep-2019 | PR-Process Settings | Process Settings-MEDMFGPR | Human Error | Human Error - Lack of communication/verification | Other/NDF | 176000.0 | 0.0 | Closure | CLOSED | 02-Nov-2019 | 03-Oct-2019 | 28. | 1 | Syringe
NC-IL081-15677 | 05-Sep-2019 | FIDELIA.CANTOR | Receipt | Supplier | PR-Incoming Non-Conforming Material | PR-Incoming Non-Conforming Material |  |  | 5241005 | Thermal Transfer Label, Roll, Unprinted, White, 5 Inch Lengt | PR-- Wrong Material | During incoming inspection was found out that supplier sent incorrect product # (Supplier part #70P* BTM4050AP correct supplier # 7AC*FAN4050 (CAH part number 5241005). Per RMPS (RM10014546) labels should be in Rolls the supplier sent folded labels. |  | 300000.0 |  | 05-Sep-2019 | 05-Sep-2019 | PR-- Wrong Material | Wrong Material-MEDMFGPR | PR-Training | Training-MEDMFGPR- Due to a lack of training by Buckeye to fully understand the part numbers and the RMPS numbers, we (Buckeye) overlooked the difference between the part number 5241005 used for the rolled labels used in Crystal Lake and the part number of 5241005-29DL used in Deland for the fan fold labels. We had recently sent an order of these to Deland, so when this order was received for Crystal Lake, we thought it was for the same fanfold labels, which was an error. | Return to Vendor | 300000.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 10-Oct-2019 | 35.
NC-IL081-15689 | 05-Sep-2019 | GRACE.AUSTIN | General | Internal | C67 | Press 67-IL081 |  |  | 31156550 | 7 GA SHARPS CONTAINER | CFF-V-Label Verification and Barcode Verification 2 | Missing Print. UDI number on Biohazard Label is missing multiple digit. | 19F21863 | 150.0 | We changed out the biohazard printer and red tagged with work order CL-3122093 to fix the bad printer. S. Schirmer 09/05/19. | 05-Sep-2019 | 05-Sep-2019 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Stuck Labels & Loose Screw |  | Rework | 150.0 | 0.0 | Closure | CLOSED | 04-Nov-2019 | 14-Oct-2019 | 39. | G7 | Container
NC-IL081-15717 | 06-Sep-2019 | GRACE.AUSTIN | General | Internal | C82 | Press 82-IL081 |  |  | 31144010 | 5 QT TP SHARPS CONT -TRANS RED | CFF-V-Quantity per Case | Missing IFU in one Finished Good found during QA Final case Inspection. Per DC00032963 | 19G09563 | 30.0 | Verified correct instruction sheets available at press. - J. Pixius 09/06/19 | 06-Sep-2019 | 06-Sep-2019 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | Human Error | Rework | 30.0 | 0.0 | Closure | CLOSED | 05-Nov-2019 | 06-Sep-2019 | 0. | G5 | Container
NC-IL081-15726 | 06-Sep-2019 | PATTI.SWISHER | General | Internal | S8P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Mixed Product 1 | Incorrect or mixed product | 19F19263X | 8.0 |  | 06-Sep-2019 | 29-Aug-2019 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Human Error | The most probable root cause includes a process inspector leaving samples at line 8 after being pulled from line 7 for evaluation, followed by an associate using the samples as back-fill after sampling the given lot. Furthermore, there is no procedure in place to control sample handling to help prevent mix up. | Scrap | 8.0 | 0.0 | Closure | CLOSED | 04-Nov-2019 | 03-Oct-2019 | 26. | Sub | Syringe
NC-IL081-15763 | 09-Sep-2019 | BERTHA.HERNANDEZ | General | Internal |  |  |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-SA-V-Mixed Product 1 | Mixed product, SR barrel found packed in F/G during rework activities on this lot. | 19G10863X | 800.0 | Start check back, contained affected product and place on hold. | 09-Sep-2019 | 07-Sep-2019 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | lodged part | SR barrel was lodged/caught somewhere within the line from a previous code and came dislodged and ended up in 8881112083 code. | Sort | 800.0 | 0.0 | Closure | CLOSED | 31-Jan-2020 | 13-Jan-2020 | 126. | 7 | Syringe
NC-IL081-15784 | 09-Sep-2019 | BENJAMIN.KELLY | General | Internal | S4A | 412A-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-M-V-Bubbles 1 | Bubble on luer tip, Cav # 10B. (Bubbles also found on luer tip of 01B & 20B during checkback process) | 19J00363X | 21500.0 | Mold Pulled for repair. Performed Lot change J.Williams 09/09/2019 | 09-Sep-2019 | 09-Sep-2019 | SFF-M-V-Bubbles1 | Bubbles in barrel tip | Clogged bubbler | Mold had a hot core due to bubbler being clogged with debris and not able to be properly cooled. | Sort | 21500.0 | 0.0 | Closure | CLOSED | 08-Nov-2019 | 03-Oct-2019 | 23. | 412 | Syringe
NC-IL081-15786 | 10-Sep-2019 | ROSEMARY.TEKSTAR | General | Internal | C14 | Press 14-IL081 |  |  | 8507SA | CONTAINER SHARPS 5QT RED | Documentation Incomplete/Missing | Failure to document machine parameters once per shift per GS10004646 Rev. 208. | 19G06263 | 240.0 |  | 09-Sep-2019 | 09-Sep-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | SS | Container
NC-IL081-15796 | 10-Sep-2019 | JESSE.SOBEL | General | Internal | S4P | Packer-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Scratch on outside of barrel; can be felt with finger. | 19J00463X' '19J00363X | 8500.0 19500.0 | Polished ram rail to remove burrs. scrapped all affected product 17 cases . J.Williams 09/10/2019. | 10-Sep-2019 | 09-Sep-2019 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Tooling/ fixture wear | Defect resulted from wear/damage on the inner rail of the ram, creating metal burrs which scratched the barrels as they traveled down the rail. | Sort' 'Scrap | 8500.0 19500.0 | 0.0 | Closure | CLOSED | 30-Oct-2019 | 11-Oct-2019 | 31. | 412 | Syringe
NC-IL081-15842 | 11-Sep-2019 | MARK.VENTURA | General | Internal |  |  |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | PR-Environmental | Insect found on inspection table located in the Rework area. | 19G10863X | 1.0 | Inspected the area for more insects and cleaned the table. Scrapped the case that staff was reworking when the insect was discovered. | 11-Sep-2019 | 11-Sep-2019 | PR-Environmental | Environmental-MEDMFGPR | PR-Environmental Control | Environmental Control-MEDMFGPR | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 17-Oct-2019 | 09-Oct-2019 | 28. | Sub | Syringe
NC-IL081-15914 | 13-Sep-2019 | ROSEMARY.TEKSTAR | General | Internal | C25 | Press 25-IL081 |  |  | 8507SA | CONTAINER SHARPS 5QT RED | Documentation Incomplete/Missing | Failed to verify/document process parameters once per shift per GS10004646 Rev. 208. | 19G06263 | 5280.0 | Process control record was documented. - G. Cruz Jr. | 12-Sep-2019 | 12-Sep-2019 | Documentation Incomplete/Missing | Documentation Incomplete/Missing | Human Error | Human Error - PS10009274 is a combined PCR of all three (3) equipment. Cell tech did not pay attention and missed the back page to fill out the process value for AM194. | Other/NDF | 5280.0 | 0.0 | Closure | CLOSED | 10-Nov-2019 | 11-Oct-2019 | 28. | SS | Container
NC-IL081-16066 | 18-Sep-2019 | BERTHA.HERNANDEZ | General | Internal | S8A | 8A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Outside Validated Parameters | Fill time out of parameters, found at 0.95 spec is 0.96-1.11 | 19J04563X' '19J04463X | 4000.0 67000.0 | Stopped and cleared the khale, all WIP scrapped at the line. F/G contained. Fill time was placed back to validated parameters. | 18-Sep-2019 | 18-Sep-2019 | Outside Validated Parameters | Outside Validated Parameters | Complex machine set-up | Machine set-up requires fine tuning | Scrap' 'Sort | 4000.0 67000.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 16-Oct-2019 | 28. | Sub | Syringe
NC-IL081-16073 | 18-Sep-2019 | BERTHA.HERNANDEZ | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Bubbles 1 | Bubble in fluid pathway of cavity 14. | 19J04463X | 81000.0 | Stopped and cleared the line. placed product on hold (BPCS Hold #19090039). | 18-Sep-2019 | 18-Sep-2019 | SFF-M-V-Bubbles 1 | Gas Bubbles in critical area such as barrel luer, fluid pathway, ability to create leakage or affect sterility or in barrel walls that during use may be interpreted as bubbles in the content (i.e. in medicine, etc…) | Obstruction | Blockage / obstruction in the gate vents causing a build up of gas. | Sort | 81000.0 | 0.0 | Closure | CLOSED | 18-Oct-2019 | 16-Oct-2019 | 28. | Sub | Syringe
NC-IL081-16092 | 19-Sep-2019 | JESSE.SOBEL | General | Internal | S2B | 2B-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-M-V-Flash/Gate 3 | Flash in non-functional area exceeds 1/16" on sleeve | 19J08163X | 79200.0 | Line stopped to clear and removed mold for repair. K. Kaecker 09/26/2019 | 19-Sep-2019 | 18-Sep-2019 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Tool PM |  | Scrap | 79200.0 | 0.0 | Closure | CLOSED | 17-Nov-2019 | 16-Oct-2019 | 27. | 3 | Syringe
NC-IL081-16093 | 19-Sep-2019 | JESSE.SOBEL | General | Internal | S2B | 2B-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | "Scratch on Sleeve can be felt with finger, Scratches on the inside of the part & Cav.'s #24 and 38 scratches on inside of the part, caused by tool damage." | 19J08163X | 136800.0 | Line stopped and cleared and all effected parts placed on hold. | 19-Sep-2019 | 18-Sep-2019 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Tool Wear |  | Scrap | 136800.0 | 0.0 | Closure | CLOSED | 17-Nov-2019 | 16-Oct-2019 | 27. | 3 | Syringe
NC-IL081-16141 | 20-Sep-2019 | JACLYN.GRISKO | General | Internal | S2R | RAM-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | Failed Qualification Test | On 9/17/2019 Dose Audit results for product family 12MLRP were reported above sterility threshold of 2/100 positive cultures. Product code 8881512878 lot 19G12963X was recorded has having 3/100 positive cultures with a conclusion that sterilization dose may not be adequate. | 19G12963X | 360.0 | 1. Calculated augmentation of Sterilization does to be increased from 15.8kGy to 17.8kGy. (See Augmentation Document 12MLRP-signed)\n2. Complied list of 12MLRP product family lots produced since date of last passing dose audit, 05/03/2019. (See 12MLRP Lots 05032019 to 09182019).\n3. Reviewed sterilization records of lots listed in See 12MLRP Lots 05032019 to 09182019 for minimum dose recorded and assessed it against the augmented dose for need of containment. (See Attached GS10000179 Dose Audit Excursion Investigation). | 20-Sep-2019 | 17-Sep-2019 | Sterility Assurance Failure | Sterility Assurance Failure | Operator in learning curve | Lack of expertise by the operator | No Issue / No Impact | 360.0 | 0.0 | Closure | CLOSED | 19-Dec-2019 | 13-Dec-2019 | 84. | 3 | Syringe
NC-IL081-16168 | 21-Sep-2019 | JESSE.SOBEL | General | Internal | C61 | Press 61-IL081 |  |  | R7417P | 12/18 GALLON SLIDE TOP LID | Documentation Not Approved | Document deviation. Red-lined for black dots on lid not present at time of molding. Will be performed during gasketing. | 19J190230 | 1.0 |  | 20-Sep-2019 | 20-Sep-2019 | [6 empty] | 0.0 | Initiation Verification | VOID
NC-IL081-16184 | 21-Sep-2019 | BERTHA.HERNANDEZ | General | Internal | S2B | 2B-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Scratch in non-fluid pathway cause by tool damage inside the sleeve (cavity#24) that can be felt with a bare finger. | 19J08163X | 27840.0 | Stopped and cleared the line, cavity 24 was blocked and affect product was placed on hold. | 21-Sep-2019 | 20-Sep-2019 | [6 empty] | 0.0 | Investigation | VOID |  |  |  | 3 | Syringe
NC-IL081-16192 | 22-Sep-2019 | JESSE.SOBEL | General | Internal | S2B | 2B-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Scratch on sleeve can be felt with finger. Cavity specific #38 | 19J08163X | 57600.0 | Stopped and cleared the line, cavity #38 was blocked and affected product was placed on hold. | 21-Sep-2019 | 21-Sep-2019 | [6 empty] | 0.0 | Investigation | VOID |  |  |  | 3 | Syringe
NC-IL081-16199 | 22-Sep-2019 | BERTHA.HERNANDEZ | General | Internal | S1A | 1A-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-SA-V-Mixed Product 1 | Mixed product 3150450 barrel mixed with 3154020 inside barrel hopper. Operator dumped a brute of 3154020 barrel into the hopper. F/G 8881112599, 19J03663X was in production at this time and confirmed to not be impacted. | 19J03663X' '19H09863X | 7200.0 4800.0 | Stopped line & cleared out the line. Scrapped all Whip product. - D. Breceda 09/22/2019 | 22-Sep-2019 | 22-Sep-2019 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Process/Method | Process/Method - Brute Kanban system - (Significant – The brute color management system did not clearly define how wip parts from one line to be used on another line for finished goods assembly should be managed. A change order needs to be made to better define the use of the colored brute system for this purpose. The pouch provided on the brutes was initially used when the parts were produced. The brute pouches were available for use on the brutes in question. | Scrap | 7200.0 4800.0 | 0.0 | Closure | CLOSED | 15-Nov-2019 | 04-Nov-2019 | 43. | 1 | Syringe
NC-IL081-16251 | 23-Sep-2019 | GRACE.AUSTIN | General | Internal | C54 | Press 54-IL081 |  |  | 8991 | 18 GA RENEWABLES HINGE LID | CFF-V-Contamination (Embedded) | Contamination (Embedded) on containers | 19G15463 | 3990.0 |  | 23-Sep-2019 | 23-Sep-2019 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | G18 | Container
NC-IL081-16266 | 24-Sep-2019 | JESSE.SOBEL | General | Internal | S2D | 2D-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Scratch that can be felt on outside of barrel. Cavity# 17B | 19H26563X | 13920.0 | Tech had to clean the gate for access drool F. Bahena 09/24/2019 | 24-Sep-2019 | 23-Sep-2019 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Machine Start Stop | Process tech found piece of drool plastic sticking to the gate. The drool plastic was dragging outside of the barrel part during mold opening resulted in scratch. The process control found to be running within validated limit. Most probable cause for the drool plastic is the degraded material as a result of machine stopped in alarm and sit for a while without being addressed | Scrap | 13920.0 | 0.0 | Closure | CLOSED | 22-Nov-2019 | 17-Oct-2019 | 23. | 3 | Syringe
NC-IL081-16267 | 24-Sep-2019 | JESSE.SOBEL | General | Internal | S2N | Hot Stamp Printer-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-SA-V-Damaged Components 3 | Broken or damaged product but still functional. Damage located on finger flange. | 19H26563X | 34560.0 | Tech adjusted the outside guard rail because it was to too tight causing it too rub and causing the string flash. F. Bahena 09/24/2019 | 24-Sep-2019 | 23-Sep-2019 | SFF-SA-V-Damaged Components 3 | Damaged Components Type 3: Damaged product/component (functional but less than desired); Scratches/drag marks on the barrel | Incorrect guide rail setup | Incorrect guide rail setup - Maintenance found the outside guide rail was too tight against the dial causing it to rub and make string flash like plastic peeling. Most probable cause for the guide rail being off position is infeed dial jam. The rail might get pulled when jam occurred outside of the infeed guide rail. | Scrap | 34560.0 | 0.0 | Closure | CLOSED | 22-Nov-2019 | 17-Oct-2019 | 23. | 3 | Syringe
NC-IL081-16364 | 26-Sep-2019 | LAKEYSHA.SIMPSON | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | CFF-V- Blisters, Bubbles and Voids | Bubbles/Blisters out of specification by the rim area of containers. | 19F21763 | 1860.0 | Technician cleaned the dirt behind the ejector plate. Fran Henry 09/25/19 | 25-Sep-2019 | 25-Sep-2019 | CFF-V-Blisters, Bubbles and Voids | Single bubble greater than or equal to 0.080”; Up to 3 bubbles between 0.060 and 0.080”; Up to 4 bubbles less than or equal to 0.060”; Up to 8 bubbles less than 0.030”; Living Hinge = Reject All | Mold Wear or Damage | Mold Wear or Damage | Sort | 1860.0 | 0.0 | Closure | CLOSED | 25-Oct-2019 | 24-Oct-2019 | 28. | G2 | Container
NC-IL081-16421 | 27-Sep-2019 | BERTHA.HERNANDEZ | General | Internal | S8P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SPBP-V-Assembly 6 | Mixed product: incorrect poly bags component found at the packer 5000109, should be 5170588 per BOM. | 19J10663X | 240.0 | Stopped the Khale, removed the incorrect poly bags and placed product on hold. | 27-Sep-2019 | 26-Sep-2019 | SFF-SPBP-V-Assembly 6 | The overshipper, bulk pack, or bundle is not properly sealed in accordance with defined packaging requirements (for example: missing tape). | Human Error | Human Error | Other/NDF | 240.0 | 0.0 | Closure | CLOSED | 26-Nov-2019 | 14-Oct-2019 | 17. | Sub | Syringe
NC-IL081-16428 | 27-Sep-2019 | BERTHA.HERNANDEZ | General | Internal | S8C | 8C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 3 | Flash in non functional area exceeds 1/16" in plungers MP-1120 Mold 58120. | 19J10663X | 7000.0 | Stopped and cleared the line, performed checkback and placed product on hold. | 27-Sep-2019 | 27-Sep-2019 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Manufacturing | Root Cause: Defect caused by the mold closing on a part that failed to eject resulting in a gap along the parting line and flashed parts | Scrap | 7000.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 15-Oct-2019 | 18. | Sub | Syringe
NC-IL081-16429 | 27-Sep-2019 | BERTHA.HERNANDEZ | General | Internal | S8K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Damaged Components 1 | Damaged product non-functional, found during the checkback of NC-IL081-16428 | 19J10663X | 12000.0 | "When Kahle jammed, operator didn't remove all damaged parts from the kahle. - D. Breceda 09/27/2019\nKahle was verified to be cleared of jammed parts. - J. Grisko 10/09/2019 (Interviewed lead for further clarification of restart activities.)" | 27-Sep-2019 | 27-Sep-2019 | SFF-SA-V-Damaged Components 1 | Damaged Components Type 1: Damaged product/components (non-functional) | Machine Failure | Machine Failure | Scrap | 12000.0 | 0.0 | Closure | CLOSED | 31-Oct-2019 | 16-Oct-2019 | 19. | Sub | Syringe
NC-IL081-16483 | 30-Sep-2019 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Water in package/on tip | Water/condensation was found in the bag of rubber tips for 4001204, LOT # 19J27-7-13, during incoming inspection for presence of water in packaging or on/in rubber tips, inspection level II, AQL 0.065 (A=0/R-1) per GS20003456 rev. 2. | 19J27-7-13 | 2320000.0 | Lot put on hold | 30-Sep-2019 | 30-Sep-2019 | Visual- Water in package/on tip | Visual- Water in package/on tip | pr-undetermined | PR-Undetermined-MEDMFGPR | Return to Vendor' 'Scrap | 2319500.0 500.0 | 0.0 | Closure | CLOSED | 30-Oct-2019 | 10-Oct-2019 | 9. | Tips | Syringe
NC-IL081-16512 | 01-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MM-054, thermocouple calibrator out of tolerance as found. Gage cannot be repaired to be back in tolerance. |  |  | Gage was decommissioned. | 01-Oct-2019 | 02-Aug-2018 | Calibration Out of Tolerance | Calibration Out of Tolerance | Out of Tolerance |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 321.
NC-IL081-16513 | 01-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | LVS-003, out of tolerance |  |  | replaced lens, passed subsequent calibration. | 01-Oct-2019 | 18-Jul-2019 | Calibration Out of Tolerance | Calibration Out of Tolerance | Out of Tolerance |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 320.
NC-IL081-16552 | 02-Oct-2019 | GRACE.AUSTIN | General | Internal | C52 | Press 52-IL081 |  |  | 8991 | 18 GA RENEWABLES HINGE LID | CFF-V-Label Verification and Barcode Verification 2 | Biohazard Label not legible, GTIN and Lot numbers are missing print. due to the line across . | 19J15163 | 770.0 | Cleaned the printer roller and removed the paper label that was stock to the roller. Berenice Villatoro 10/02/19 | 02-Oct-2019 | 01-Oct-2019 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Human Error | Human Error- Associate did not remove remaining paper labels from printer roller when the labels were changed. | Rework | 770.0 | 0.0 | Closure | CLOSED | 14-Jan-2020 | 07-Jan-2020 | 97. | G18 | Container
NC-IL081-16599 | 03-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | SR-717, Line 8 in-line check weigher will not calibrate. Previous calibration occurred on 03/26/2019. |  |  | Cream City Scale is investigating the cause of the issue. | 03-Oct-2019 | 26-Sep-2019 | Calibration Out of Tolerance | Calibration Out of Tolerance | Out of Tolerance |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 318.
NC-IL081-16601 | 03-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Damage Gauge | MD-0770, 6" ruler was discovered slightly bent at normally scheduled calibration on 08/26/2019. Ruler passed visual inspection (calibration) and was decommissioned and discarded to eliminate future risk. Previous calibration was on 09/28/2016. |  |  |  | 03-Oct-2019 | 26-Aug-2019 | Calibration Damage Gauge | Calibration Damage Gauge | Damaged Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 318.
NC-IL081-16604 | 03-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Damage Gauge | SR-891, Torque driver tip separated from body exposing the load cell on 08/26/2019. Previous calibration date was 09/06/2018. |  |  | Sent unit out for evaluation on 08/26/2019. | 03-Oct-2019 | 26-Aug-2019 | Calibration Damage Gauge | Calibration Damage Gauge | PR-CALIBRATION | CALIBRATION-MEDMFGPR |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 318.
NC-IL081-16606 | 03-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Past Due | LVS-009 was not able to be calibrated as discovered during normally scheduled calibration activity on 09/13/2019. This was due to the admin level account (which is required for calibration) was locked out. Previous calibration date was 08/15/2019. |  |  | Red-tagged and labelled the unit with a "Do not use until repaired, tested, and calibrated" yellow sticker. Entered IT ticked INC3936616 and placed unit in IT repair queue. | 03-Oct-2019 | 13-Sep-2019 | Calibration Past Due | Calibration Past Due | Past Due |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 318.
NC-IL081-16608 | 03-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0669, 6" ruler is lost. It could not be located during normally scheduled calibration activities on 09/25/2019. Previous calibration was on 09/28/16. |  |  |  | 03-Oct-2019 | 25-Sep-2019 | Calibration Lost Gauge | Calibration Lost Gauge | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 318.
NC-IL081-16609 | 03-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0695, 6" ruler is lost. It could not be located during normally scheduled calibration activities on 09/25/2019. Previous calibration was on 09/28/16. |  |  |  | 03-Oct-2019 | 25-Sep-2019 | Calibration Lost Gauge | Calibration Lost Gauge | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 318.
NC-IL081-16610 | 03-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0756, 6" ruler is lost. It could not be located during normally scheduled calibration activities on 09/25/2019. Previous calibration was on 09/28/16. |  |  |  | 03-Oct-2019 | 25-Sep-2019 | Calibration Lost Gauge | Calibration Lost Gauge | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 318.
NC-IL081-16667 | 06-Oct-2019 | ANGELICA.AGUILAR | General | Internal | C33 | Press 33-IL081 |  |  | WP8935 | 12 GA HEAT STK LID DOOR &RAILS | CFF-V-Packaging 1 | Incorrect quantity of lids packed into the box #2. Should be 37 pieces instead 34 pieces per standard work. | 19F188879 | 2072.0 | Stopped the machines, di an End of Run and performed a Product Change. B. Villatoro | 06-Oct-2019 | 06-Oct-2019 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | Human Error - Associate failed to follow standard work instructions and was working on "auto pilot". See attached human error investigation form. | Rework | 2072.0 | 0.0 | Closure | CLOSED | 04-Dec-2019 | 17-Oct-2019 | 11. | G18 | Container
NC-IL081-16672 | 06-Oct-2019 | BERTHA.HERNANDEZ | General | Internal | S2D | 2D-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-M-V-Bubbles 1 | Bubble in barrel wall that may be interpreted as bubble in content. | 19J02763X | 4800.0 | Stopped and cleared barrel press and printer, performed check back on F/G and WIP. | 06-Oct-2019 | 06-Oct-2019 | SFF-M-V-Bubbles 1 | Gas Bubbles in critical area such as barrel luer, fluid pathway, ability to create leakage or affect sterility or in barrel walls that during use may be interpreted as bubbles in the content (i.e. in medicine, etc…) | Tooling/ fixture damaged | Damaged tooling/ fixture causing process variation Injection Molding (Significant – Injection Speed too high and vents clogged in the mold ). The observed injection speed at the time of the defect was confirmed to be within validated parameters. When the injection speed was lowered by the tech is also was confirmed to be within validated parameters. The lowering of the injection speed offsets the effect of the clogged vents. | Scrap | 4800.0 | 0.0 | Closure | CLOSED | 29-Nov-2019 | 12-Nov-2019 | 37. | 3 | Syringe
NC-IL081-16702 | 07-Oct-2019 | GRACE.AUSTIN | General | Internal | C82 | Press 82-IL081 |  |  | 8982 | 2 GA HINGED CHEMO - YELLOW | CFF-V-Contamination | Resin gels embedded in the bottom of the container. | 19H21363 | 900.0 | End of Run 8982. Ken Wiltse\nThe lot was cut short due to the NCR. The station was switched to produce product code 047 under a different process. Steven Hansen | 07-Oct-2019 | 07-Oct-2019 | CFF-V-Contamination | Carbon or White Specks greater than 0.080" (Surfaces A, B, C, & D) | Process/Method | The molding process is controlled via a specified process in accordance with GS10004646. The process controlled does not include the RPM recovery or barrel heats. Not controlling these parameters on the PCR allowed for "un-melted pellets" as the material in the barrel did not soak as long as necessary. | Other/NDF | 900.0 | 0.0 | Closure | CLOSED | 04-Feb-2020 | 12-Feb-2020 | 128. | G2 | Container
NC-IL081-16706 | 07-Oct-2019 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Water in package/on tip | Water/condensation was found in 8 bags of rubber tips for 4001204, lot 19K04-1-12, box # 161,193,211,212 233, 305, 315, and 317 during incoming inspection on 10/7/19 and 10/8/19, failing incoming inspection for presence of water in package or on/in rubber tips, Inspection Level II, AQL 0.065, (A=0/R=1) per GS20003456, Rev. 2. | 19K04-1-12 | 4320000.0 |  | 07-Oct-2019 | 07-Oct-2019 | Visual- Water in package/on tip | Visual- Water in package/on tip | PR-Undetermined | PR-Undetermined-MEDMFGPR. To be documented in SCAR CAPA-IL081-01268 | Return to Vendor | 4320000.0 | 0.0 | Closure | CLOSED | 06-Nov-2019 | 14-Oct-2019 | 6. | Tips | Syringe
NC-IL081-16716 | 08-Oct-2019 | LAKEYSHA.SIMPSON | General | Internal | S8F | Form, Fill Sealer-IL081 |  |  | 412SE | ENFIT 12 ML SYRINGE | SSF-PPSP-V-SEAL INTEGRITY 1 | Seal integrity, void allows air into sealed package. Cavity specific 2 & 7 | 19G27363X | 35040.0 | Stopped the line, performed check back and placed all affected product on Hold. | 07-Oct-2019 | 07-Oct-2019 | SFF-PPSP-V-Seal Integrity 1 | Void allows air into sealed package | Dirty Vacuum Head |  | Sort | 35040.0 | 0.0 | Closure | CLOSED | 30-Nov-2019 | 18-Oct-2019 | 10. | 8 | Syringe
NC-IL081-16719 | 08-Oct-2019 | LAKEYSHA.SIMPSON | General | Internal | S3N | Hot Stamp Printer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-M-V-Grease/Lint/Particulate 1 | Particulate; Extraneous matter in fluid pathway. | 19J02763X | 72960.0 | Ricardo has aligned the in-feed dial with the mandrel. - F. Bahena 10/07/19 | 07-Oct-2019 | 07-Oct-2019 | SFF-M-V-Grease/Lint/Particulate 1 | Grease/Lint/Particulate Type 1: Particulate in fluid pathway | Mandrel hitting barrel | The in feed dial was not centered. | Scrap | 72960.0 | 0.0 | Closure | CLOSED | 31-Dec-2019 | 27-Dec-2019 | 80. | 3 | Syringe
NC-IL081-16802 | 10-Oct-2019 | LAKEYSHA.SIMPSON | General | Internal | C20B | Press 20B-IL081 |  |  | 8303SA | 1 QT TRANSPORTABLE CONT | Outside Validated Parameters | Cooling time out of parameters. Specification for cooling time should be 5.00-8.00 seconds recorded at 9.00 seconds. | 19G18163 | 1800.0 | Machine cooling time was within parameters 7.00 But document incorrectly on PCR. PCR was corrected by Tech. and press restarted with S?U inspection. Oleg Ratinsky 10/09/19 | 10-Oct-2019 | 10-Oct-2019 | Documentation Incorrect/Unclear | Documentation Incorrect/Unclear | Human Error | Human Error | Other/NDF | 1800.0 | 0.0 | Closure | CLOSED | 09-Nov-2019 | 06-Nov-2019 | 27. | G1 | Container
NC-IL081-16817 | 10-Oct-2019 | BERTHA.HERNANDEZ | General | Internal | S4A | 412A-IL081 |  |  | R7642B | Barrel, 12 mL | PR- Documentation Incomplete | Product consumed prior to release; unable to perform GS20003063 requirements. | 19K00263X' '19K00963X | 144000.0 | Place product on hold. | 10-Oct-2019 | 06-Oct-2019 | PR- Documentation Incomplete | Documentation Incomplete-MEDMFGPR | Lot not released | The last lot of 412SE on the production schedule was skipped, this put the lots being produced out of order. Line 8 had caught up to the lot of barrels that was currently being produced off of 412 molding press. | Other/NDF | 144000.0 | 0.0 | Closure | CLOSED | 31-Jan-2020 | 13-Jan-2020 | 95. | 8 | Syringe
NC-IL081-16818 | 10-Oct-2019 | BERTHA.HERNANDEZ | General | Internal | S8K | Kahle-IL081 |  |  | 8881112015 | 12ML SYRINGE PURPLE EN-FIT TIP | Documentation Not Approved | Used not yet approved rubber tips per BOM; Rayco used instead of West. | 19K00963X' '19K00863X | 144000.0 67000.0 | Placed product on hold. | 10-Oct-2019 | 09-Oct-2019 | Documentation Not Approved | Documentation Not Approved | Operator error | Operator failed to confirm rubber tip component number on the bill of materials when supplying the line. The Rayco part #4101801 is an alternate/ replacement for the West part #4001801 but the CDP had not been approved and the Bill of Material had not been updated. | Hold | 144000.0 67000.0 | 0.0 | Closure | CLOSED | 28-Feb-2020 | 18-Feb-2020 | 131. | 8 | Syringe
NC-IL081-16830 | 10-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out for Repair | LVS-003 would not scan. Issue reported by QA technician, verified by calibration technician. IT ticket INC3968525 entered for repair of unit. The LVS-003 barcode verifier was delivered to the IT repair queue. A calibration alert was issued on 10/10/2019. |  |  |  | 10-Oct-2019 | 19-Sep-2019 | Calibration Out for Repair | Calibration Out for Repair | Out for Repair |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 311.
NC-IL081-16889 | 11-Oct-2019 | JACLYN.GRISKO | General | Internal | S8P | Packer-IL081 |  |  | CM285160000 | Unit Box 12ml Soft Pack printed - Covidien | Supplier Misidentification | During the start up verification of product 1181200777, Lot# 19G16963X it was noticed that the inner carton CM285160000 contained an incorrect identification with a missing zero. | 494113 | 16800.0 | Assessment of product impact completed with no impact to product found. | 11-Oct-2019 | 06-Jun-2019 | Supplier Misidentification | Supplier Misidentification | Process/Method | Process/Method- A mistake was made by OEC Graphics when making the printing plate. Instead of putting four 0’s in the ID number on the printing plate, three 0’s were put on. This was then missed by our Print Plate evaluator in Material Control, during the double check process when he is comparing the print plate to the print plate layout spec. | Other/NDF | 16800.0 | 0.0 | Closure | CLOSED | 29-Nov-2019 | 05-Nov-2019 | 25.
NC-IL081-16900 | 12-Oct-2019 | MARK.VENTURA | General | Internal | S1A | 1A-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | PR-Environmental | During weekly cleaning, a spider was found on floor of line 1 north CME. The spider was discovered under the conveyor for barrels next to press 1 A. Adding to the NC- spider found on line 8 south CME inspection table. | 19K00563x | 1.0 | Spider removed from area line 1. spider removed from area line 8 | 12-Oct-2019 | 12-Oct-2019 | Contamination - Insect/Rodent | Contamination - Insect/Rodent | PR-Not Applicable | Not Applicable-MEDMFGPR | No Issue / No Impact | 1.0 | 0.0 | Closure | CLOSED | 06-Dec-2019 | 15-Nov-2019 | 34. | 7 | Syringe
NC-IL081-16928 | 14-Oct-2019 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Water in package/on tip | Water/condensation was found in the bag of rubber tips for 4001204, lot 19K11-1-12 box # 106 during incoming inspection on 10/14/19, failing incoming inspection for presence of water in package or on/in rubber tips, Inspection Level II, AQL 0.065, (A=0/R=1) per GS20003456, Rev. 2. | 19K11-1-12 | 4320000.0 |  | 14-Oct-2019 | 14-Oct-2019 | Visual- Water in package/on tip | Visual- Water in package/on tip | PR-Undetermined | PR-Undetermined-MEDMFGPR | Return to Vendor | 4320000.0 | 0.0 | Closure | CLOSED | 14-Nov-2019 | 28-Oct-2019 | 14. | Tips | Syringe
NC-IL081-16985 | 15-Oct-2019 | BERTHA.HERNANDEZ | General | Internal | S7C | 7C-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-M-V-Shorts 1 | Short in fluid pathway of barrel cavity #41C. | 19K04063X | 4800.0 | Stopped and cleared the line, performed check back, no defects found in F/G, all WIP scrapped at the line. | 15-Oct-2019 | 15-Oct-2019 | SFF-M-V-Shorts 1 | Shorts in fluid pathway or in area where Sterility is required | Blockage | Blockage / obstruction in the gate | Scrap | 4800.0 | 0.0 | Closure | CLOSED | 31-Jan-2020 | 13-Jan-2020 | 90. | 7 | Syringe
NC-IL081-17002 | 16-Oct-2019 | JESSE.SOBEL | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Out of Spec Visual Inspection (Visual Defect) | Additional rework sheets for 19E02563X added. | 19E02563X' '19E02463X | 40000.0 264000.0 | Line stopped and effected parts placed on hold. | 16-Oct-2019 | 15-Oct-2019 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Injection Molding | Inherent variation | Sort | 40000.0 264000.0 | 0.0 | Closure | CLOSED | 22-Jun-2020 | 15-May-2020 | 212. | Sub | Syringe
NC-IL081-17020 | 16-Oct-2019 | BENJAMIN.KELLY | General | Internal | S6P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Documentation Incomplete/Missing | Revision does not match template. Document was not retrieved through Agile per GS00100009. | 19J14063X | 335000.0 | No product affected. Non-standard S.O. continued to run on hold. | 16-Oct-2019 | 16-Oct-2019 | Documentation Incomplete/Missing | Documentation Incomplete/Missing | People/Human Error/Training |  | Other/NDF | 335000.0 | 0.0 | Closure | CLOSED | 15-Nov-2019 | 01-Nov-2019 | 15. | Sub | Syringe
NC-IL081-17021 | 16-Oct-2019 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | PT00015840 | Folding Carton, Printed, MONOJECT 12 mL Syringe, Regular Tip | Incorrect/Inadequate Labeling | During incoming inspection on 10/16/19 found out that the product label is missing the PO # (821936), this information is required per RMPS ( DC00035235 rev. 0) section 2.1. | 341519 | 25200.0 | Product put on hold. | 16-Oct-2019 | 16-Oct-2019 | Incorrect/Inadequate Labeling | Incorrect/Inadequate Labeling | Process/Method | Process/Method. Old customer service representative started the ordering process for this PO, and then the new customer service representative took over the account. | Other/NDF | 25200.0 | 0.0 | Closure | CLOSED | 20-Dec-2019 | 16-Dec-2019 | 60.
NC-IL081-17028 | 16-Oct-2019 | BENJAMIN.KELLY | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SPBP-V-Labeling/Print 1 | Extra labels at end of run label verification. | 19G17063X | 230400.0 | Completed End of Run, unable to reconcile labels for the inner cartons. Product placed on hold pending rework. - P. Myerson 10/16/2019 | 16-Oct-2019 | 16-Oct-2019 | SFF-SPBP-V-Labeling/Print 1 | Missing Lot Number /Label | Miss-communication | The Weber printer applicator on line 8 softpack was repaired mid-lot. The inner carton labels had already been pre-printed by the production clerk. The operators started using the labels on the line instead of closing out the lot of pre-printed labels per FM20007686 | Sort | 230400.0 | 0.0 | Closure | CLOSED | 08-Nov-2019 | 24-Oct-2019 | 8. | 8 | Syringe
NC-IL081-17034 | 16-Oct-2019 | ROSEMARY.TEKSTAR | General | Internal | C25 | Press 25-IL081 |  |  | 8507SA | CONTAINER SHARPS 5QT RED | CFF-V-Label Verification and Barcode Verification 5 | Incorrect lot number printed on label. Lot number printed as 59H11863, should read 19H11863. | 19H11863 | 480.0 |  | 16-Oct-2019 | 10-Oct-2019 | CFF-V-Label Verification and Barcode Verification 5 | OS - Bad/Missing pad print | Human Error | Human Error | Rework | 480.0 | 0.0 | Closure | CLOSED | 27-Jan-2020 | 22-Jan-2020 | 98. | SS | Container
NC-IL081-17035 | 16-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Past Due | QA-0019; Calibration overdue. (12&18 Gallon Min FIP width gauge) |  |  |  | 16-Oct-2019 | 15-Aug-2019 | Calibration Past Due | Calibration Past Due | Past Due |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 305.
NC-IL081-17041 | 16-Oct-2019 | BENJAMIN.KELLY | General | Internal |  |  |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Particulate 1 | SCAR received from Customer. Extraneous matter in the fluid pathway. SCAR # 19-016 | 18K10263X' '18K10663X' '18J10663X' '18J17063X | 1.0 |  | 16-Oct-2019 | 16-Oct-2019 | PR- Contaminated raw material | Contaminated raw material-MEDMFGPR | PR-Not Applicable | Not Applicable-MEDMFGPR | Return to Vendor | 1.0 | 0.0 | Closure | CLOSED | 22-Nov-2019 | 15-Nov-2019 | 29. | Sub | Syringe
NC-IL081-17050 | 17-Oct-2019 | JESSE.SOBEL | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Out of Spec Visual Inspection (Visual Defect) | Damaged plungers, functional but less than desired. | 19E02463X | 17000.0 | Post product release, effected product placed on hold. | 17-Oct-2019 | 15-Oct-2019 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Machine Failure | Machine Failure | Scrap | 17000.0 | 0.0 | Closure | CLOSED | 25-Jan-2020 | 15-Jan-2020 | 90. | Sub | Syringe
NC-IL081-17116 | 18-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Past Due | QA-0330, Calibration overdue. Item quarantined before due date. Became overdue waiting for vendor activation in "Buy Now". Sent to alternate calibration vendor on 10/17/2019. |  |  |  | 18-Oct-2019 | 18-Oct-2019 | Calibration Over Due | Calibration Over Due | PR-Personnel | Personnel-MEDMFGPR |  |  | 0.0 | Closure | CLOSED | 23-Jan-2020 | 15-Jul-2020 | 271.
NC-IL081-17117 | 18-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Past Due | QA-347, Calibration overdue. Item quarantined before due date. Became overdue waiting for vendor activation in "Buy Now". Sent to alternate calibration vendor on 10/17/2019. |  |  |  | 18-Oct-2019 | 18-Oct-2019 | Calibration Over Due | QA-347 became overdue waiting for vendor activation in the BuyNow program. Sent to alternate calibration vendor on 10/17/2019. The BuyNow system was implemented by Crystal Lake Calibration in October 2019. At that time, not all Calibration suppliers were entered and available for use. | PR-Purchasing Control | Purchasing Control-MEDMFGPR |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 304.
NC-IL081-17231 | 22-Oct-2019 | MARK.VENTURA | General | Internal | S4A | 412A-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-M-V-Bubbles 1 | Bubble in fluid pathway | 19K03363X | 99000.0 | Production stopped, Line cleared | 22-Oct-2019 | 22-Oct-2019 | SFF-M-V-Bubbles 1 | Gas Bubbles in critical area such as barrel luer, fluid pathway, ability to create leakage or affect sterility or in barrel walls that during use may be interpreted as bubbles in the content (i.e. in medicine, etc…) | Tool Cleaning |  | Sort | 99000.0 | 0.0 | Closure | CLOSED | 30-Mar-2020 | 23-Mar-2020 | 152. | 412 | Syringe
NC-IL081-17233 | 22-Oct-2019 | MARK.VENTURA | General | Internal | S4R | RAM-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-SA-V-Damaged Components 3 | Damage outside of barrel, Functional but less that desired | 19K03363X | 99000.0 | Stopped production, cleared line | 22-Oct-2019 | 22-Oct-2019 | SFF-SA-V-Damaged Components 3 | Damaged Components Type 3: Damaged product/component (functional but less than desired); Scratches/drag marks on the barrel | Machine Failure | Machine Failure | Sort | 99000.0 | 0.0 | Closure | CLOSED | 30-Mar-2020 | 23-Mar-2020 | 152. | 412 | Syringe
NC-IL081-17234 | 22-Oct-2019 | JACLYN.GRISKO | General | Internal | S1R | RAM-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | Incorrect/Incomplete Assembly | RGA; CASE-2019-00089396 Incomplete/Incorrect assembly, Plunger/tip not fully seated into barrel. | 19G16163X | 496800.0 | N/A - Post production NCR | 22-Oct-2019 | 22-Oct-2019 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Machine |  | Rework | 496800.0 | 0.0 | Closure | CLOSED | 19-Feb-2020 | 08-May-2020 | 198. | 1 | Syringe
NC-IL081-17236 | 22-Oct-2019 | MARK.VENTURA | General | Internal |  |  |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-M-V-Flash/Gate 3 | Flash on plunger ejector pins exceed 1/16" | 19K03363X | 99000.0 | Stopped production, cleared line | 22-Oct-2019 | [7 empty] | 0.0 | Initiation Verification | VOID |  |  |  | 412 | Syringe
NC-IL081-17377 | 26-Oct-2019 | JESSE.SOBEL | General | Internal | S6P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Environmental Monitoring | Insect found on floor near door on line 6. No product impacted. Finished good and lot number used as placeholder for NCR | 19K16963X | 1.0 | Clean floor and hopper with alcohol . F. Brena 10-25-2019 | 25-Oct-2019 | 25-Oct-2019 | Contamination - Insect/Rodent | Contamination - Insect/Rodent | pr-not applicable | Not Applicable-MEDMFGPR | No Issue / No Impact | 1.0 | 0.0 | Closure | CLOSED | 14-Jan-2020 | 07-Jan-2020 | 73. | Sub | Syringe
NC-IL081-17380 | 26-Oct-2019 | MARK.VENTURA | General | Internal | S8K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Particulate 2 | Excessive lubricant outside of barrel on finger flange | 19K12663X' '19K12563X' '19K12463X' '19K12163X' '19K12263X' '19K12363X | 46000.0 335000.0 | Lot cut short. Maintenance replaced heads 1,2,3 and 4 due to excessive drool on heads, purged for air after change.. K Kaecker | 26-Oct-2019 | 26-Oct-2019 | SFF-SA-V-Particulate 2 | Particulate Type 2: Extraneous matter outside of syringe; Excessive lubricant in or on the syringe | Machine Failure | CAPA-IL081-01128 has been investigated which identified causes for excess silicone. See this CAPA for the investigation. | Sort | 46000.0 335000.0 | 0.0 | Closure | CLOSED | 28-May-2020 | 16-Sep-2020 | 325. | Sub | Syringe
NC-IL081-17392 | 28-Oct-2019 | ROSEMARY.TEKSTAR | General | Internal | C61 | Press 61-IL081 |  |  | WP7417P | 8618RC LID | Out of Spec Product Functional | Failed leak test, position "A". | 19K190387 | 1250.0 | Completed a head change and completed a leak test. JR. Cruz. | 28-Oct-2019 | 26-Oct-2019 | CFF-F-Assembly 2 | Leak Test Failure | Process Not Robust | Process Not Robust | Scrap | 1250.0 | 0.0 | Closure | CLOSED | 04-Jul-2020 | 29-May-2020 | 214. | G18 | Container
NC-IL081-17397 | 28-Oct-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0588; FIP gasket gage; Gage could not be located for calibration. |  |  |  | 28-Oct-2019 | 22-Oct-2019 | Calibration Lost Gauge | Calibration Lost Gauge | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 293.
NC-IL081-17400 | 28-Oct-2019 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Water in package/on tip | Water/condensation was found in the bag of rubber tips for 4001204, lot 19K25-1-12 (pallet 9) Zone # 318 during incoming inspection on 10/28/19, failing incoming inspection for presence of water in package or on/in rubber tips, Inspection Level II, AQL 0.065, (A=0/R=1) per GS20003456, Rev. 2. | 19K25-1-12 | 4320000.0 |  | 28-Oct-2019 | 28-Oct-2019 | Visual- Water in package/on tip | Visual- Water in package/on tip | PR-Undetermined | PR-Undetermined-MEDMFGPR | Rework | 4320000.0 | 0.0 | Closure | CLOSED | 27-Nov-2019 | 12-Nov-2019 | 14. | Tips | Syringe
NC-IL081-17457 | 29-Oct-2019 | BENJAMIN.KELLY | General | Internal | S8A | 8A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Out of Spec Dimensional (Poor Dimensional Conformance) | Luer out of Spec. | 19K15763X | 68000.0 | Stopped and cleared line | 29-Oct-2019 | 29-Oct-2019 | Out of Spec Dimensional (Poor Dimensional Conformance) | Out of Spec Dimensional (Poor Dimensional Conformance) | Human Error | Human Error\n\nTool Room Technician (Significant – Improperly machined luer tips were used to repair 58118. This is most likely caused by the improper storage of components within the toolboxes. There was a historically NCR which removed all of the improperly machined luer tip components. The corrective and preventative actions from this historical NCR would not have been sufficient to remove the unauthorized storage of components within the toolboxes. This appears to have been a one off occurrence based on the timeframe since the historical NCR, no trend was observed.) | Other/NDF | 68000.0 | 0.0 | Closure | CLOSED | 16-Feb-2020 | 05-Feb-2020 | 98. | Sub | Syringe
NC-IL081-17555 | 01-Nov-2019 | MARIA.ROSALES | General | Internal | S7P | Packer-IL081 |  |  | 8881112985 | 10ML STERILE WATER SYR BULK NS | Fails UDI Grade Test | Fails UDI Grade Test | 19K10763X | 38000.0 | Stopped Production, Cleared line & Placed effected product on hold. | 31-Oct-2019 | 31-Oct-2019 | Fails UDI Grade Test | Fails UDI Grade Test | Sensors: Failure, Dirty, Positioning | Sensors: Failure, Dirty, Positioning | Rework | 38000.0 | 0.0 | Closure | CLOSED | 29-Jan-2020 | 06-Jan-2020 | 66. | 7 | Syringe
NC-IL081-17562 | 01-Nov-2019 | MARIA.ROSALES | General | Internal | S2P | Packer-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | Incorrect expiration date on over shipper label.er | Incorrect expiration date on over shipper label ( use by and expiration date should have the same date but they are different. Use by is 2024-10-30 and expiration date is 2024-10-31. | 19K10263X | 43680.0 | Stopped production, cleared line and placed effected product on hold. | 31-Oct-2019 | 31-Oct-2019 | SFF-SPBP-V-Labeling/Print 3 | Print does not match drawing | IT/Computer Systems | IT/Computer Systems | Rework | 43680.0 | 0.0 | Closure | CLOSED | 30-Jan-2020 | 15-Jan-2020 | 75. | 3 | Syringe
NC-IL081-17618 | 04-Nov-2019 | JESSE.SOBEL | General | Internal | S8B | 8B-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Scratch on cavity 11E that can be felt with finger on outside of barrel. | 19K16363X' '19K16263X | 15000.0 5000.0 | Line Stopped and cleared. Began check back procedure & all effected product placed on hold. | 04-Nov-2019 | 03-Nov-2019 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Machine Failure | Machine Failure | Scrap | 15000.0 5000.0 | 0.0 | Closure | CLOSED | 18-Jan-2020 | 10-Jan-2020 | 67. | Sub | Syringe
NC-IL081-17653 | 05-Nov-2019 | MARK.VENTURA | General | Internal | S2B | 2B-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SA-V-Damaged Components 3 | Scratch on outside of barrel. Can be felt with finger cavity 21B. | 19K18263X | 17760.0 | Cavity 21B was blocked, all affected product was scraped. F. Bahena 11/05/2019 | 05-Nov-2019 | 05-Nov-2019 | SFF-SA-V-Damaged Components 3 | Damaged Components Type 3: Damaged product/component (functional but less than desired); Scratches/drag marks on the barrel | Machine Failure | Machine Failure | Scrap | 17760.0 | 0.0 | Closure | CLOSED | 24-Jan-2020 | 16-Jan-2020 | 72. | 3 | Syringe
NC-IL081-17733 | 06-Nov-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Past Due | LVS-013, damaged repairable, will become calibration overdue 11/09/2019. Unit is generating false failures according to production personnel. Due to staff hours reduction, it is not possible to investigate the issue at this time. The unit is labelled "do not use until repaired, tested, and calibrated" and is quarantined in the calibration lab. |  |  |  | 06-Nov-2019 | 30-Oct-2019 | Calibration Past Due | Calibration Past Due | Past Due |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 284.
NC-IL081-17798 | 07-Nov-2019 | BERTHA.HERNANDEZ | General | Internal | S8B | 8B-IL081 | Equipment ID | 58119-IL081 | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-M-V-Plugs/Broken Pins 2 | Missing pins crated a plugged opening on cavity 31E | 19G23363X | 2400.0 | Line cleared of all sub-assemblies and barrels. 2 brutes of barrels from 8B scrapped. Cavity Blocked. P. Myerson 11/07/2019 | 07-Nov-2019 | 07-Nov-2019 | SFF-M-V-Plugs/Broken Pins 2 | Missing pins created a plugged opening --Defect occurred when the pin broke off in cavity #31E of the 58119 tool, in machine 8B, on line 8. The tool room was brought out to examine the tool. The tip snapped off when the tool closed. | Tooling/ fixture damaged | Damaged tooling/ fixture causing process variation Injection Molding (Significant – Defect occurred when the pin broke off in cavity #31E of the 58119 tool, in machine 8B, on line 8. The tool room was brought out to examine the tool. The tip snapped off when the tool closed. The possible causes of this condition are:\n\nPlastic residue staying in the cavity when closed\nInsufficient Loctite adhesive\nMis-alignment of mold\n\nAlthough the root cause cannot be determined at this time, the evidence points to plastic residue residing in the mold. \u200bThe other tips on the mold were visually inspected for application of loctite, with no issues found. The mold was determined to be level and if the mold was not level, there would be multiple tip breakages that would have occurred. The molds are cleaned once per shift but it is possible that the area with plastic residue buildup was missed during the cleaning process. | Scrap | 2400.0 | 0.0 | Closure | CLOSED | 22-Nov-2019 | 12-Nov-2019 | 5. | 8 | Syringe
NC-IL081-17814 | 07-Nov-2019 | ROSEMARY.TEKSTAR | General | Internal | C42 | Press 42-IL081 | Equipment ID | R7046L-IL081 | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-V-Flash (Non-Functional Area), Gate | String flash at gate area exceeding 1". | 19H12763' '19H05663' '19G12363 | 6480.0 17060.0 49940.0 | Trimming gate instructions were placed on the job stand for trimming stringy gate. Work Order was created, CL-3141890. - Berenice Villatoro 11/08/2019 | 07-Nov-2019 | 06-Nov-2019 | CFF-V-Flash (Non-Functional Area), Gate | Sharp Flash; Surfaces A & B: greater than or equal to 0.080”; Surface C: greater than or equal to 0.125”; All Surfaces: more than 3 locations in total | Cavity Insert Diameter undersized | The R7046L tool was evaluated under work order CL-3143251. The cavity insert gate diameter was .001 undersized, restricting full movement of the valve stem. Over time the constant and consistent impact of the stem caused varying amounts of damage to the gate area which produced the gate strings. See attached MRB with investigation for further details. | Other/NDF | 49940.0 17060.0 6480.0 | 0.0 | Closure | CLOSED | 08-Jan-2020 | 07-Jan-2020 | 60. | SS | Container
NC-IL081-17815 | 07-Nov-2019 | JACLYN.GRISKO | General | Internal | S8K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Damaged Components 1 | Incident Description from Customer:\nSCAR-19-017: During production, crushed syringes were found at the filler or after inspection machine. Due to the high number of syringes found, a 100% inspection of the filled syringes had to be done. The remaining unused syringes were put on hold (about 124,000 syringes): some unopened boxes were verified and damaged syringes were found. See attached pictures. \n\nSCAR-19-018: During production, eight lots of 12cc syringes were found with issues like crushed / damaged syringes, holes in barrels, missing plunger and molding issue at the flange. The details are in the attached table. See attached pictures. | 19K17963X' '19K15863X' '19K15963X' '19K16063X' '19K16163X' '19K16263X' '19K16363X' '19K12463X' '19K16963X' '19K12563X' '19K12663X' '19K12963X' '19K15763X' '19J14163X' '19J14563X' '19K12163X' '19K12263X' '19K12363X' '19J14863X' '19K20663X' '19K20763X' '19K20863X' '19K16663X' '19E02463X' '19E02563X' '19K12763X' '19J04563X' '19F10063X' '19J10863X' '19J10963X' '19K20563X' '19K12863X' '19J10663X' '19J13663X' '19J13563X' '19J13463X' '19J04463X' '19K16763X' '19K17063X' '19J14063X' '19K17163X' '19K17563X' '19K17663X' '19J10763X' '19J04663X' '19J10263X' '19J14763X' '19J14963X' '19K16863X' '19J04163X' '19J10363X' '19J10563X' '19J10463X' '19J10163X' '19J04763X | 335000.0 15000.0 46000.0 177000.0 309000.0 48000.0 334000.0 264000.0 162000.0 238000.0 274000.0 298000.0 312000.0 24000.0 124000.0 115000.0 154000.0 179000.0 199000.0 257000.0 | Reference investigation. | 07-Nov-2019 | 15-Sep-2019 | SFF-SA-V-Damaged Components 1 | Damaged Components Type 1: Damaged product/components (non-functional) | Machine Failure | Machine Failure | Hold' 'Sort | 124000.0 115000.0 154000.0 179000.0 199000.0 15000.0 335000.0 46000.0 48000.0 177000.0 309000.0 238000.0 334000.0 264000.0 162000.0 257000.0 274000.0 298000.0 312000.0 24000.0 | 7440.0 6900.0 9240.0 10740.0 11940.0 900.0 20100.0 2760.0 2880.0 10620.0 18540.0 14280.0 20040.0 15840.0 9720.0 15420.0 16440.0 17880.0 18720.0 1440.0 | Closure | CLOSED | 05-Jan-2020 | 23-Dec-2019 | 45. | Sub | Syringe
NC-IL081-17896 | 11-Nov-2019 | ROSEMARY.TEKSTAR | General | Internal | C77 | Press 77-IL081 |  |  | 31314886 | 3 GA GATORGUARD - TRANS RED | CFF-V-Label Verification and Barcode Verification 8 | Label position out of spec., positioned too high. Specifications: 1 5/8" +/- 3/16" from top of rim to full line. | 19H26963 | 276.0 | Process technician adjusted labeler to have label placement within specification. O. Ratinskii | 11-Nov-2019 | 11-Nov-2019 | 1010 | INCORRECT LABELING | Incorrect Labeling | Incorrect Labeling | Rework | 276.0 | 0.0 | Closure | CLOSED | 07-Jan-2020 | 07-Jan-2020 | 56. | GG | Container
NC-IL081-17898 | 11-Nov-2019 | ROSEMARY.TEKSTAR | General | Internal | C54 | Press 54-IL081 | Equipment ID | R7479C-IL081 | 8938 | CONTAINER SHARPS RED SLIDE LID DISP 18GL | CFF-V-Flash (Non-Functional Area), Gate | String flash at gates exceeding 1.00". | 19J06863' '19H22763' '19E16963 | 17970.0 9995.0 2475.0 | Pulled tool for repair. - J. Pixius 11/25/2019 for 11/11/2019 | 11-Nov-2019 | 30-Oct-2019 | CFF-V-Flash (Non-Functional Area), Gate | Sharp Flash; Surfaces A & B: greater than or equal to 0.080”; Surface C: greater than or equal to 0.125”; All Surfaces: more than 3 locations in total | Oversized cavity insert gate | The cavity insert gate diameters are .004 to .009 oversized allowing molten plastic to pull through the gate shutoff during mold open. See attached MRB for further details. | Other/NDF | 2475.0 17970.0 9995.0 | 0.0 | Closure | CLOSED | 31-Jan-2020 | 07-Jan-2020 | 56. | G18 | Container
NC-IL081-17940 | 12-Nov-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Over Due | QA-0171C, overdue due to staff hours reduction. |  |  |  | 12-Nov-2019 | 07-Nov-2019 | Calibration Past Due | QA-0171C became overdue waiting for vendor activation in the BuyNow program. | PR-Purchasing Control | The BuyNow system was implemented by Crystal Lake Calibration in October 2019. At that time, not all Calibration suppliers were entered and available for use. |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 278.
NC-IL081-17960 | 13-Nov-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Over Due | MD-0099, Calibration Overdue, due to reduction in staff hours. |  |  |  | 13-Nov-2019 | 07-Nov-2019 | Calibration Past Due | Calibration Past Due | PR-Personnel | Personnel-MEDMFGPR |  |  | 0.0 | Closure | CLOSED | 29-Apr-2020 | 10-Apr-2020 | 149.
NC-IL081-17963 | 13-Nov-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Over Due | QA-0275, Calibration Overdue, due to staff hours reduction. |  |  |  | 13-Nov-2019 | 07-Nov-2019 | Calibration Past Due | QA-0275 became overdue due to staff hours reduction. | PR-Personnel | Personnel-MEDMFGPR |  |  | 0.0 | Initiation Verification | VOID | 31-Jul-2020
NC-IL081-17971 | 13-Nov-2019 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Water in package/on tip | Water/condensation was found in the bag of rubber tips for 4001204, lot 19L08-1-12, box # 176,209,279 and 392 during incoming inspection on 11/13/19, failing incoming inspection for presence of water in package or on/in rubber tips, Inspection Level II, AQL 0.065, (A=0/R=1) per GS20003456, Rev. 2. | 19L08-1-12 | 4320000.0 |  | 13-Nov-2019 | 13-Nov-2019 | Visual- Water in package/on tip | Visual- Water in package/on tip | PR-Undetermined | PR-Undetermined-MEDMFGPR | Rework | 4320000.0 | 0.0 | Closure | CLOSED | 13-Dec-2019 | 26-Nov-2019 | 13. | Tips | Syringe
NC-IL081-17996 | 14-Nov-2019 | ROSEMARY.TEKSTAR | General | Internal | C81 | Press 81-IL081 |  |  | 31142222 | 2 GAL CONTAINER, RED | CFF-V-Quantity per Case | Finished good package missing 2 lids. | 19L04063 | 760.0 | Reviewed standard work with associate on proper packaging. S. Schirmer | 14-Nov-2019 | 14-Nov-2019 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | The Manufacturing Associate failed to follow Standard Work FM20007214 and failed to place the required amount of lids in each finished good case because they miscounted while multitasking. | Rework | 760.0 | 0.0 | Closure | CLOSED | 14-Dec-2019 | 15-Nov-2019 | 1. | G2 | Container
NC-IL081-18002 | 14-Nov-2019 | ROSEMARY.TEKSTAR | General | Internal | RSN | INCORRECT RESIN-IL081 |  |  | RM165 | Resin, Greenwood blend, Pre - colored Red | Raw Material - Does Not Meet Functional Test Specification | Batch history record revealed failed impact testing. | 00190646 | 22085.0 | All affected contained under BPCS hold 19110038. | 14-Nov-2019 | 14-Nov-2019 | Raw Material - Does Not Meet Functional Test Specification | Raw Material - Does Not Meet Functional Test Specification | Machine Failure | Machine Failure | Scrap | 22085.0 | 0.0 | Closure | CLOSED | 30-Apr-2020 | 17-Apr-2020 | 154. | Recy | Container
NC-IL081-18063 | 16-Nov-2019 | LAKEYSHA.SIMPSON | General | Internal |  |  |  |  | WP8536 | 8533SA LID | Missing PCR Readings | Process Tech failed to verify validated parameters once per shift per GS10004646. | 19H16936 | 0.0 |  | 16-Nov-2019 | [7 empty] | 0.0 | Initiation Verification | VOID |  |  |  | SS | Container
NC-IL081-18083 | 18-Nov-2019 | BERTHA.HERNANDEZ | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SPIC-V-Labeling/Print 1 | Incorrect expiration date on inner cartons labels. Found 2024-11-06, correct expiration date is 2024-11-16. Error was discovered at end of run label verification. | 19L02263X | 912.0 | Incorrect Exp. date found at EOR. line was cleared and the end of run was completed. Correct information was programmed for the new lot. _ P. Myerson 11-17-19 | 17-Nov-2019 | 17-Nov-2019 | SFF-SPIC-V-Labeling/Print 1 | Lot Number and Expiration Date (where applicable) is incomplete or missing | Human Error | Human Error\nThe normal manufacturing printing/labeling process utilizes the inline labeler/printer which at the time of this defect was not functioning. The manual labeling process is an approved alternate method. During the production run the amount of labels that had been programmed ran out, causing the operator to have to reprogram. When the operator re-programmed they did not verify the correct expiration date with the current lot control. The operator has to manually in-put the EXP information. It is unclear as to whom re-programmed the printer so a human error form could not be completed. | Rework | 912.0 | 0.0 | Closure | CLOSED | 16-Feb-2020 | 10-Feb-2020 | 84. | 8 | Syringe
NC-IL081-18100 | 18-Nov-2019 | ROSEMARY.TEKSTAR | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | Out of Spec Visual Inspection (Visual Defect) | Finished good case missing 2 lids. | 19H05863 | 180.0 | Retrained associated on FM20008307. B. Villatoro. | 18-Nov-2019 | 18-Nov-2019 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Human Error | The Manufacturing Associate failed to follow Standard Work FM20008307 section 3 which was available at the press workstation when this incident occurred. Human Error. | Rework | 180.0 | 0.0 | Closure | CLOSED | 25-Jan-2020 | 22-Jan-2020 | 64. | G2 | Container
NC-IL081-18131 | 19-Nov-2019 | BEATRIZ.GUILLEN | General | Internal | C81 | Press 81-IL081 |  |  | 31142222 | 2 GAL CONTAINER, RED | CFF-V-Label Verification and Barcode Verification 8 | Missing Biohazard Label and Label Placement is out of specification. QA Final case inspection, found missing label on one side of containers, and biohazard label upside down on the other side. | 19L11363 | 800.0 | Technician Adjusted the container stop on labeler. Steve Schirman 11/19/19 | 19-Nov-2019 | 19-Nov-2019 | CFF-V-Label Verification and Barcode Verification 8 | BH - Incorrect label placement | Mispositioned Label Stop |  | Rework | 800.0 | 0.0 | Closure | CLOSED | 19-Dec-2019 | 20-Nov-2019 | 1. | G2 | Container
NC-IL081-18151 | 19-Nov-2019 | MARK.VENTURA | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 1 | Drag mark that can be felt with bare finger Mold 58114/MP-1136 cavity 4B | 19l03463X' '19L03363X | 45000.0 43000.0 | Cleared line, place product on hold | 19-Nov-2019 | 19-Nov-2019 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 1 | Drag Marks/Gouges/Nicks Scratches Type 1: Material removed - Depth is measurable. Scratch inside Fluid Pathway - No Defects Allowed (Scratches/Scuffs on the inside of the barrel are not acceptable; Drag Marks that can be felt with a bare finger are not acceptable; Drag marks extended along the entire barrel length that can be felt with a bare finger) | Damaged gate | Due to normal wear and tear the gate was damaged / worn. | Scrap | 45000.0 43000.0 | 0.0 | Closure | CLOSED | 19-Jan-2020 | 10-Jan-2020 | 51. | Sub | Syringe
NC-IL081-18155 | 20-Nov-2019 | MARIA.ROSALES | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Flash in fluid pathway exceeds 1/64 | Flash in fluid pathway exceeds 1/64" mold 58110/MP-1134 cavity 46C & 37 | 19L03363X' '19L03263X' '19L03163X' '19L03063X' '19L02963X' '19L02863X' '19K21063X' '19K20963X' '19K20463X' '19K20363X | 43000.0 335000.0 | Closed cavity 37 and 46C and scrap all affected product. F. Bahena | 19-Nov-2019 | 19-Nov-2019 | Flash in fluid pathway exceeds 1/64 | Flash in fluid pathway exceeds 1/64 | Tooling/ fixture wear | Wear on tooling causes variation | Sort | 335000.0 43000.0 | 0.0 | Closure | CLOSED | 30-Apr-2020 | 09-Jun-2021 | 567. | Sub | Syringe
NC-IL081-18173 | 20-Nov-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Over Due | QA-335; Calibration overdue, due to staff hours reduction. |  |  |  | 20-Nov-2019 | 07-Nov-2019 | Calibration Past Due | Calibration Past Due | PR-Personnel | Personnel-MEDMFGPR |  |  | 0.0 | Initiation Verification | VOID | 31-Jul-2020
NC-IL081-18263 | 22-Nov-2019 | BENJAMIN.KELLY | General | Internal | S8K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Out of Spec Visual Inspection (Visual Defect) | Mfg. Check back found defective parts after operator observed parts getting damaged at the plunger insertion station of the Kahle. Parts also found missing Rubber plunger tips. | 19K16463X | 335000.0 | Line stopped and cleared, Mfg. began check back & Engineering shut down for further review. Engineering released back into production 11/22/19. - P. Myerson 11-22-19 | 22-Nov-2019 | 21-Nov-2019 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Machine Failure | Machine Failure | Sort | 265000.0 70000.0 | 0.0 | Closure | CLOSED | 21-Jan-2020 | 13-Jul-2020 | 234. | Sub | Syringe
NC-IL081-18292 | 24-Nov-2019 | MARIA.ROSALES | General | Internal | S8D | 8D-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 3 | Flash in Non-Functional area exceeds 1/16. ( Plunger cavity specific #38 E ) Mold 58119/MP-1121 | 19K16563X | 65000.0 | Stopped the line and presses. Closed cavity 38E and cut the lot short. L Reed 11/24/2019 | 24-Nov-2019 | 23-Nov-2019 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Blockage | Blockage | Sort | 65000.0 | 0.0 | Closure | CLOSED | 31-Jan-2020 | 15-Apr-2020 | 143. | Sub | Syringe
NC-IL081-18321 | 25-Nov-2019 | GRACE.AUSTIN | General | Internal | C74 | Press 74-IL081 |  |  | 8985 | 8 GA HINGED LID CHEMO -YELLOW | CFF-V-Packaging 1 | Missing absorbent Pads | 19J05763 | 180.0 | Verified packaging per Standard Work was correct and reviewed with associate. Ken Wiltse 11/25/19 | 25-Nov-2019 | 25-Nov-2019 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | Human Error | Rework | 180.0 | 0.0 | Closure | CLOSED | 24-Jan-2020 | 16-Dec-2019 | 20. | G8 | Container
NC-IL081-18332 | 26-Nov-2019 | MARIA.ROSALES | General | Internal | S8K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Damaged Components 1 | Damaged product ( Non-Functional ) found during rework of NC 18292 | 19K16563X | 2000.0 | QA/MFG performed checkback with no defects found | 26-Nov-2019 | 25-Nov-2019 | SFF-SA-V-Damaged Components 1 | Damaged Components Type 1: Damaged product/components (non-functional) | Machine Failure | Machine Failure | Sort | 2000.0 | 0.0 | Closure | CLOSED | 31-Jul-2020 | 27-Jul-2020 | 244. | Sub | Syringe
NC-IL081-18418 | 30-Nov-2019 | BENJAMIN.KELLY | General | Internal | S8C | 8C-IL081 | Equipment ID | 58120-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Burns/Gas Marks | Burn mark on thumb pad of plunger; Mold 58120, cavity 35D | 19L12463X | 8000.0 | Cleared line of all barrels and sub assemblies. Cavity was closed. Scrapped 8 cases. - P. Myerson 11/30/2019 | 30-Nov-2019 | 30-Nov-2019 | SFF-M-V-Burns/Gas Marks | Gas or Burn Marks on the part that renders it cosmetically unacceptable | Mold Wear or Damage | Mold Wear or Damage | Scrap | 8000.0 | 0.0 | Closure | CLOSED | 31-Dec-2019 | 04-Dec-2019 | 3. | Sub | Syringe
NC-IL081-18422 | 01-Dec-2019 | BENJAMIN.KELLY | General | Internal | S6C | 6C-IL081 | Equipment ID | 58114-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 1 | Flash in fluid pathway (Luer Tip) longer than 1/64th. Cav. 38B of mold 58114 | 19L12063X | 1.0 9600.0 | Cleared line of all barrels and sub-assemblies. Cavity blocked, work order# CL-3150166. - P. Myerson 12/01/2019 | 01-Dec-2019 | 01-Dec-2019 | SFF-M-V-Flash/Gate 1 | Flash Type 1: Flash in Fluid Path (Luer Tip) exceeds 1/64” (.016”). Does not meets drawing specification | Mold Wear or Damage | Mold Wear or Damage -Root Cause: – Defect resulted from core deflection. Core deflection is most likely caused due to force acted upon the core and not due to the operation of the mold. The two main contributors to force upon the core are (1) increased injection pressure or (2) parts/plastic stuck in the mold gate/wear housing affecting the flow of material. This force causes a shift in the core tip to wear insert shut off which allows plastic to get behind the wear bushing. Once this occurs, the defect will remain and only proceed to get worse over time as more plastic accumulates behind the wear bushing. It is unknown at this time which of these events contributed to this defect) | Scrap | 9600.0 1.0 | 0.0 | Closure | CLOSED | 31-Dec-2019 | 04-Dec-2019 | 2. | Sub | Syringe
NC-IL081-18447 | 02-Dec-2019 | BENJAMIN.KELLY | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 3 | Flash in non-functional area exceeding 1/16th". shoulder area Cavity # 34 Mold 58110/MP-1134 | 19L12163X | 1.0 | Cavity blocked and all affected product scrapped. Line restarted. . D. Brecede 12/02/2019 | 02-Dec-2019 | 02-Dec-2019 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Wear | Injection Molding (Significant MP-1134 (6A) running 58110 mold cavity 34 has wear / damage to the roof insert/cavity block shut off surfaces. The wear damage was determined to be cavity specific and not impact the other cavities in the mold. The most likely root cause of the damage occurring is built up material in the roof insert. ) | Scrap | 1.0 | 0.0 | Closure | CLOSED | 31-Jan-2020 | 08-Jan-2020 | 36. | Sub | Syringe
NC-IL081-18449 | 02-Dec-2019 | ROSEMARY.TEKSTAR | General | Internal | C52 | Press 52-IL081 |  |  | 8939 | CONTAINER 18 GAL CHEMO W SLIDE | Out of Spec Visual Inspection (Visual Defect) | String flash exceeding 1.00" at rim. | 19J04863 | 1235.0 | Placed trim instructions at the press. CL-3150729 was written to correct the tool when pulled. M. Ozimek | 02-Dec-2019 | 02-Dec-2019 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Mold Wear or Damage | Mold Wear or Damage - Plastic got behind the stripper plate causing string flash off of the container handle. | Rework | 1235.0 | 0.0 | Closure | CLOSED | 14-Feb-2020 | 10-Feb-2020 | 69. | G18 | Container
NC-IL081-18475 | 03-Dec-2019 | GRACE.AUSTIN | General | Internal | C81 | Press 81-IL081 |  |  | 31142222 | 2 GAL CONTAINER, RED | CFF-V-Quantity per Case | Missing Lids, missing biohazard label on one side of the containers, Product not properly packed per Packaging standard work per FM20007214 and DC00032963. | 19L11363 | 280.0 | Confirmed that proper standard work for labeling, packing and skid configuration were at the job stand. S. Schirmer | 03-Dec-2019 | 03-Dec-2019 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | Human Error | Rework | 280.0 | 0.0 | Closure | CLOSED | 01-Feb-2020 | 23-Dec-2019 | 19. | G2 | Container
NC-IL081-18511 | 04-Dec-2019 | GRACE.AUSTIN | General | Internal | C52 | Press 52-IL081 |  |  | 8998PG2 | SHARPS CONTAINER 18GAL RED | CFF-V-Label Verification and Barcode Verification 2 | Biohazard Label Print incomplete, Missing print on the Lot number and address information. | 19H28163 | 245.0 | Changed Ribbon and Label Roll, Tech also adjusted the pressure plate on the Printer Heat. Steve Schirmer 12/04/19 | 04-Dec-2019 | 04-Dec-2019 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Pressure Bar out of position |  | Rework | 245.0 | 0.0 | Closure | CLOSED | 02-Feb-2020 | 20-Jan-2020 | 47. | G18 | Container
NC-IL081-18517 | 04-Dec-2019 | PATTI.SWISHER | General | Internal | S8D | 8D-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Incorrect Document Used | MFG documented hourly inspections for plunger molds 58120 and 58121 on FM20007681 that was lined out\nwith no deviation | 19L12763X | 1.0 |  | 04-Dec-2019 | 03-Dec-2019 | Incorrect Document Used | Incorrect Document Used | Human Error | Human Error |  |  | 0.0 | Closure | CLOSED | 17-Mar-2020 | 17-Mar-2020 | 103. | Sub | Syringe
NC-IL081-18589 | 06-Dec-2019 | ROSEMARY.TEKSTAR | General | Internal | C54 | Press 54-IL081 |  |  | 8991 | 18 GA RENEWABLES HINGE LID | CFF-V-Short | Short/gas mark on handle. | 19K02563 | 115.0 |  | 06-Dec-2019 | 06-Dec-2019 | CFF-V-Short | Part is not completely formed; Missing plastic | Filter | Installed Filter on the nozzle was causing melt material not to fill mold completely resulting in little shorts on one of the handles. | Rework | 115.0 | 0.0 | Closure | CLOSED | 29-Feb-2020 | 19-Feb-2020 | 74. | G18 | Container
NC-IL081-18593 | 07-Dec-2019 | ROSEMARY.TEKSTAR | General | Internal | Raw Material | Raw Material |  |  | RM165 | Resin, Greenwood blend, Pre - colored Red | Unapproved Material | Untested/unreleased RM165 released to potential production silo. without testing or QA approval. | 190716 | 1700.0 | Production silo was not accessible to the manufacturing floor at time of impact. Quarantined released RM165 material from the unreleased RM165. Purged Silo 4 containing unreleased material into gaylords, placed on hold. Product contained under BPCS hold number 19120018 until testing could be completed and assessed for compliance. | 06-Dec-2019 | 06-Dec-2019 | Other | Operator working in autopilot due to familiar task. Transferred untested/unreleased material from recycler silo to manufacturing silo #5. | Man | Human error. | Other/NDF | 1700.0 | 0.0 | Closure | CLOSED | 04-May-2020 | 09-Apr-2020 | 124. | Recy | Container
NC-IL081-18601 | 09-Dec-2019 | MARIA.ROSALES | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 1 | Flash in Non-Functional area exceeds 1/16" finger flange. Cavity 25 Mold 58110/MP-1134 | 19L15463X | 65000.0 | Line was stopped after discovery of defect. Cleared out hopper and Khale. Closed cavity 25 on press 6A. Cut lot short and created work order # CL-3152012. Logan Reed 12/08/2019 | 08-Dec-2019 | 08-Dec-2019 | SFF-M-V-Flash/Gate 1 | Flash Type 1: Flash in Fluid Path (Luer Tip) exceeds 1/64” (.016”). Does not meets drawing specification | Mold Wear or Damage | Mold Wear or Damage | Sort | 65000.0 | 0.0 | Closure | CLOSED | 31-Jan-2020 | 28-Jan-2020 | 50. | Sub | Syringe
NC-IL081-18619 | 09-Dec-2019 | ROSEMARY.TEKSTAR | General | Internal | C42 | Press 42-IL081 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-V-Short | Shorts on handles and rims. | 19H27463 | 180.0 | Found that the hopper slide was partially closed. We had a technician open and secure it. S. Schirmer | 09-Dec-2019 | 09-Dec-2019 | CFF-V-Short | Part is not completely formed; Missing plastic | Loose Hopper Slide |  | Rework | 180.0 | 0.0 | Closure | CLOSED | 07-Feb-2020 | 09-Jan-2020 | 30. | SS | Container
NC-IL081-18643 | 10-Dec-2019 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Past Due | Calibration past due, gage ID: 5. Calibration past due, due to a reduction in staff hours. |  |  |  | 10-Dec-2019 | 05-Dec-2019 | Calibration Past Due | Calibration Past Due | Past Due |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 251.
NC-IL081-18775 | 14-Dec-2019 | MARIA.ROSALES | General | Internal |  |  |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-M-V-Plugs/Broken Pins 2 | Broken pin cavity #46C | 19L16563X | 1.0 | Stopped the line, Closed the cavity (46C), created W/O CL-3153132. Scrap affected work in process components.\n L. Reed 12/13/19 | 13-Dec-2019 | 13-Dec-2019 | SFF-M-V-Plugs/Broken Pins 2 | Missing pins created a plugged opening | Mold Wear or Damage | Mold Wear or Damage | Scrap | 1.0 | 0.0 | Closure | CLOSED | 29-Feb-2020 | 21-Feb-2020 | 69. | 7 | Syringe
NC-IL081-18780 | 14-Dec-2019 | BERTHA.HERNANDEZ | General | Internal | S4R | RAM-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-SA-V-Damaged Components 3 | Damaged components (functional but less than desired) big scratch outside of barrel body. | 19L13763X | 1500.0 | Removed all barrel components, Started press, Inspected molded barrels. Ran barrels through printer after adjustments to the fingers. L. Reed 12/18/19 | 14-Dec-2019 | 14-Dec-2019 | SFF-SA-V-Damaged Components 3 | Damaged Components Type 3: Damaged product/component (functional but less than desired); Scratches/drag marks on the barrel | Equipment variation requiring maintenance adjustments | Equipment variation requiring maintenance adjustments - Maintenance personnel found three areas that could have been a contributing factor to this defect. As the parts were transporting through the Ram AM-222 they were getting caught due to a gouge on the screw, there was a burr on the rail and a burr on the exit tube. | Sort | 1500.0 | 0.0 | Closure | CLOSED | 13-Mar-2020 | 04-Mar-2020 | 81. | 412 | Syringe
NC-IL081-18789 | 15-Dec-2019 | JESSE.SOBEL | General | Internal | S8B | 8B-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | Out of Validated Parameter - Equipment Failure | Charge recovery time out of parameter for MP-1133. | 19M00463X | 115200.0 | Charge Recovery Time Adjusted back into parameters. K. Kaecker 12/15/19 | 15-Dec-2019 | 15-Dec-2019 | Out of Validated Parameter - Equipment Failure | Out of Validated Parameter - Equipment Failure | Human Error | Human Error | Other/NDF | 115200.0 | 0.0 | Closure | CLOSED | 30-Apr-2020 | 08-Apr-2020 | 115. | 8 | Syringe
NC-IL081-18792 | 15-Dec-2019 | ANGELICA.AGUILAR | General | Internal | C61 | Press 61-IL081 |  |  | WP7417P | Slide Top, Foam In Place Lid, 12 and 18 Gallon Container, PG | Failed Leak Test | Leak test failed on position C | 19K190421 | 10650.0 | Cleared out the cell. Changed head and cleaned the ISO dosing valve. B. Villatoro | 15-Dec-2019 | 11-Dec-2019 | CFF-F-Assembly 2 | Leak Test Failure | Process Not Robust | Process Not Robust | Scrap | 10650.0 | 0.0 | Closure | CLOSED | 04-Jul-2020 | 29-May-2020 | 165. | G18 | Container
NC-IL081-18894 | 19-Dec-2019 | LAKEYSHA.SIMPSON | General | Internal | C31 | Press 31-IL081 |  |  | 8979MW | 2 GA RENEWABLES MULTI PURPOSE | CFF-V-Warping, Deformation and Pulling | Deformation/Pulling on the rim area cavity # 1 | 19M08563 | 3120.0 | Shutdown the line, removed affected material to QAHOLD. Re-stared the machine and change the ejector position.\nMinda Sailor 12/19/19 | 19-Dec-2019 | 18-Dec-2019 | CFF-V-Warping, Deformation and Pulling | Distortion affecting functionality | Mold Wear or Damage | Mold Wear or Damage | Sort | 3120.0 | 0.0 | Closure | CLOSED | 28-Feb-2020 | 18-Feb-2020 | 61. | G2 | Container
NC-IL081-18943 | 21-Dec-2019 | LAKEYSHA.SIMPSON | General | Internal | C45 | Press 45-IL081 | Equipment ID | R7460C-IL081 | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-V-Short | Shorts found on cavity #2 of R7060C, used to make product code WP8536. Molded component produced at press 45 and assembled at AM-171. | 19K19263 | 16990.0 2130.0 | Adjustments made: 0.02 of an in more of plastic, 4.00 to 4.50 2nd stage fill, .5 to 1.0 hold speed. - Melissa Gonzales 12/21/2019 | 20-Dec-2019 | 20-Dec-2019 | CFF-V-Short | Part is not completely formed; Missing plastic | Mold Wear or Damage | Mold Wear or Damage | Sort | 16990.0 2130.0 | 0.0 | Closure | CLOSED | 19-Feb-2020 | 20-Feb-2020 | 61. | SS | Container
NC-IL081-18946 | 21-Dec-2019 | PATTI.SWISHER | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Broken/Damaged 2 | Damaged product functional but less than desired, scratches/scuff marks that can be felt with a bare finger | 19M07063X | 39000.0 | Found a stuck part in the kahle, Cleared kahle and packer of all barrels and sub assemblies. P. Myerson 12/21/2019 | 21-Dec-2019 | 21-Dec-2019 | SFF-M-V-Broken/Damaged 2 | Broken or damaged product but still Functional | Misaligned assembly transfer | Misaligned assembly transfer | Scrap | 39000.0 | 0.0 | Closure | CLOSED | 30-Jan-2020 | 28-Jan-2020 | 38. | Sub | Syringe
NC-IL081-18947 | 22-Dec-2019 | JESSE.SOBEL | General | Internal | S8D | 8D-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-M-V-Broken/Damaged 2 | Broken plunger tip, but still functional. Cavity 36E mold 58120 | 19M07863X | 9600.0 | Stopped press, Tech found stuck part in mold, Removed and cleaned mold and restarted. D. Breceda 12/21/19 | 21-Dec-2019 | 21-Dec-2019 | SFF-M-V-Broken/Damaged 2 | Broken or damaged product but still Functional | Mold Wear or Damage | Mold Wear or Damage | Scrap | 9600.0 | 0.0 | Closure | CLOSED | 29-Feb-2020 | 28-Feb-2020 | 68. | 8 | Syringe
NC-IL081-18950 | 22-Dec-2019 | PATTI.SWISHER | General | Internal | S8D | 8D-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | Deficient Documentation | Expiry Date on shop order shown as 99/99/99 | 19M07963X | 1.0 | programmed labels for 5 years from date of start up. J. Pixius 12/22/2019 | 22-Dec-2019 | 22-Dec-2019 | Deficient Documentation |  | Incorrect Date | The shop order was printed with an incorrect expiry date reading “99/99/99.” Since this was the only shop order available at the time, the decision was made to run on a non-conformance. Everything else on the shop order was correct. | No Issue / No Impact | 1.0 | 0.0 | Closure | CLOSED | 20-Feb-2020 | 09-Jan-2020 | 17. | 8 | Syringe
NC-IL081-18978 | 26-Dec-2019 | ROSEMARY.TEKSTAR | General | Internal | C67 | Press 67-IL081 |  |  | 8931 | 12 GA HINGED CHEMO - YELLOW | CFF-V-Label Verification and Barcode Verification 12 | Incorrect biohazard label PT0006653 used. Correct label is PT00006646. | 19K05963 | 1040.0 | Brought correct biohazard labels to press and changed them out. J. Pixius | 26-Dec-2019 | 24-Dec-2019 | CFF-V-Label Verification and Barcode Verification 12 | BH - Wrong label | Human Error | Human Error | Rework | 1040.0 | 0.0 | Closure | CLOSED | 24-Feb-2020 | 29-Jan-2020 | 34. | G12 | Container
NC-IL081-18998 | 29-Dec-2019 | MARQUES.CALIMEE | General | Internal | C25 | Press 25-IL081 |  |  | 8507SA | CONTAINER SHARPS 5QT RED | CFF-V-Short | Shorts on locking tab of R7501D. WP8506 Lids was used into FG 8507SA | 19L04963 | 1700.0 | Process Tech. processed out the short. Guillermo Cruz 12/28/19 | 28-Dec-2019 | 28-Dec-2019 | CFF-V-Short | Part is not completely formed; Missing plastic | Mold Wear or Damage | Mold Wear or Damage | Rework | 1700.0 | 0.0 | Closure | CLOSED | 27-Mar-2020 | 04-Mar-2020 | 66. | SS | Container
NC-IL081-19000 | 29-Dec-2019 | BERTHA.HERNANDEZ | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Cracking 1 | Crack/hole in shoulder of barrel cavity 14 mold 58110. | 19M10963X' '19M10863X | 17000.0 | Maintenance Tech blocked defective cavity. No work order required. Ryan Hill 12/29/2019 | 29-Dec-2019 | 29-Dec-2019 | SFF-M-V-Cracking 1 | Cracking Type 1: No Crack allowed: leakage or loss of sterility as risk | Mold Wear or Damage | Mold Wear or Damage | Scrap | 17000.0 | 0.0 | Closure | CLOSED | 17-Apr-2020 | 23-Mar-2020 | 84. | Sub | Syringe
NC-IL081-19003 | 30-Dec-2019 | JESSE.SOBEL | General | Internal | S8C | 8C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Out of Spec Visual Inspection (Visual Defect) | burn on plunger rendering it cosmetically unacceptable. Cavity 35D/mold 58118/ MP-1132 | 19M16963X | 29000.0 | Maintenance Tech. closed cavity. No work order required. Ryan Hill 12-30-2019 | 30-Dec-2019 | 29-Dec-2019 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Mold Wear or Damage | Mold Wear or Damage | Scrap | 29000.0 | 0.0 | Closure | CLOSED | 17-Apr-2020 | 23-Mar-2020 | 84. | Sub | Syringe
NC-IL081-19026 | 02-Jan-2020 | BERTHA.HERNANDEZ | General | Internal | S8B | 8B-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Contamination/Specks 1 | Contamination/ Specks; inclusions in fluid pathway. | 4502570 | 31000.0 | Stopped and cleared the line of all barrels and sub-assys. Ran 6 consecutive shots from the press and verified no further contamination issues. Scrapped 31 cases of finished goods. P. Myerson 01/02/20 | 02-Jan-2020 | 02-Jan-2020 | SFF-M-V-Contamination/Specks 1 | Contaminations/Specks Type 1: Inclusions within the molded plastic Barrel luer, fluid pathway or ability to affect sterility - no defects allowed regardless of size or qty. | Burnt plastic | Defect resulted from plastic build up on the barrel and screw which burnt upon startup of the injection molding press after being down over 24 hours. The plastic cools and adheres to the barrel and screw, the burnt plastic breaks loose as the press runs and results in speck on the parts. | Scrap | 31000.0 | 0.0 | Closure | CLOSED | 15-Mar-2020 | 12-Mar-2020 | 70. | Sub | Syringe
NC-IL081-19056 | 03-Jan-2020 | BERTHA.HERNANDEZ | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | PR-Parameters out of specification | Back pressure out of parameters, found at 300psi during QA audit spec is 400-800psi. | 19M17263X | 50000.0 | Stopped and cleared the line, process tech changed back pressure back to validated parameter. changed the lot. P. Myerson 01/03/20 | 03-Jan-2020 | 03-Jan-2020 | PR-Parameters out of specification | Parameters out of specification-MEDMFGPR | Human Error | Process tech lowered the back pressure in order to purge the material due to contamination. The machine screw will not move backwards at a setting of 600 when in manual mode. Once he was done purging the press he forgot to change the setting back to the validated parameters. | Scrap | 50000.0 | 0.0 | Closure | CLOSED | 15-Mar-2020 | 12-Mar-2020 | 69. | Sub | Syringe
NC-IL081-19058 | 03-Jan-2020 | GRACE.AUSTIN | General | Internal | C70 | Press 70-IL081 |  |  | 065 | 1 QUART PHLEBOTOMY CONTAINER | Documentation Incorrect | Documented Low Wall Thickness of 0.064" bottom area cavity #4. Minimum is 0.065". Checkback revealed this was a documentation error only, no product affected. | 19K13763 | 1.0 | No line start-up correction, error found during EOR DHR inspection. Checkback revealed this was a documentation error only, no product affected. | 03-Jan-2020 | 03-Jan-2020 |  | Documentation error | Human Error | Human error. The associate failed to follow LS100000072, Wall Thickness Inspection, which calls for measuring material when the probe is in the stand, whenever possible to ensure it is perpendicular to the surface material. This allows the probe to be perpendicular to the surface of the material. When using the probe outside of the stand, the opportunity for error is induced causing false readings. |  |  | 0.0 | Closure | CLOSED | 02-Feb-2020 | 14-Jan-2020 | 10.
NC-IL081-19150 | 07-Jan-2020 | BERTHA.HERNANDEZ | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Damaged Components 3 | Damaged product / component (functional but less than desired). SCAR-20-001 customer found barrels slightly crushed (pinched). | 19L11663X | 335000.0 |  | 07-Jan-2020 | 07-Jan-2020 | 1160 | DAMAGED PRODUCT | Machine Failure | Machine Failure | Other/NDF | 335000.0 | 0.0 | Closure | CLOSED | 31-Mar-2020 | 12-Mar-2020 | 64. | Sub | Syringe
NC-IL081-19248 | 10-Jan-2020 | MARIA.ROSALES | General | Internal | S8B | 8B-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | Out of Spec Visual Inspection (Visual Defect) | Drag mark can be felt with a bare finger. Cavity 23E | 20A03563X | 24960.0 | Stopped line, blocked cavity 23E. created W/O CL-3164337. Discarded all barrels downstream from the press. L Reed 01/10/2020 | 10-Jan-2020 | 10-Jan-2020 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Mold Wear or Damage | Mold Wear or Damage | Scrap | 24960.0 | 0.0 | Closure | CLOSED | 13-Mar-2020 | 28-Feb-2020 | 49. | 8 | Syringe
NC-IL081-19271 | 10-Jan-2020 | PATTI.SWISHER | General | Internal | S3D | 3D-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | Out of Spec Visual Inspection (Visual Defect) | Short in area where sterility is required(Cap). 3D/MP-1110 mold 58040 cavity specific D1 , D2, 11-3 and 11-4 | 20A02863X | 8640.0 | Line stopped. Cleared all WIP caps and sub assemblies. Tech. made adjustment to process. K. Kaecker 01/10/2020 | 10-Jan-2020 | 10-Jan-2020 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Mold Heats |  | Sort | 8640.0 | 0.0 | Closure | CLOSED | 09-Feb-2020 | 20-Jan-2020 | 10. | 3 | Syringe
NC-IL081-19373 | 14-Jan-2020 | GRACE.AUSTIN | General | Internal | C52 | Press 52-IL081 |  |  | 8991 | 18 GA RENEWABLES HINGE LID | CFF-V-Label Verification and Barcode Verification 2 | Bad Print on Biohazard labels. Missing Print on UDI and lot numbers | 19L06863 | 525.0 | Removed and Red tagged Printer LP-127. Replaced with Green Tagged Printer LD-107. Did a label Verifiaction. Steve Schirmer 01/14/20 | 14-Jan-2020 | 14-Jan-2020 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Faulty Recoil Mechanism | Technician found that the Recoil mechanism was not working properly and needed to be repaired. | Rework | 525.0 | 0.0 | Closure | CLOSED | 19-Feb-2020 | 17-Feb-2020 | 34. | G18 | Container
NC-IL081-19435 | 15-Jan-2020 | MARK.VENTURA | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Documentation Missing | Failure to document PCR once per 12 hours. Line 6, press 6C/58114 | 20A07363X | 1.0 | Press history on machine verified for no process changes and PCR inspection in parameters. K. Kaecker 01/15/2020 | 15-Jan-2020 | 14-Jan-2020 | Documentation Missing | Documentation Missing | Human Error | The Technical operator failed to follow GS10004646 Rev 208 Section 6.3 which requires the PCR readings to be recorded and verified once per shift. The technical operator failed to take PCR readings on the 58114 MP-1136 per GS10004646. | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 14-Feb-2020 | 23-Jan-2020 | 8. | Sub | Syringe
NC-IL081-19510 | 17-Jan-2020 | GRACE.AUSTIN | General | Internal | C25 | Press 25-IL081 |  |  | 8507SA | CONTAINER SHARPS 5QT RED | CFF-F-Welding | Weld Strength Test Failed on one side of the Lids | 19L13363 | 720.0 280.0 | Work Order was created CL-3165757 for the encoder to be locked down again. Berenice Villatoro, 01/17/20 | 17-Jan-2020 | 17-Jan-2020 | CFF-F-Welding | Weld Strength | Defective | Defective components of welder needed replacement. | Scrap | 720.0 280.0 | 0.0 | Closure | CLOSED | 30-Apr-2020 | 24-Apr-2020 | 98. | SS | Container
NC-IL081-19543 | 18-Jan-2020 | GRACE.AUSTIN | General | Internal | C77 | Press 77-IL081 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-V-Incorrect Items | Incorrect IFU. CL63R9881A was packed. should be using CL63R8334A per BOM. Found during In-process Start-up.\nPress previously down from 2nd inspection. | 19L10063 | 540.0 | "Check Back was completed, Removed wrong IFU's from Job Stand and Replaced wrong IFU's with correct IFU's in all cases affected. Steve Schirmer 01/18/20" | 18-Jan-2020 | 18-Jan-2020 | CFF-V-Incorrect Items | Incorrect IFU/Absorbent pads/Other Kit items/Other Labels in the Overshipper Box | Human Error | The incorrect IFU’s were brought to the press due to human error. | Rework | 540.0 | 0.0 | Closure | CLOSED | 18-Mar-2020 | 28-Feb-2020 | 41. | SS | Container
NC-IL081-19706 | 24-Jan-2020 | GRACE.AUSTIN | General | Internal | C74 | Press 74-IL081 |  |  | 8608RC | 8 GA RCRA HAZARDOUS WASTE | Procedure Not Followed | Failed to follow procedure QS00002152. Deviation run without Non-standard shop order completed. |  |  | N/A, not required for Non-product NCRs. | 24-Jan-2020 | 24-Jan-2020 | Procedure Not Followed | Procedure Not Followed | Incomplete completion of documentation | Incomplete completion of documentation |  |  | 0.0 | Closure | CLOSED | 31-Mar-2020 | 12-Mar-2020 | 48. | G8 | Container
NC-IL081-19719 | 24-Jan-2020 | ROSEMARY.TEKSTAR | General | Internal | C74 | Press 74-IL081 | Equipment ID | R7449A-IL081 | 8608RC | 8 GA RCRA HAZARDOUS WASTE | Improper Test Procedure | Non-standard shop order 189702 (Product code WP7449) was run on hold under BPCS hold number 20010021. This was run to perform an ECR from Maximo Work Order CL-3125368 for the addition of the channel to the R7449A mold. While running this ECR, it was noticed that the gauging requirements for a channeled component were not yet updated. | 19J05863 | 2310.0 | Production was stopped and change orders were submitted to correct this issue. EC129566 was created to update the PCR to reflect only channeled tool requirements. EC129604 was created to update the inspection procedures and inspection form. - J. Grisko 02/24/2020 | 24-Jan-2020 | 24-Jan-2020 | Improper Test Procedure | Improper Test Procedure | Process/Method | Process/Method | Other/NDF | 2310.0 | 0.0 | Closure | CLOSED | 31-Mar-2020 | 25-Feb-2020 | 31. | G8 | Container
NC-IL081-19737 | 26-Jan-2020 | JESSE.SOBEL | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Contamination/Specks 2 | Group of specks greater than 4 mm^2 on cavity 37.Barrel mold 58110 in MP-1134 | 20A11663X' '20A11563X | 22000.0 12000.0 | Tech blocked defective cavity. Work Order# CL-3171243. R. Hill 01/26/2020 | 26-Jan-2020 | 25-Jan-2020 | SFF-M-V-Contamination/Specks 2 | Contamination/Specks Type 2/3: Inclusions in the non-fluid pathway. Length: Greater than 2 millimeters (0.078 inch), Area: Greater than 3.00 squared mm, or Groups or clusters: Sum of all bubble areas accumulated to a total of 4.00 squared mm shall be considered defective. Use Tappi chart. | Blocked Vents | Blocked Vents | Scrap | 22000.0 12000.0 | 0.0 | Closure | CLOSED | 15-Mar-2020 | 23-Mar-2020 | 57. | Sub | Syringe
NC-IL081-19739 | 26-Jan-2020 | ANGELICA.AGUILAR | General | Internal | C31 | Press 31-IL081 |  |  | 31300878 | 2 GAL CONTAINER, TRANS YELLOW | CFF-V-Packaging 1 | Missing (IFU) instruction sheet into finished product | 19L00463 | 280.0 | Retrained the associate on DC00047470 Standard Work. Berenice Villaroto. 01/26/20 | 26-Jan-2020 | 26-Jan-2020 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | The Manufacturing Associate failed to follow Standard Work DC00047470 section 3 which was available at the press workstation when this incident occurred. Human Error. | Sort | 280.0 | 0.0 | Closure | CLOSED | 26-Mar-2020 | 27-Mar-2020 | 60. | G2 | Container
NC-IL081-19740 | 26-Jan-2020 | BERTHA.HERNANDEZ | General | Internal | S8P | Packer-IL081 |  |  | 8881540122 | 12ML SMARTIP MED-PREP CANNULA | SFF-SPBP-V-Labeling/Print 1 | Label reconciliation; missing inner cartons labels at end of run label verification. | 20A12163X | 168480.0 | Completed End of Run, Removed all labels from the line. placed product on hold. P. Myerson 01/26/2020 | 26-Jan-2020 | 26-Jan-2020 | SFF-SPBP-V-Labeling/Print 1 | Missing Lot Number /Label | Human Error | Root Cause: Inconsistent printer issues resulting in blank/not legible labels being destroyed during the run which affected the label count. | Rework | 168480.0 | 0.0 | Closure | CLOSED | 31-Mar-2020 | 05-Mar-2020 | 38. | 8 | Syringe
NC-IL081-19755 | 27-Jan-2020 | BERTHA.HERNANDEZ | General | Internal | S8F | Form, Fill Sealer-IL081 |  |  | 8881540122 | 12ML SMARTIP MED-PREP CANNULA | SFF-PPSP-D-Seal Width | Seal integrity; seal has a void doesn\'t meets the 3mm (1/8") width on package # 9. | 20A12263X | 12480.0 | Cleared line of all sub assembly packages. Gasket was replaced. Affected product was scrapped and repackaged. P. Myerson 01/30/2020 | 27-Jan-2020 | 27-Jan-2020 | SFF-PPSP-D-Seal Width | Seal Width | Part Out of Pocket | Part Out of Pocket | Sort | 12480.0 | 0.0 | Closure | CLOSED | 30-Apr-2020 | 08-Apr-2020 | 72. | 8 | Syringe
NC-IL081-19774 | 27-Jan-2020 | BERTHA.HERNANDEZ | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-F-Functional | Functional, leakage cavity 24 Mold 58110, MP-1134/6A | 20A15263X | 198000.0 | Cleared line of all barrels and sub-assemblies, blocked cavity. P. Myerson 01/27/2020 | 27-Jan-2020 | 27-Jan-2020 | SFF-M-F-Functional | Functional Leak Test | ejection method | Barrels were being ejected off the press onto a hard plastic conveyor. There was nothing to help soften their decent. This caused a small defect to occur on the edge of the tip of the barrel. The tip was dented slightly. | Scrap | 198000.0 | 0.0 | Closure | CLOSED | 26-Apr-2020 | 24-Apr-2020 | 87. | Sub | Syringe
NC-IL081-19799 | 28-Jan-2020 | GRACE.AUSTIN | General | Internal | C52 | Press 52-IL081 |  |  | 8871 | 18 GA PHARMSAFETY - SLIDE TOP | CFF-V-Label Verification and Barcode Verification 17 | Wrinkled and Damaged Biohazard Labels packed in FG cases. | 19J06063 | 445.0 | Checked Back was completed and a Human error Form was issued to the associate responsible. Steve Schirmer 01/28/20 | 28-Jan-2020 | 28-Jan-2020 | CFF-V-Label Verification and Barcode Verification 17 | OS - Wrinkled labels | Human Error | Human error. The associate failed to follow Standard Work Labeling procedure, which states to smooth labels with hands, after labeling. | Rework | 445.0 | 0.0 | Closure | CLOSED | 28-Mar-2020 | 17-Feb-2020 | 20. | G18 | Container
NC-IL081-19823 | 29-Jan-2020 | ROSEMARY.TEKSTAR | General | Internal | C31 | Press 31-IL081 |  |  | 8982 | 2 GA HINGED CHEMO - YELLOW | Out of Spec Product Functional | Failed Lid Fit Test during scheduled manufacturing inspection. | 20A08263' '19M02063' '19K17463' '190589' '190750' '19L00963' '19L02663' '19E08663 | 620.0 12950.0 6610.0 16128.0 6144.0 1202.0 17360.0 680.0 |  | 28-Jan-2020 | 28-Jan-2020 | CFF-F-Molding 3 | Lid Fit | PR-Undetermined | PR-Undetermined-MEDMFGPR; Root cause will be determined under CAPA | Rework | 620.0 12950.0 6610.0 16128.0 6144.0 1202.0 17360.0 680.0 | 0.0 | Closure | CLOSED | 27-Apr-2020 | 21-Apr-2020 | 83. | G2 | Container
NC-IL081-19848 | 29-Jan-2020 | GRACE.AUSTIN | General | Internal | C67 | Press 67-IL081 |  |  | 31156550 | 7 GA SHARPS CONTAINER | Manual- Lid Fit failed | Failed Lid Fit inspection, due to the broken one lefter on the lids WP4810, it was use into FG. | 19J12763' '20A191416 | 5830.0 600.0 | R7567C pulled for repair (Broken Lifter) under Work Order# CL-3165164; 1st Articles inspection passed after repair. Ken Wiltse. 02/04/20 | 29-Jan-2020 | 29-Jan-2020 | Manual- Lid Fit failed | Manual- Lid Fit failed | Broken Insert | One of the inserts on the R7567C Tool broke and needed to be repaired. | Scrap | 5830.0 600.0 | 0.0 | Closure | CLOSED | 28-Apr-2020 | 27-Mar-2020 | 57. | G7 | Container
NC-IL081-19886 | 30-Jan-2020 | BERTHA.HERNANDEZ | General | Internal | S8K | Kahle-IL081 |  |  | 4502500 | SYR 12ml WFF Magellan Sub | SFF-SA-V-Mixed Product 2 | Incomplete assembly of product, plunger not fully inside barrel. | 20A16963X | 31000.0 | Cleared line of all sub-assemblies, adjusted the plunger inspect wheel by lowering. P. Myerson 01/30/2020 | 30-Jan-2020 | 30-Jan-2020 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Loose lock nut | When the mechanical maintenance person investigated the cause of this defect, he found that the lock nut on the adjustment was loose and the adjustment for the short plunger had moved. | Scrap | 31000.0 | 0.0 | Closure | CLOSED | 29-Feb-2020 | 03-Feb-2020 | 3. | 8 | Syringe
NC-IL081-19922 | 01-Feb-2020 | ALEJANDRA.CASTANEDA | General | Internal | C54 | Press 54-IL081 |  |  | 8617RC | 18 GAL RCRA HINGE TOP | CFF-V-Mixed Product 1 | Using incorrect lids WP8998 per BOM should be using WP8998P. | 19L19663 | 3010.0 | Removed all WP8998 lids and then used the correct WP8998P lids. S. Schirmer | 01-Feb-2020 | 01-Feb-2020 | CFF-V-Mixed Product 1 | Incorrect Container/Lids in the Overshipper Box | Human Error | Human Error | Rework | 3010.0 | 0.0 | Closure | CLOSED | 31-Mar-2020 | 12-Mar-2020 | 40. | G18 | Container
NC-IL081-19934 | 02-Feb-2020 | MARIA.ROSALES | General | Internal | S0A | S0A-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | Documentation Missing | Failed to perform start-up checklist DC00007100 for 0A, MP-1123/58116 plunger. | 20A17463X | 1.0 | Perform a start up for press 0A, first article and PCR readings were completed without a start up being completed. L. Reed 02/02/2020 | 01-Feb-2020 | 01-Feb-2020 | Documentation Missing | Documentation Missing | Human Error | Human Error | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 17-Apr-2020 | 23-Mar-2020 | 50. | 1 | Syringe
NC-IL081-19935 | 02-Feb-2020 | MARIA.ROSALES | General | Internal | S8N | Pad Printer-IL081 |  |  | 1181200777 | NLC NO MOVE (020814)SYR,MONOJECT SOFTPACK,12CC,LL, | SFF-P-V-Graduations 3 | Inconsistent graduations lines. Graduation lines do not have the same thickness. | 20A16063X | 168960.0 | Stopped the line, cleared all affected parts from line and kanban, cut the lot short and installed new blanket on the printer. Verified that the print was acceptable. L. Reed 02/01/2020 | 02-Feb-2020 | 01-Feb-2020 | SFF-P-V-Graduations 3 | Type 3: Inconsistent graduation lines. Graduation lines do not have same thickness | Print Pad Issue | Print Pad Issue | UseAsIs | 168960.0 | 16896.0 | Closure | CLOSED | 30-Apr-2020 | 17-Apr-2020 | 75. | 8 | Syringe
NC-IL081-20037 | 05-Feb-2020 | ANGELICA.AGUILAR | General | Internal | S8A | 8A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Out of Spec Visual Inspection (Visual Defect) | Embedded particulate and extraneous matter in barrel. mold 58118/MP-1132/8A | 20A21963X | 1.0 | Clear line of all barrels and sub-assemblies. checked shot from press and it was acceptable. P. Myerson 02/05/20 | 05-Feb-2020 | 05-Feb-2020 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Particulate | Defect appears to be embedded cardboard molded into the barrel. Most probable cause would be that a piece of cardboard was in the resin when the material was loaded into the silo and transported to the molding press. | Scrap | 1.0 | 0.0 | Closure | CLOSED | 05-May-2020 | 08-Apr-2020 | 62. | Sub | Syringe
NC-IL081-20064 | 06-Feb-2020 | JACLYN.GRISKO | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | PR- Cartons not loaded correctly on pallet or corner beads not installed | Notified by Sterilizer that the palatalization of the 777 code is inadequate for their process. The product is hanging off the sides of the pallet and jamming in the production area. The current configuration is 8 bundles (cubes) to a layer and five layers high. | 20A16263X' '20A16163X | 63.0 480.0 | No immediate actions required as this product is through the physical production process and pending release. | 06-Feb-2020 | 05-Feb-2020 | PR- Cartons not loaded correctly on pallet or corner beads not installed | Cartons not loaded correctly on pallet or corner beads not installed-MEDMFGPR | Carton loading | The pallet configuration was reviewed, and it was determined that the current pallet configuration is susceptible to excessive overhang if during transportation excessive force is applied (i.e. a traffic incident). | Rework | 63.0 480.0 | 0.0 | Closure | CLOSED | 06-Apr-2020 | 02-Apr-2020 | 56. | 8 | Syringe
NC-IL081-20143 | 09-Feb-2020 | ANGELICA.AGUILAR | General | Internal | C77 | Press 77-IL081 |  |  | 31314886 | 3 GA GATORGUARD - TRANS RED | CFF-D-Molding 1 | Container cav.1 with wall thickness out of spec. in the gate area. | 19M03663 | 1836.0 | Tool was pulled for repair, work order CL-3174619 | 08-Feb-2020 | 08-Feb-2020 | CFF-D-Molding 1 | Wall Thickness | Gate pin damage | The most probable root cause is a damage gate pin due to normal wear and tear conditions. | Rework | 1440.0 396.0 | 0.0 | Closure | CLOSED | 08-May-2020 | 28-Apr-2020 | 79. | GG | Container
NC-IL081-20144 | 09-Feb-2020 | ALEJANDRA.CASTANEDA | General | Internal | C54 | Press 54-IL081 |  |  | 8989CRH | 18 GAL CHEMO RED HINGE LID | CFF-V-Label Verification and Barcode Verification 10 | Containers not labeled on both side, missing one label. Per standard work FM20009172 containers should be labeled double side. | 19M14363 | 560.0 | Switched program to label both sides of containers. Guillermo Cruz. 02/08/20 | 08-Feb-2020 | 08-Feb-2020 | CFF-V-Label Verification and Barcode Verification 10 | BH - Missing label | People/Human Error/ | People/Human Error/ | Rework | 560.0 | 0.0 | Closure | CLOSED | 17-Apr-2020 | 23-Mar-2020 | 43. | G18 | Container
NC-IL081-20196 | 11-Feb-2020 | ROSEMARY.TEKSTAR | General | Internal | C63 | Press 63-IL081 |  |  | WP8998P | 8931PG2 LID | Out of Spec Product Functional | Lid lacking corona treatment, causing poor foam adhesion. | 190951' '19M09263 | 1890.0 140.0 | Reattached a broken wire to the corona that was causing the corona treatment to fail. Jr. Cruz | 10-Feb-2020 | 10-Feb-2020 | Out of Spec Product Functional | Out of Spec Product Functional | Machine Failure | Machine Failure | Rework | 140.0 1890.0 | 0.0 | Closure | CLOSED | 10-Apr-2020 | 03-Apr-2020 | 52. | G12 | Container
NC-IL081-20270 | 12-Feb-2020 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 5170588 | Poly Bag, Gusseted, 22.5 Inch x 16.25 Inch x 42.00 Inch | Packaging Damaged | On 2/12/20 1 case ( 225ea) was found damaged by warehouse personnel, failed incoming inspection for shipping damaged. | UP90005397 | 225.0 | Product placed on hold. | 12-Feb-2020 | 12-Feb-2020 | Visual- Shipping Damage | Visual- Shipping Damage | PR-Communication | Communication-MEDMFGPR | Scrap | 225.0 | 0.0 | Closure | CLOSED | 27-Mar-2020 | 09-Mar-2020 | 25. | IQA | Incoming
NC-IL081-20307 | 13-Feb-2020 | GRACE.AUSTIN | General | Internal | C67 | Press 67-IL081 |  |  | 31143665 | 10 GA SHARPS CONTAINER | CFF-V-Flash (Non-Functional Area), Gate | Flash on the Gate area of the containers more than 1" inch out of specification. Flash exceeding specification. | 20A00663 | 720.0 | The associate was trained on following standard work for Rework instructions for flash trimming FM20011812. Berenice Villatoro 02/13/20 | 13-Feb-2020 | 13-Feb-2020 | CFF-V-Flash (Non-Functional Area), Gate | Sharp Flash; Surfaces A & B: greater than or equal to 0.080”; Surface C: greater than or equal to 0.125”; All Surfaces: more than 3 locations in total | Worn gate pins | The gate pins were worn, not allowing proper shut off. This is a normal expectation of the manufacturing process. | Rework | 720.0 | 0.0 | Closure | CLOSED | 13-Apr-2020 | 02-Apr-2020 | 49. | G10 | Container
NC-IL081-20314 | 13-Feb-2020 | ROSEMARY.TEKSTAR | General | Internal | C67 | Press 67-IL081 |  | R7567C-IL081 | 31143665 | 10 GA SHARPS CONTAINER | CFF-V-Flash (Functional Area) | Flash exceeding specifications in functional area of WP4810 lid. | 20A00663' '191103' '19J12763' '19J06363' '19M12063 | 32440.0 1.0 24600.0 10800.0 4800.0 | Pulled tool for repair. - J. Pixius | 13-Feb-2020 | 13-Feb-2020 | CFF-V-Flash (Functional Area) | Sharp Flash; Hinge Pins and Holes; All other areas: Less than or equal to 0.060" and no more than 3 locations per part. | Parting line damage | Damage to the parting line caused insert pocket to get plastic. Tool was pulled for repair CL-3175761. | Other/NDF | 32440.0 24600.0 10800.0 4800.0 1.0 | 0.0 | Closure | CLOSED | 13-Apr-2020 | 10-Apr-2020 | 57. | G10 | Container
NC-IL081-20316 | 13-Feb-2020 | BERTHA.HERNANDEZ | General | Internal | S1A | 1A-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-M-V-Deformation 2 | Deformation on barrels caused by thermolator being off. | 20A17963X | 12000.0 | Cleared line of all barrels and sub-assemblies. scrapped 15 cs. Thermolater found to be off, restarted. P. Myerson 02/13/2020 | 13-Feb-2020 | 13-Feb-2020 | SFF-M-V-Deformation 2 | The dimension of the barrel end is different than specification | Human Error | Human Error - An object was placed near the control panel for the thermolator which caused the thermolator switch to be bumped and turned off. | Scrap | 12000.0 | 0.0 | Closure | CLOSED | 31-Mar-2020 | 24-Feb-2020 | 11. | 1 | Syringe
NC-IL081-20382 | 16-Feb-2020 | LAKEYSHA.SIMPSON | General | Internal |  |  |  |  | 8507SAS | CAH BRAND SHELL | failed mold repair dimensional inspection | Failed mold repair dimensional inspection. | 20A01163 | 480.0 |  | 15-Feb-2020 | [7 empty] | 0.0 | Initiation Verification | VOID
NC-IL081-20425 | 17-Feb-2020 | FIDELIA.CANTOR | Inspection | Supplier | INC | Incoming-INC-IL081 |  |  | 5000231 | Gusseted Polybag, Unprinted, 13.5 Inch x 14.25 Inch x 5.5 In | Visual- Contamination (not for rubber tips) | During incoming inspection of 500231, lot 19002297019042D, on 2/17/20, found 2cs with black, loose, foreign contamination, failing incoming inspection AQL is S-4, 0.065 (A-0, R-1). | 19002297019042D | 25200.0 | Placed product on hold. | 17-Feb-2020 | 17-Feb-2020 | Visual- Contamination (not for rubber tips) | Visual- Contamination (not for rubber tips) | Failure to Follow Procedures/Policies | Failure to Follow Procedures/Policies- Operator failed to clean draw rollers prior to production run. | Return to Vendor | 25200.0 | 0.0 | Closure | CLOSED | 27-Mar-2020 | 26-Mar-2020 | 37.
NC-IL081-20463 | 18-Feb-2020 | ROSEMARY.TEKSTAR | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | Out of Spec Visual Inspection (Visual Defect) | Finished good case missing four lids. | 19L09263 | 240.0 | Train associated on SW FM20008307 for packaging. B. Villatoro | 18-Feb-2020 | 18-Feb-2020 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Human Error | The Manufacturing Associate failed to follow Standard Work FM20008307 section 3 which was available at the press workstation when this incident occurred. Human Error. | Rework | 240.0 | 0.0 | Closure | CLOSED | 18-Apr-2020 | 23-Mar-2020 | 33. | G2 | Container
NC-IL081-20631 | 23-Feb-2020 | PATTI.SWISHER | General | Internal | S7C | 7C-IL081 |  |  | 8881907003 | 10ML ORAL MED SYRINGE (AMBER) | SFF-M-V-Deformation 2 | Deformed shoulder cavity #45 (concave) mold 58020-03 | 20B03363X | 1.0 | Cleared line of all barrels and sub assemblies, Tech polished pins. P. Myerson 02/23/2020 | 23-Feb-2020 | 23-Feb-2020 | SFF-M-V-Deformation 2 | The dimension of the barrel end is different than specification | Dirty pins | Build up / dirty pins, which caused parts to stick on the mold. | Scrap | 1.0 | 0.0 | Closure | CLOSED | 30-Apr-2020 | 13-Apr-2020 | 50. | 7 | Syringe
NC-IL081-20633 | 23-Feb-2020 | PATTI.SWISHER | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 3 | Flash in non functional area. finger flange cavity # 16B | 20B14763X | 5000.0 | Cleared all barrels and sub assemblies from the line Tech cleared out a piece of stuck plastic from the mold. P. Myerson 02/23-20 | 23-Feb-2020 | 23-Feb-2020 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Resin Build up | Defect resulted from flash build up around the cavity block. | Scrap | 5000.0 | 0.0 | Closure | CLOSED | 30-Apr-2020 | 13-Apr-2020 | 49. | Sub | Syringe
NC-IL081-20648 | 24-Feb-2020 | GRACE.AUSTIN | General | Internal |  |  |  |  | WP7449 | Hotair Staked Chemotherapy Lid, 8 Gallon | CFF-F-Assembly 2 | Failed Leak Test Inspection, position C | 20A191530' '19K01863 | 1110.0 2460.0 | 1. Increased flow rate for gasket material. Proper gasket height confirmed by inspection gauges. B. Dominguez.\n2. Change head and mixing disk. Guillermo Cruz 02/24/20 | 24-Feb-2020 | 24-Feb-2020 | CFF-F-Assembly 2 | Leak Test Failure | Dirty Mixing chamber | Dirty mixing chamber affected the flow of the gasket rate | Scrap | 2460.0 1110.0 | 0.0 | Closure | CLOSED | 24-May-2020 | 04-May-2020 | 70. | G8 | Container
NC-IL081-20753 | 27-Feb-2020 | GRACE.AUSTIN | General | Internal | C42 | Press 42-IL081 |  |  | 8836SA | 3GA PHARMASTAR - TRANS WHITE | CFF-V-Label Verification and Barcode Verification 2 | Incorrect Lot number printed on Overshipper labels, (10)20A191390 it should be printed (10) 20A05863 per Shop Order | 20A05863 | 900.0 | Reprogram the labeler with correct shop order. B. Doninguez | 27-Feb-2020 | 27-Feb-2020 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Human Error | The manufacturing process is not in line and the shop order allocation is transposed for the stations manufacturing the applicable components. The Manufacturing Associate used the incorrect shop order (WIP) to make the FG label and did not verify after printing. This resulted in the incorrect information being on the FG label. | Rework | 900.0 | 0.0 | Closure | CLOSED | 27-Apr-2020 | 03-Apr-2020 | 36. | SS | Container
NC-IL081-20798 | 28-Feb-2020 | GRACE.AUSTIN | General | Internal |  |  |  |  | V8870 | NOVAPLUS 18 GA HINGED PHARMA | CFF-D-Molding 1 | Wall Thickness out of specification 0.052 minimum 0.060 | 19L07063 | 1370.0 | Re-hung tool and re-connected water. Ken Wiltse 02/28/2020 | 28-Feb-2020 | 28-Feb-2020 | CFF-D-Molding 1 | Wall Thickness | Water lines not open | Four water lines were not opened after the back half of the mold was clamped up causing no water flow to the circuits. | Rework | 1370.0 | 0.0 | Closure | CLOSED | 28-Apr-2020 | 08-Apr-2020 | 40. | G18 | Container
NC-IL081-20822 | 01-Mar-2020 | MARK.VENTURA | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SPBP-V-Labeling/Print 1 | Label count reconciliation failed | 20B16663X | 229440.0 | Completed end of run, Product placed on hold. Started new lot. P. Myerson 02/28/20 | 01-Mar-2020 | 28-Feb-2020 | SFF-SPBP-V-Labeling/Print 1 | Missing Lot Number /Label | Human Error | Manual labeling is highly affected by personnel and is prone to human error. This is an approved process, but is not utilized as frequently as the preferred in-line labeling system. | Sort | 229440.0 | 0.0 | Closure | CLOSED | 30-Apr-2020 | 27-Mar-2020 | 25. | 8 | Syringe
NC-IL081-20934 | 04-Mar-2020 | GRACE.AUSTIN | General | Internal | C74 | Press 74-IL081 |  |  | 8851 | 8GA PHARMASAFETY WASTE SLIDE | CFF-F-Assembly 2 | Leak Test Failed on lids WP7450BL, Lids were used into Finished Good product 8851. | 19H12063' '19H190037 | 1580.0 1.0 | Corona tray was cleaned and filter for ionizing fan was replaced. Work Order no. CL-3183622. B. Villatoro | 04-Mar-2020 | 04-Mar-2020 | CFF-F-Assembly 2 | Leak Test Failure | Lack of PM | Lack of PM contributed to excessive build up on the corona ionizer fan filter and corona tray cause the gasket to deviate from the corona treatment on the lid. | Scrap | 1580.0 1.0 | 0.0 | Closure | CLOSED | 02-Jun-2020 | 08-May-2020 | 64. | G8 | Container
NC-IL081-20942 | 05-Mar-2020 | ROSEMARY.TEKSTAR | General | Internal | C91 | Press 91-IL081 |  |  | WP7478 | CHEMO CAP DISC ASSM | Documentation Incorrect | Using incorrect shop order at press. Using shop order number 190855, should be using 190856. | 19L190856 | 1.0 | Removed incorrect shop order and replaced with correct shop order. G. Cruz, Jr. | 04-Mar-2020 | 25-Feb-2020 | Documentation Incorrect |  | Human Error | Human Error |  |  | 0.0 | Closure | CLOSED | 03-Apr-2020 | 26-Mar-2020 | 21. | G8 | Container
NC-IL081-20947 | 05-Mar-2020 | JESSE.SOBEL | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SPBP-V-Labeling/Print 1 | unable to reconcile Labels at end of run | 20B25163X | 230400.0 | Completed End of run, removed all labels from line. Product placed on hold. R. Hill 03/05/2020 | 05-Mar-2020 | 04-Mar-2020 | SFF-SPBP-V-Labeling/Print 1 | Missing Lot Number /Label | Method not validated | Lack of process validation | Sort | 230400.0 | 0.0 | Closure | CLOSED | 14-Apr-2020 | 26-Mar-2020 | 21. | 8 | Syringe
NC-IL081-21045 | 10-Mar-2020 | JESSE.SOBEL | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SPBP-V-Labeling/Print 1 | Lot number on cap printing incorrect. Reads "20B03263X" when should be "20C03263X" | 20C03263X | 24480.0 | Cut lot short and product placed on hold. R. Hill 03/10/2020 | 10-Mar-2020 | 10-Mar-2020 | SFF-SPBP-V-Labeling/Print 1 | Missing Lot Number /Label | Human Error | Human Error | Sort | 24480.0 | 0.0 | Closure | CLOSED | 19-Apr-2020 | 01-Apr-2020 | 22. | 3 | Syringe
NC-IL081-21100 | 11-Mar-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | SR-1099, torque limiter calibrated to wrong value and with wrong units. OOT. |  |  |  | 11-Mar-2020 | 11-Mar-2020 | Calibration Out of Tolerance | Calibration Out of Tolerance | Gauge Out of Tolerance |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 158.
NC-IL081-21114 | 12-Mar-2020 | LAKEYSHA.SIMPSON | General | Internal |  |  |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SSF-PPSP-V-LOT NUMBER 2 | Incorrect expiration date on the cap. Reads 2025-03-10 instead of 2025-03-09. | 20C03363X | 70080.0 | Cut the lot short, affected case put on hold and placed in the hold area, cleared out the cap print on packer. L. Reed 03/12/2020 | 11-Mar-2020 | 10-Mar-2020 | SSF-PPSP-V-LOT NUMBER 2 |  | Human Error | Human Error | Sort | 70080.0 | 0.0 | Closure | CLOSED | 20-Apr-2020 | 16-Apr-2020 | 35. | 3 | Syringe
NC-IL081-21188 | 13-Mar-2020 | GRACE.AUSTIN | General | Internal | RSN | INCORRECT RESIN-IL081 |  |  | RM165 | Resin, Greenwood blend, Pre - colored Red | PR-- PR--oduct testing fails | Batch History Record Revealed Failed Impact Testing | 00191840 | 15740.0 |  | 13-Mar-2020 | 13-Mar-2020 | Impact resistance | Failed impact resistance testing at all three heights. | RM164 percentage low | Feed rate of RM164 Kraton 1.5%; well below validated 2.8% feed rate. | Scrap | 15740.0 | 0.0 | Closure | CLOSED | 25-May-2020 | 04-May-2020 | 52. | Recy | Container
NC-IL081-21196 | 14-Mar-2020 | JESSE.SOBEL | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag mark in non-fluid pathway. cavity 37 mold 58110 | 20C04663X' '20C04463X | 134000.0 335000.0 | Cav. 37 blocked, all affected product on hold. D. Breceda | 13-Mar-2020 | 13-Mar-2020 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Damaged gate | During the filling process, the gate pin retracts which allows material to flow around the gate pin into the cavity. Over time material can collect on the gate pin changing the ability for the gate pin to function.) | Sort | 134000.0 335000.0 | 0.0 | Closure | CLOSED | 11-Jun-2020 | 15-Jul-2020 | 123. | Sub | Syringe
NC-IL081-21201 | 14-Mar-2020 | MARQUES.CALIMEE | General | Internal | C54 | Press 54-IL081 |  |  | 8618RC | 18 GA RCRA HAZARDOUS WASTE | CFF-V-Incorrect Items | "The updated pre-kitted labels(with the new 'Hazardous Waste Pharmaceuticals' label) were not used in packaging." | 20B26463 | 11480.0 | Started using the label kit with all labels in the kit. - Guillermo Cruz 3/14/2020 | 14-Mar-2020 | 09-Mar-2020 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | G18 | Container
NC-IL081-21216 | 16-Mar-2020 | MARIA.ROSALES | General | Internal |  |  |  |  | 8881112015 | 12ML SYRINGE PURPLE EN-FIT TIP | Deficient Documentation | Incorrect expiration date on over shipper label not per shop order. Over shipper label reads 2025-03-09 should be 2025-03-12. | 20C05563X | 123.0 | Post production NCR- current lot expiration date was verified to be correct. M. Buck 03/16/2020 | 16-Mar-2020 | 11-Mar-2020 | Deficient Documentation |  | Documentation | Documentation | Rework | 123.0 | 0.0 | Closure | CLOSED | 25-May-2020 | 11-May-2020 | 56. | 8 | Syringe
NC-IL081-21230 | 16-Mar-2020 | PATTI.SWISHER | General | Internal | S8K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Tracking/Trending | Hole,Split or crack in barrel which permits leakage | 20B03963X | 10.0 |  | 16-Mar-2020 | 17-Feb-2020 | SFF-SA-V-Damaged Components 1 | Damaged Components Type 1: Damaged product/components (non-functional) | Machine Failure | Machine Failure | Saleable | 10.0 | 0.0 | Closure | CLOSED | 14-May-2020 | 04-May-2020 | 48. | Sub | Syringe
NC-IL081-21234 | 16-Mar-2020 | PATTI.SWISHER | General | Internal | S8K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Tracking/Trending | The dimension of the barrel end is different than specification. Scar-20-007 | 19F09563X' '19F09663X' '19E01863X' '19H10063X' '19H18963X' '19H19063X' '19H18863X' '19E02263X' '19H19163X | 1.0 |  | 16-Mar-2020 | 20-Jun-2019 | SFF-M-V-Deformation 2 | The dimension of the barrel is different than specification | PR-Not Applicable | Not Applicable-MEDMFGPR | Scrap | 1.0 | 0.0 | Closure | CLOSED | 14-May-2020 | 04-May-2020 | 48. | Sub | Syringe
NC-IL081-21293 | 18-Mar-2020 | GRACE.AUSTIN | General | Internal | C33 | Press 33-IL081 |  |  | PT00016495 | HINGED CHEMO LID, 12 & 18 GAL | Out of Spec Visual Inspection (Visual Defect) | Channel is not continuous, breaks appear in channel. | 20A191353 | 15000.0 |  | 18-Mar-2020 | 18-Mar-2020 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Verification of retention screws | The PM Tooling Technician most likely did not verify the ejector plate retention screws were torqued during the 12/7/19 preventive maintenance service. Over time the retention screws backed out, causing the ejector plates to separate. The polarized ejector pins floated free from their key slots allowing them to rotate during operation, misaligning the track. | Rework | 15000.0 | 0.0 | Closure | CLOSED | 17-May-2020 | 03-Apr-2020 | 16. | G12 | Container
NC-IL081-21294 | 18-Mar-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Damage Gauge | LVS-007, damaged. Cannot log into the PC to calibrate or to use the barcode verifier. |  |  |  | 18-Mar-2020 | 18-Mar-2020 | Calibration Damage Gauge | Calibration Damage Gauge | Damaged Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 151.
NC-IL081-21296 | 18-Mar-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Damage Gauge | LVS-009, Damaged, cannot log into the pc to use or to calibrate the barcode verifier. |  |  |  | 18-Mar-2020 | 18-Mar-2020 | Calibration Damage Gauge | Calibration Damage Gauge | Damaged Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 151.
NC-IL081-21297 | 18-Mar-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0711, Lost 6" ruler. |  |  |  | 18-Mar-2020 | 18-Mar-2020 | Calibration Lost Gauge | Calibration Lost Gauge | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 151.
NC-IL081-21299 | 18-Mar-2020 | FIDELIA.CANTOR | Inspection | Supplier | INC | Incoming-INC-IL081 |  |  | 5000231 | Gusseted Polybag, Unprinted, 13.5 Inch x 14.25 Inch x 5.5 In | Visual- Contamination (not for rubber tips) | On 3/18/20, production informed QA of small black, loose foreign contamination observed on bags of 5000231, lot 19001897019016C. 461 bags remaining in the case were contained. | 19001897019016C | 461.0 | Placed product on hold. | 18-Mar-2020 | 03-Mar-2020 | Visual- Contamination (not for rubber tips) | Visual- Contamination (not for rubber tips) | Failure to Follow Procedures/Policies | Failure to Follow Procedures/Policies. Operator failed to clean draw rollers prior to production run.\n\nThis is the second occurrence of contamination within the last 12 months. This lot was the lot received prior to lot 19002297019042D of 5000231 from NC-IL081-20425, and after speaking with the supplier the root cause for this issue is typically due to the operator failed to clean draw rollers prior to production run. | Scrap | 461.0 | 0.0 | Closure | CLOSED | 17-Apr-2020 | 26-Mar-2020 | 7.
NC-IL081-21312 | 19-Mar-2020 | LAKEYSHA.SIMPSON | General | Internal | 131 | Cell 131-IL081 |  |  | 8506SA | CONTAINER BIOHAZARD 5 QT | CFF-Welding | Failed weld strength inspection on lids WP8506 , lids was used into Finish Good 8506SA. | 19M03863 | 1920.0 | Increase the air and amplitude to the welders to create stronger welds. Michael Buck 03/20/2020 | 19-Mar-2020 | 19-Mar-2020 | CFF-Welding |  | Machine Failure | Machine Failure | Rework | 1920.0 | 0.0 | Closure | CLOSED | 18-Apr-2020 | 10-Apr-2020 | 22. | SS | Container
NC-IL081-21421 | 23-Mar-2020 | BERTHA.HERNANDEZ | General | Internal | S6A | 6A-IL081 | Equipment ID | 58110-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Deformation 1 | Pulled ring stop and gate vestige, found during 1st article inspection of Lot#20C09963X, no F/G packed and all the WIP was scrapped at the line. Pulled ring stop cavities 10,28,32,42. gate vestige cavities 24,37338,39,48 | 20C09863X | 67000.0 | Cleared line of all barrels and sub assemblies. Slowed down the cycle time. P. Myerson 03/23/20 | 23-Mar-2020 | 23-Mar-2020 | SFF-M-V-Deformation 1 | The ring stop has been pulled from the inside dimension of the barrel possibly rendering the barrel unusable | Blocked Bubbler |  | Scrap | 67000.0 | 0.0 | Closure | CLOSED | 23-May-2020 | 11-May-2020 | 48. | Sub | Syringe
NC-IL081-21433 | 24-Mar-2020 | JESSE.SOBEL | General | Internal | S3C | 3C-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Scratch on sleeve can be felt with a bare finger. Cavity# 58. . | 20C08963X | 36960.0 | Scrapped contaminated product from line before starting new lot. R. Hill 03/23/2020 | 24-Mar-2020 | 23-Mar-2020 |  |  |  |  | Scrap |  | 0.0 | Investigation | VOID |  |  |  | 3 | Syringe
NC-IL081-21448 | 24-Mar-2020 | GRACE.AUSTIN | General | Internal | C56 | Press 56-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | CFF-V-Label Verification and Barcode Verification 2 | Incomplete Print on Biohazard Label. Missing some print on Product code and Lot number. | 20B09063 | 330.0 | Associate changed label roll. Ken Wiltse, 03/24/20 | 24-Mar-2020 | 24-Mar-2020 | CFF-V-Label Verification and Barcode Verification 2 |  | Wrinkled Label |  | Sort | 330.0 | 0.0 | Closure | CLOSED | 23-May-2020 | 15-Apr-2020 | 22. | G8 | Container
NC-IL081-21486 | 25-Mar-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | 054; Scale was out of tolerance as found during normally scheduled calibration activity on 03/19/2020. Calibration passed after adjustment. |  |  |  | 25-Mar-2020 | 19-Mar-2020 | Calibration Out of Tolerance | Calibration Out of Tolerance | Out of Tolerance |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 144.
NC-IL081-21488 | 25-Mar-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | 053; Scale out of tolerance as found during normally scheduled calibration activity on 03/19/2020. Scale passed calibration after adjustment. |  |  |  | 25-Mar-2020 | 19-Mar-2020 | Calibration Out of Tolerance | Calibration Out of Tolerance | Out of Tolerance |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 144.
NC-IL081-21492 | 25-Mar-2020 | FIDELIA.CANTOR | General | Internal | C61 | Press 61-IL081 |  |  | WP7417P | 8618RC LID | PR-- Missing PR--int | During PQ 3 run, production found lids with missing dots on one side of lids for product code WP7417P. not per standard work DC00034290. | 20C192075 | 175.0 | Program primer to print dots. Justin Pixius. 03/30/20 | 25-Mar-2020 | 25-Mar-2020 | PR- Missing Print | Missing print-MEDMFGPR | Process Not Robust | Process Not Robust | Scrap | 175.0 | 0.0 | Closure | CLOSED | 04-Jul-2020 | 29-May-2020 | 64. | G18 | Container
NC-IL081-21512 | 26-Mar-2020 | GRACE.AUSTIN | General | Internal | C89 | Press 89-IL081 |  |  | 31353603 | 5 QT GATORGUARD JR - TRANS RED | CFF-V-Packaging 1 | Packed Incorrect quantity, missing 2 ea. containers per box, should be packed with 14 ea of containers per Standard Work and BOM | 20A04963 | 238.0 | Retrained associate on packaging form FM20008551. Berenice Villatoro. 3/26/20 | 26-Mar-2020 | 26-Mar-2020 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | The Manufacturing Associate failed to follow Standard Work for packaging FM20008551 section 2 which was available at the press workstation when this incident occurred. Human Error. | Rework | 238.0 | 0.0 | Closure | CLOSED | 25-May-2020 | 06-May-2020 | 41. | GG | Container
NC-IL081-21516 | 26-Mar-2020 | GRACE.AUSTIN | General | Internal | CME | CME-IL081 |  |  |  |  | Documentation Missing | Failure to follow line clearance and Start procedure. GS10000122. Documentation missing No product impacted. Found during run of 4001204, Lot# TSTB031920. |  |  | N/A as the NCR is non-product. | 26-Mar-2020 | 26-Mar-2020 | Procedure Not Followed | Procedure Not Followed | Procedure not followed | Procedure not followed |  |  | 0.0 | Closure | CLOSED | 20-May-2020 | 05-May-2020 | 40.
NC-IL081-21525 | 26-Mar-2020 | BERTHA.HERNANDEZ | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Deformation 1 | Pulled ring stop on cavity 42 mold 58110. | 20C11063X' '20C10963X | 9000.0 191000.0 | Clear line, cavity blocked. D. Breceda 03/26/2020 | 26-Mar-2020 | 26-Mar-2020 | SFF-M-V-Deformation 1 | The ring stop has been pulled from the inside dimension of the barrel possibly rendering the barrel unusable | Shot Size |  | Scrap' 'Sort | 9000.0 191000.0 | 0.0 | Closure | CLOSED | 25-May-2020 | 29-Sep-2020 | 187. | Sub | Syringe
NC-IL081-21555 | 27-Mar-2020 | GRACE.AUSTIN | General | Internal | C61 | Press 61-IL081 |  |  | WP7417P | Slide Top, Foam In Place Lid, 12 and 18 Gallon Container, PG | CFF-D-Molding 3 | Failed Gasket Height inspection during PQ3 Run. | 20C192075 | 175.0 | Performed Line Clearance to that Engineering could trouble shoot. Justin Pixius. 04/07/20 | 27-Mar-2020 | 27-Mar-2020 | CFF-F-Molding 8 | Failed gasket height/width | Process Not Robust | Process Not Robust | Scrap | 175.0 | 0.0 | Closure | CLOSED | 04-Jul-2020 | 29-May-2020 | 62. | G18 | Container
NC-IL081-21569 | 30-Mar-2020 | LAKEYSHA.SIMPSON | General | Internal |  |  |  |  | 8989PG2 | SHARPS CONTAINER 18GAL | CFF-V-Assembly 2 | Using PT00013694 lids from RAMPF into Finished Good 8989PG2 that are not released from incoming. Lids was used before inspection completed. | 20B28663 | 655.0 | Did End of Run and Started new Job with different lids. Steve Schirmer 04/07/20 | 29-Mar-2020 | 29-Mar-2020 | CFF-V-Assembly 2 | Disc Assembly Inspection | SOP |  | Other/NDF | 655.0 | 0.0 | Closure | CLOSED | 28-May-2020 | 08-May-2020 | 39. | G18 | Container
NC-IL081-21609 | 31-Mar-2020 | PATTI.SWISHER | General | Internal | S7C | 7C-IL081 |  |  | 8881907003 | 10ML ORAL MED SYRINGE (AMBER) | SFF-M-F-Functional | Failed Light transmission Testing LS10000773 | 20B15663X | 143000.0 |  | 31-Mar-2020 | 24-Feb-2020 | SFF-M-F-Functional | Functional Leak Test | Unclear Procedure | Process control procedure (PS20003316 Amber Barrel Process Control) is in place for startup/shut down and troubleshooting. The verbiage used is unclear for the process set up. | Scrap | 143000.0 | 0.0 | Closure | CLOSED | 30-May-2020 | 27-May-2020 | 57. | 7 | Syringe
NC-IL081-21624 | 01-Apr-2020 | MARIA.ROSALES | General | Internal | S8C | 8C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Burns/Gas Marks | Burn marks on the plunger cavity 23D Mold 58120/8C | 20C14763X | 13000.0 | Cleared line of all plungers and sub assemblies. adjust shot size from 2.80-2.55 and injection velocity from 1.70-1.00. P. Myerson 03/31/2020 | 31-Mar-2020 | 31-Mar-2020 | SFF-M-V-Burns/Gas Marks | Gas or Burn Marks on the part that renders it cosmetically unacceptable | Shot size | Shot size was not adjusted enough during blocking of mold cavities. | Scrap | 13000.0 | 0.0 | Closure | CLOSED | 31-May-2020 | 15-Apr-2020 | 14. | Sub | Syringe
NC-IL081-21642 | 01-Apr-2020 | BERTHA.HERNANDEZ | General | Internal | S2D | 2D-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-M-V-Cracking 1 | Hole allowed leakage or loss of sterility on barrel cavity# 47B mold 58115/2D. | 20C13463X | 45600.0 | Cleared line of all barrels and sub assemblies. Blocked cavity 47B. P. Myerson | 01-Apr-2020 | 31-Mar-2020 | SFF-M-V-Cracking 1 | Cracking Type 1: No Crack allowed: leakage or loss of sterility as risk | Insufficient fill | Insufficient material caused by the plastic setting up before the end of part causing a short. This is most likely due to the gate pin being stuck in the forward position | Scrap | 45600.0 | 0.0 | Closure | CLOSED | 30-Jun-2020 | 14-May-2020 | 42. | 3 | Syringe
NC-IL081-21725 | 05-Apr-2020 | BERTHA.HERNANDEZ | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Damaged Components 3 | Damaged component (plunger) functional but less than desired. | 20C19863X | 1000.0 | Cleared line of all sub assemblies. Scrapped affected case. Removed caught parts from the kahle. P. Myerson 04/05/2020 | 05-Apr-2020 | 05-Apr-2020 | SFF-SA-V-Damaged Components 3 | Damaged Components Type 3: Damaged product/component (functional but less than desired); Scratches/drag marks on the barrel | Caught part | There was a caught part in the kahle dial system which was striking the plungers as they moved around the dials. | Scrap | 1000.0 | 0.0 | Closure | CLOSED | 05-Jun-2020 | 01-Jun-2020 | 56. | Sub | Syringe
NC-IL081-21763 | 07-Apr-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Damage Gauge | MD-645, damaged un-repairable. Received for inspection in the calibration lab. Discovered damaged load cell, and determined to be un-repairable. Unit would power up, but would only display 0.0 when force was applied (both in tension and compression). |  |  |  | 07-Apr-2020 | 07-Apr-2020 | Calibration Damage Gauge | Calibration Damage Gauge | Damaged Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 131.
NC-IL081-21824 | 09-Apr-2020 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | PT00013692 | Gasketed Lid, Single Pass, 8 Gallon | Incorrect/Inadequate Labeling | During incoming inspection on 4/09/20, for product code PT00013692 from Rampf, 41 labels with the incorrect PO # 822877 and 3 labels with incorrect PO# 822830 and incorrect lot #165 were identified. Correct PO # is 822985 and correct lot # 171 per Certificate of Compliance. | 171 | 8820.0 |  | 09-Apr-2020 | 09-Apr-2020 | PR-Labeling Incorrect | Labeling Incorrect-MEDMFGPR | Inadequate Training | Inadequate Training | Saleable | 8820.0 | 0.0 | Closure | CLOSED | 09-May-2020 | 14-Apr-2020 | 5. | G8 | Container
NC-IL081-21842 | 10-Apr-2020 | BERTHA.HERNANDEZ | General | Internal | S6P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | PR-Other | Failed Endotoxin specification of <0.3 EU/Device per GS20003250. | 20C27063X | 335000.0 | Contained lot, re-pull samples again. | 10-Apr-2020 | 05-Apr-2020 | Out of Spec Chemical/Biological (Lab Testing) | Out of Spec Chemical/Biological (Lab Testing) | Handling/Storage | Handling/Storage | UseAsIs | 335000.0 | 0.0 | Investigation Approval | VOID | 21-May-2020 |  |  | Sub | Syringe
NC-IL081-21848 | 13-Apr-2020 | LAKEYSHA.SIMPSON | General | Internal | C19 | Press 19-IL081 |  |  | 8506SA | CONTAINER BIOHAZARD 5 QT | CFF-F-Welding | Failed weld strength inspection on Lids WP8506. Lids was used and packed into Finish Good | 20B07863 | 2280.0 | Increase the amperage on the welder. Michael Buck. 04/13/20 | 12-Apr-2020 | 12-Apr-2020 | CFF-F-Welding | Weld Strength | Machine Failure | Machine Failure | Accept | 2280.0 | 0.0 | Closure | CLOSED | 12-Jul-2020 | 05-Jun-2020 | 53. | SS | Container
NC-IL081-21854 | 13-Apr-2020 | PATTI.SWISHER | General | Internal | S8K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Damaged Components 1 | Damaged product/components Non Functional | 20A22163X | 19.0 |  | 10-Apr-2020 | 06-Feb-2020 | SFF-SA-V-Damaged Components 1 | Damaged Components Type 1: Damaged product/components (non-functional) | Machine Failure | Machine Failure | Saleable | 19.0 | 0.0 | Closure | CLOSED | 11-Jun-2020 | 04-May-2020 | 21. | Sub | Syringe
NC-IL081-21871 | 14-Apr-2020 | LAKEYSHA.SIMPSON | General | Internal | C85 | Press 85-IL081 |  |  | WP3204 | Snap Cap and Lid Assembly, Biohazard | Documentation Missing | Failed to perform Auxiliary PCR readings at Start-up. Auxiliary PCR was not print out. | 20B191665 | 1.0 | Auxiliary PCR was printed out and place by the Press. All parameters were within the limits. Oleg Ratinskiy. 04/16/20 | 14-Apr-2020 | 13-Apr-2020 | Documentation Missing | Documentation Missing | Incomplete completion of documentation | Incomplete completion of documentation | Other/NDF | 1.0 | 0.0 | Closure | CLOSED | 13-Jul-2020 | 26-Jun-2020 | 73. | G2 | Container
NC-IL081-21909 | 15-Apr-2020 | MARQUES.CALIMEE | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 1 | Hard flash on functional area exceeding 1/64th". Mold 58114/6C cavity 37B & 45B | 20D06863X | 120000.0 | Stopped line, cleared hopper, kahle and disposed of WIP from on 6C. Performed lot change. Restarted the line. L. Reed 04/15/2020 | 15-Apr-2020 | 14-Apr-2020 | SFF-M-V-Flash/Gate 1 | Flash Type 1: Flash in Fluid Path (Luer Tip) exceeds 1/64” (.016”). Does not meets drawing specification | Mold Wear or Damage | Mold Wear or Damage | Accept | 120000.0 | 0.0 | Closure | CLOSED | 15-Jun-2020 | 17-Jun-2020 | 63. | Sub | Syringe
NC-IL081-21944 | 16-Apr-2020 | ALEJANDRA.CASTANEDA | General | Internal | C64 | Press 64-IL081 |  |  | WP8850 | 8850 LID | CFF-F-Molding 8 | Corona Failure- Gasket not adhered to the lids. | 20C192148 | 1716.0 | Reattached corona power wire to ensure gasket adhesion. Guillermo Cruz. 04/15/20. | 15-Apr-2020 | 15-Apr-2020 | CFF-F-Molding 8 | Failed gasket height/width | Machine Failure | Machine Failure | Accept | 1716.0 | 0.0 | Closure | CLOSED | 31-Jul-2020 | 09-Jul-2020 | 84. | G8 | Container
NC-IL081-22011 | 18-Apr-2020 | BERTHA.HERNANDEZ | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | NLC NO MOVE (020814)SYR,MONOJECT SOFTPACK,12CC,LL, | SFF-PPSP-V-Seal Integrity 2 | Seal integrity, seal has a void doesn\'t meet the 3MM (1/8") width. | 20D07963X | 12960.0 | Cleared line of all packaging, changed paper roll, scrapped 27 cases. P. Myerson 04/18/2020 | 18-Apr-2020 | 18-Apr-2020 | SFF-PPSP-V-Seal Integrity 2 | Seal is intact but less than 3 mm (1/8 inch) width seal. For 60 mL and 35 mL products, less than 5 mm width. | Film mis-aligned | The film had moved to one side on the Multivac and folded over which caused a double layer of material. This caused it not to seal properly. | Scrap | 12960.0 | 1296.0 | Closure | CLOSED | 18-Jun-2020 | 01-Jun-2020 | 44. | 8 | Syringe
NC-IL081-22016 | 19-Apr-2020 | BERTHA.HERNANDEZ | General | Internal | S7R | RAM-IL081 |  |  | 6492743 | 12ML LUER LOCK SYRINGE, PRINTED | SFF-SA-V-Mixed Product 2 | Incomplete assembly of product , upside rubber tip. | 20D07363X | 800.0 | Cleared Line, scrapped box 116, replaced cables on sensors. P. Myerson 04/19/2020 | 19-Apr-2020 | 19-Apr-2020 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Ram improper operation | In-line vibratory rail not vibrating due to wires from on/off switch not making contact at terminal strip. Wires at quick connect at the vibratory block barely making contact. \nHigh plunger sensor not working properly\nPlunger inspect station not detecting defect | Scrap | 800.0 | 0.0 | Investigation | VOID | 19-Jun-2020 |  |  | 7 | Syringe
NC-IL081-22017 | 19-Apr-2020 | BERTHA.HERNANDEZ | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Shorts 1 | Shorts in fluid pathway of cavity 9. of mold 58110/MP-1134/6A | 20D13863X | 14400.0 | All affected product scrapped, Cavity blocked. P. Myerson | 19-Apr-2020 | 18-Apr-2020 | SFF-M-V-Shorts 1 | Shorts in fluid pathway or in area where Sterility is required | Blockage in gate pin | Cavity 9 was shorting due to a blockage in the gate pin. Blocked cavity. | Reject | 14400.0 | 0.0 | Closure | CLOSED | 19-Jun-2020 | 01-Jun-2020 | 43. | Sub | Syringe
NC-IL081-22019 | 20-Apr-2020 | MARQUES.CALIMEE | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Drag Marks/ Gouges/Nicks Scratches 2 | Drag marks on outside of barrel of cavity 18 B that can be felt with finger. On the first article. 6C/58114 | 20D14163X' '20D14063X | 92000.0 70000.0 | Tech cleared the gate and we cleared the line. F Bahena 04/20/2020 | 20-Apr-2020 | 19-Apr-2020 | SFF-M-V-Drag Marks/ Gouges/Nicks Scratches 2 |  | Mold Wear or Damage | Mold Wear or Damage | Accept | 70000.0 92000.0 | 0.0 | Closure | CLOSED | 08-Jun-2020 | 23-Jun-2020 | 64. | Sub | Syringe
NC-IL081-22044 | 20-Apr-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Damage Gauge | "LVS-011, Damaged, needs repair. During normally scheduled calibration activities, unit was found with the LED's level of illumination insufficient to capture an image. The unit was unable to yield any results." |  |  |  | 20-Apr-2020 | 17-Apr-2020 | Calibration Damage Gauge | Calibration Damage Gauge | Damaged Gauge |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 118.
NC-IL081-22086 | 21-Apr-2020 | MARK.VENTURA | General | Internal | S8K | Kahle-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SA-V-Particulate 2 | Silicone Group Average on heads 2 and 10 exceeds the maximum allowable silicone limit specification by weight | 20D15863X | 230400.0 | Place product on QA Hold and pull samples from BEG, MID and END to send out for extraction. | 21-Apr-2020 | 20-Apr-2020 | SFF-SA-F-Spraymation | Spraymation | Process Complexity |  | Reject | 230400.0 | 0.0 | Closure | CLOSED | 31-Aug-2020 | 26-Aug-2020 | 126. | 8 | Syringe
NC-IL081-22107 | 21-Apr-2020 | PATTI.SWISHER | General | Internal | S8K | Kahle-IL081 |  |  | 1181200777K | 10ML STANDARD LUER LOCK SYRING | SFF-SA-V-Particulate 2 | Excessive silicone on head 5 during weekly spraymation testing | 20D07663X | 230400.0 | Place product on QA Hold. Pull samples from BEG, MID and END send out for extraction testing. | 21-Apr-2020 | 15-Apr-2020 | SFF-SA-F-Spraymation | Spraymation | Process Complexity |  | Reject | 230400.0 | 0.0 | Closure | CLOSED | 18-Sep-2020 | 26-Aug-2020 | 126.
NC-IL081-22125 | 22-Apr-2020 | GRACE.AUSTIN | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | CFF-V-Label Verification and Barcode Verification 10 | Missing Biohazard Label on one side of the containers. Per Standard Work required label both side of the containers. | 20B18263 | 720.0 | Isolated Questionable Product, Refereed to Standard Work and Labeled per Standard Work. Steve Schrimer, 04/22/20 | 22-Apr-2020 | 22-Apr-2020 | CFF-V-Label Verification and Barcode Verification 10 | BH - Missing label | Human Error | Human Error | Accept | 720.0 | 0.0 | Closure | CLOSED | 30-Jun-2020 | 17-Jun-2020 | 56. | G2 | Container
NC-IL081-22172 | 23-Apr-2020 | GRACE.AUSTIN | General | Internal | C42 | Press 42-IL081 |  |  | PT00016490 | 12/18 GAL. HINGED CHEMO LID | CFF-V-Label Verification and Barcode Verification 13 | Non Gasket lids are sent to Rampf for gasket application and then return to site for consumption into Finished Good product. During processing of PO# 823131, Rampf found additional labels on the corrugate sent by Crystal lake which contain the non-gasketed lids. | 20C192382 | 15000.0 | Product was contained at Rampf with a memo provided to request the removal of all additional labels found. J. Grisko 4/23/2020 | 23-Apr-2020 | 23-Apr-2020 | CFF-V-Label Verification and Barcode Verification 13 | OS - Wrong label | Human error | Human error. There is no way to trace back to human that caused issue. | Saleable | 15000.0 | 0.0 | Closure | CLOSED | 23-May-2020 | 01-May-2020 | 8. | G12 | Container
NC-IL081-22215 | 27-Apr-2020 | GRACE.AUSTIN | General | Internal | C42 | Press 42-IL081 |  |  | 85221R | 3GA ALWAYS OPEN INROOM -T RED | CFF-V-Label Verification and Barcode Verification 10 | Missing Biohazard label on one side of the containers, Required label both side per standard work FM20008357. | 20D22063 | 260.0 | Reload label program to the printer. Ken Wiltse. 4/29/20 | 27-Apr-2020 | 27-Apr-2020 | CFF-V-Label Verification and Barcode Verification 10 | BH - Missing label | Machine Failure | Machine Failure | Accept | 260.0 | 0.0 | Closure | CLOSED | 31-Jul-2020 | 26-Jun-2020 | 60. | G3 | Container
NC-IL081-22238 | 27-Apr-2020 | ANGELICA.AGUILAR | General | Internal | C56 | Press 56-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | Unapproved Material | Un-tested RM165 from day silo 1 was released into the approved resin silo #5. Press 56 was consuming RM165 at time of release. | 20C02063' '00192601 | 1080.0 11885.0 | 1. The valve on day silo 1 was closed.\n2. Manufacturing performed an end of run on the production job at press 56 and switched to a job with a different material. | 27-Apr-2020 | 27-Apr-2020 | Unapproved Material | Unapproved Material | Air valve malfunction | Air valve malfunction | Saleable | 1080.0 11885.0 | 0.0 | Closure | CLOSED | 27-May-2020 | 01-May-2020 | 3. | G8 | Container
NC-IL081-22260 | 28-Apr-2020 | PATTI.SWISHER | General | Internal | S4R | RAM-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | Validation - Process Not to Protocol Requirements | This nonconformance serves to document the occurrence of a deviation to the validated tip/syringe assembly process (Agile validation XXX). All lots impacted by this deviation from process will be recorded in this non-conformance and their disposition will be vetted by its disposition.\n\nCrystal Lake site Operations, Engineering, and QA personnel were involved in the decision to deviate, as recorded here. | 20C19063X | 1.0 | "Immediate Actions:\n1. Review causing event (mechanical failure of line XX tip hopper (complete)\n2. Review of options to return line to service (complete)\n3. Completion of product impact review of the alternate process engaged (complete, attached)\n\nDisposition of the lots listed in this nonconformance shall include acceptable review of:\n1. The attached, signed product impact statement\n2. A review of the completed work order to repair the tip hopper (investigation section of this nonconformance may refer to cause of malfunction listed in work order)\n3. Acceptable results for all testing requirements for affected lots\n4. Acceptable completion of all standard product release requirements\n-----\n\nTech's made an adjustment to the rubber tip bowel sensor. P. Myerson 04/28/2020" | 28-Apr-2020 | 28-Apr-2020 | Non-Compliant w Procedure | Non-Compliant w Procedure | Sensor | Rubber tip bowl sensor not working to advance tips to bowl. | Accept | 1.0 | 0.0 | Closure | CLOSED | 28-Jun-2020 | 10-Jun-2020 | 43. | 412 | Syringe
NC-IL081-22265 | 28-Apr-2020 | BERTHA.HERNANDEZ | General | Internal | S8P | Packer-IL081 |  |  |  |  | PCRINC | Failed to performed PCR inspection FM20007657 once per shift per GS10004646. Found during production of product 8881540122, Lot# 20D16163X. NCR will be tracked via BPCS hold number 20040050. |  |  | No product impact. Line has been down since 0800 04/27/20 | 28-Apr-2020 | 27-Apr-2020 | PCRINC |  | Human Error | Human Error |  |  | 0.0 | Closure | CLOSED | 28-Jun-2020 | 20-May-2020 | 22.
NC-IL081-22331 | 30-Apr-2020 | MARK.VENTURA | General | Internal | S8K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-F-Spraymation | silicone group average on heads 9 and 12 exceed maximum allowable silicone limit specification by weight | 20D18463X | 132000.0 | Line purged and changed heads. K. Kaecker 05/05/2020 | 30-Apr-2020 | 30-Apr-2020 | SFF-SA-F-Spraymation | Spraymation | Process Complexity |  | Accept | 132000.0 | 0.0 | Closure | CLOSED | 31-Aug-2020 | 03-Aug-2020 | 94. | Sub | Syringe
NC-IL081-22340 | 01-May-2020 | MARK.VENTURA | General | Internal | Tumblers | Tumblers-IL081 |  |  |  |  | Missing Labeling | Hold labels not applied on product from non standard shop order.\n\nTracking/trending of NCR will be maintained via BPCS hold numbers 20050011, 20050012, and 20050013. DHRs for products 4006202 (SO# A000175043), 4101801 (SO# 20C13-8), and 4001204 (SO# YL150919-011) were used to generate the BPCS hold numbers. |  |  |  | 01-May-2020 | 22-Apr-2020 | Missing Labeling | Missing Labeling | Non-Compliant w Procedure | Non-Compliant w Procedure: This is a one off event caused by Human Error. The human was distracted and running on auto pilot. |  |  | 0.0 | Closure | CLOSED | 30-Jun-2020 | 19-Jun-2020 | 48.
NC-IL081-22383 | 05-May-2020 | FIDELIA.CANTOR | Inspection | Internal | C33 | Press 33-IL081 |  |  | PT00013694 | Gasketed Lid, Hinged, 12 and 18 Gallon Container, PGII | CFF-V-Broken, Cracking, and Crazing | During Incoming inspection on 5/4/20 the Leak test failed due to cracked hinge on the lid (R7515A/PT00016495 ) this product code was produced in house on press 33 lot # 20C192158. | 176 | 90.0 | N/A - Post Production | 04-May-2020 | 04-May-2020 | CFF-V-Broken, Cracking, and Crazing | Broken, Cracking, and Crazing all surfaces | Failure to Follow Procedures/Policies | Failure to Follow Procedures/Policies | Reject | 90.0 | 0.0 | Closure | CLOSED | 31-Aug-2020 | 03-Aug-2020 | 90. | IQA | Incoming
NC-IL081-22394 | 05-May-2020 | ANGELICA.AGUILAR | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-PPSP-V-Labeling/Print 2 | Lot number, illegible print on cap | 20D22963X | 23040.0 | Changed print cartridge to new one. K. Kaecker 05/05/2020 | 05-May-2020 | 05-May-2020 | SFF-PPSP-V-Labeling/Print 2 | Printing is smeared and not legible | Ink Cartridge | The head of the ink cartridge had a buildup of ink causing the ink not to flow properly. | Reject' 'Accept | 19143.0 3897.0 | 0.0 | Closure | CLOSED | 03-Aug-2020 | 13-Jul-2020 | 69. | 3 | Syringe
NC-IL081-22451 | 07-May-2020 | LAKEYSHA.SIMPSON | General | Internal | S1B | 1B-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-M-V-Flash/Gate | "Flash in fluid pathway exceeds 1/64'' cavity 41, Mold 58140-01/ MP1126" | 20D26963X | 3200.0 | Line was stopped,, cleared all product from the line. Created W.O.# CL-3204774 to correct cavity defect. Blocked cavity 41, Cut the lot short. L. Reed 05/06/2020 | 07-May-2020 | 06-May-2020 | SFF-M-V-Flash/Gate 1 | Flash Type 1: Flash in Fluid Path (Luer Tip) exceeds 1/64” (.016”). Does not meets drawing specification | Mold Wear or Damage | Mold Wear or Damage | Reject | 3200.0 | 0.0 | Closure | CLOSED | 30-Jun-2020 | 04-Jun-2020 | 28. | 1 | Syringe
NC-IL081-22476 | 07-May-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Damage Gauge | "WSV-004; Unit won't power-up. Failure occured shortly after calibration. Failure mode precludes the delivery of incorrect results." |  |  |  | 07-May-2020 | 07-May-2020 | Calibration Damage Gauge | Calibration Damage Gauge | DAMAGED GAUGE |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 14-Aug-2020 | 99.
NC-IL081-22517 | 09-May-2020 | MARK.VENTURA | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Shorts 1 | Shorts in fluid pathway or in area where sterility is required. Cavity 27B. Shorts on shoulder. Mold 58114, 6C, MP-1136 | 20D28063X | 1000.0 | Cleared line of parts to scrap and then Tech increased Temp to correct short. K. Kaecker 05/14/2020 | 09-May-2020 | 09-May-2020 | SFF-M-V-Shorts 1 | Shorts in fluid pathway or in area where Sterility is required | Blockage | Blockage / obstruction in the gate pin | Reject | 1000.0 | 0.0 | Closure | CLOSED | 07-Aug-2020 | 06-Jul-2020 | 57. | Sub | Syringe
NC-IL081-22520 | 11-May-2020 | MARIA.ROSALES | General | Internal | CME | CME-IL081 |  |  |  |  | Calibration Expired | Luer gage calibration due date was expired. Found during manufacturing of product 8881112083, Lot# 20D27263X. | 20D27263X | 1.0 | N/A for non-product | 10-May-2020 | 10-May-2020 | PR-Calibration/Maintenance | Calibration/Maintenance-MEDMFGPR | Documentation | Documentation |  |  | 0.0 | Closure | CLOSED | 07-Jan-2021 | 18-Aug-2020 | 99.
NC-IL081-22552 | 12-May-2020 | BERTHA.HERNANDEZ | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Contamination/Specks 1 | Contamination on barrel, mold 58114. | 20D28663X | 1000.0 | Scrapped case 93 and cleared line. P. Myerson 05/12/2020 | 12-May-2020 | 12-May-2020 | SFF-M-V-Contamination/Specks 1 | Contaminations/Specks Type 1: Inclusions within the molded plastic Barrel luer, fluid pathway or ability to affect sterility - no defects allowed regardless of size or qty. | Machine Failure | Machine Failure | Reject | 1000.0 | 0.0 | Closure | CLOSED | 21-Jun-2020 | 19-Jun-2020 | 38. | Sub | Syringe
NC-IL081-22678 | 16-May-2020 | MARIA.ROSALES | General | Internal | S2B | 2B-IL081 | Equipment ID | 58019-IL081 | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-M-V-Flash/Gate3 | Flash in non-functional area exceeds 1/16" Mold 58019, MP-1115, RE001170443 Validation | 20E05663X | 82080.0 | No Correction at the time mold 58019 not in use. L. Reed 05/16/2020 | 15-May-2020 | 15-May-2020 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Process/Method | Process/Method | Reject | 82080.0 | 0.0 | Closure | CLOSED | 31-Aug-2020 | 14-Aug-2020 | 90. | 3 | Syringe
NC-IL081-22679 | 16-May-2020 | LAKEYSHA.SIMPSON | General | Internal | C33 | Press 33-IL081 |  |  | R7515P | HINGED CHEMO LID, 12 & 18 GAL | CFF-V-Packaging 1 | Product is not packaged per approved packing standards. | 20E192983 | 2335.0 | Update to formula sheet for alternate packing in rout. Micheal Buck. 05/20/20 | 16-May-2020 | 16-May-2020 | CFF-V-Packaging 1 | Product is not packaged as per approved packaging standards. | Packaging/Cartoning | Production made the call to change the packaging due to the current packaging causing bends in the lid that would not run through our gasketing process. | Accept | 2335.0 | 0.0 | Closure | CLOSED | 15-Jun-2020 | 21-May-2020 | 5.
NC-IL081-22688 | 16-May-2020 | BERTHA.HERNANDEZ | General | Internal | C89 | Press 89-IL081 |  |  | 31353603 | 5 QT GATORGUARD JR - TRANS RED | CFF-V-Labeling - Product & Overshipper 1 | Missing lot number on biohazard label. Found during the QA Final Case Inspection | 20C07063 | 1596.0 | Cleaned up printer head and roller changed out the ribbon. Bernice Villatoro, 05/16/20 | 16-May-2020 | 16-May-2020 | CFF-V-Labeling - Product & Overshipper 1 | Product and Lot Number is incomplete or missing | Uneven platen roller | The uneven platen roller of the printer caused inconsistent print pressure. | Accept | 1596.0 | 0.0 | Closure | CLOSED | 15-Jul-2020 | 22-Jun-2020 | 36. | GG | Container
NC-IL081-22689 | 16-May-2020 | BERTHA.HERNANDEZ | General | Internal | S2N | Hot Stamp Printer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-P-V-Numbering 1 | Missing one or more numerals and alpha on barrel print. | 20E03663X | 2400.0 | Cleared line of all printed barrels and sub assemblies. Scrapped 5 affected case. Tech adjusted the MK anilex roller for MK to transfer to the cliche. P. Myerson 05/16/2020 | 16-May-2020 | 16-May-2020 | SFF-P-V-Numbering 1 | Type 1: Missing 1 or more numeral(s) | In-adequate ink transfer | Adjustments made to the gap on the anilox roller was too much, causing ink to not transfer to the cliché. | Reject | 2400.0 | 0.0 | Closure | CLOSED | 16-Jul-2020 | 24-Jun-2020 | 38. | 3 | Syringe
NC-IL081-22750 | 20-May-2020 | ANGELICA.AGUILAR | General | Internal | S8C | 8C-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | PR-Parameters out of specification | Press MP-1120 (8C/58120-02) out of parameters. Fill time 0.97 and should be 1.00-1.10 | 20D22163X | 201600.0 | Press brought back into parameters and parts cleared from line. K. Kaecker 05/20/2020 | 20-May-2020 | 20-May-2020 | PR-Parameters out of specification | Parameters out of specification-MEDMFGPR | People/Human Error/ | People/Human Error/ | Accept | 201600.0 | 0.0 | Closure | CLOSED | 29-Jun-2020 | 20-Aug-2020 | 92. | 8 | Syringe
NC-IL081-22804 | 22-May-2020 | BERTHA.HERNANDEZ | General | Internal | S8F | Form, Fill Sealer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-PPSP-V-Seal Integrity 1 | Void allows air into sealed package cavity 18. | 20D22463X | 11040.0 | Replaced gasket, Work Order CL-3208356. Product reworked. P. Myerson 05/22/2020 | 22-May-2020 | 22-May-2020 | SFF-PPSP-V-Seal Integrity 1 | Void allows air into sealed package | Damaged Gasket | Damaged Gasket | Accept | 11040.0 | 0.0 | Closure | CLOSED | 01-Jul-2020 | 24-Jun-2020 | 33. | 8 | Syringe
NC-IL081-22817 | 24-May-2020 | MARK.VENTURA | General | Internal | CME | CME-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | PR-Environmental | Leak coming from ceiling of the CME packing area. | 20D22463X | 1.0 | area cleaned, bucket placed under leak. M. Ventura 05/24/2020 | 24-May-2020 | 24-May-2020 | Environmental Monitoring | Environmental Monitoring | Equipment | Equipment | Accept | 1.0 | 0.0 | Closure | CLOSED | 16-Sep-2020 | 17-Jul-2020 | 54. | 8 | Syringe
NC-IL081-22842 | 26-May-2020 | BERTHA.HERNANDEZ | General | Internal | S3D | 3D-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-M-V-Flash/Gate 3 | Flash in non-functional area exceeds 1/16" on cap component. | 20E07363X | 15360.0 | Cleared line of all caps and sub assemblies. changed lots. Tech adjusted probe tip 480-500, cooling 9.00-10.00, Speed V41 70-95 P. Myerson 05/26/2020 | 26-May-2020 | 26-May-2020 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Machine settings | Defect occurred due to being too hot and possibly the decompression not working. Caused by normal wear and tear and possible dirty gates. The original parameters were in spec. | Accept' 'Reject | 13279.0 1121.0 960.0 | 0.0 | Closure | CLOSED | 26-Jul-2020 | 29-Jun-2020 | 33. | 3 | Syringe
NC-IL081-22857 | 27-May-2020 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Loose Particulate | During Incoming inspection on 5/26/20 white particles were found on part 4001204, rubber plunger tips, received from Kokoku Rubber Inc. (13ea) box # 140,245,353,68,319,425,430 and 29, failing Incoming inspection. AQL is N-II, 0.25 A-7 R-8. | 20E22-1-12 | 4320000.0 |  | 27-May-2020 | 26-May-2020 | Visual- Loose Particulate | Visual- Loose Particulate | PR-Undetermined | PR-Undetermined-MEDMFGPR | Reject | 4320000.0 | 0.0 | Closure | CLOSED | 31-Jul-2020 | 03-Jun-2020 | 7. | Tips | Syringe
NC-IL081-22909 | 29-May-2020 | MARK.VENTURA | General | Internal |  |  |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-SA-V-Mixed Product 1 | Incorrect rubber tip used. Incorrect tip #4101801 used instead of #4006202 | 20E06863X | 36000.0 | line stopped, line cleared, tips removed J.S 05/29/2020 | 29-May-2020 | 29-May-2020 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Human Error | The Technical operator failed to check the shop order to verify that the correct rubber tips were being utilized. | Reject | 36000.0 | 0.0 | Closure | CLOSED | 28-Jul-2020 | 22-Jun-2020 | 24. | 1 | Syringe
NC-IL081-22969 | 02-Jun-2020 | MARK.VENTURA | General | Internal | S8F | Form, Fill Sealer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-PPSP-V-Seal Integrity 2 | "Seal integrity. Seal intact but less than 3 mm (1/8') width. cavity #20" | 20D22363X | 5760.0 | Cleared line, opened W.O.# CL-3214111 to replace gasket. K. Kaecker 06/02/2020 | 02-Jun-2020 | 02-Jun-2020 | SFF-PPSP-V-Seal Integrity 2 | Seal is intact but less than 3 mm (1/8 inch) width seal. For 60 mL and 35 mL products, less than 5 mm width. | Part Disorientation | Part disorientated causing it to get stuck on heat plate where it seals when going through multivac creating the void. | Accept | 5760.0 | 0.0 | Closure | CLOSED | 01-Aug-2020 | 17-Jul-2020 | 44. | 8 | Syringe
NC-IL081-22996 | 03-Jun-2020 | PATTI.SWISHER | General | Internal | QA | QA-IL081 |  |  |  |  | Incorrect Document Used | Audit of laminated documents issued to the label cage (aka Standard Work Library) performed 05/22/2020 found incorrect (old) revisions of three documents (DC00037181 Rev 0, FM20008544 Rev 1, and FM20008561 Rev 0). Having incorrect revisions of documents on file in the Standard Work Library could result in the incorrect revision being issued to the production floor as part of a shop order packet.\n\nTracking/trending of this NCR will be maintained using BPCS hold number 20060005. DHR for product 8507SA, Lot# 20C17263 was used to open the BPCS hold number, no product impacted. |  |  | Destroyed wrong revision documents as correct revision documents were already present. See attached immediate action objective evidence. | 03-Jun-2020 | 22-May-2020 | Documentation Incorrect | Documentation Incorrect | Inadequate Procedure | Inadequate Procedure |  |  | 0.0 | Closure | CLOSED | 31-Jul-2020 | 16-Jul-2020 | 42.
NC-IL081-23020 | 04-Jun-2020 | GRACE.AUSTIN | General | Internal | C46 | Press 46-IL081 | Equipment ID | R7461E-IL081 | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-V-Short | Manufacturing found a short on the final locking tab of the WP8536 during their inspection. This lid has an alternate name of PT00096177. The affected lot of WP8536 was running in line for Finished Good 8537SA, Lot# 20C17363 and V8536, Lot# 20C16163. No Kanban was created at this time. The production run of this lid prior to this job was under PT00096177, Lot# 20C192291. There is no Kanban on this lot in the warehouse as it was fully consumed into 8533SA. | 20C17363' '20C18463' '20C16163 | 28060.0 2350.0 130.0 24830.0 | 1) Lower the cut off from 1.75 to 1.65 and the limit velocity was raised from .53 to 1.0. Berenice Villatoro. 06/04/20\n2) Quality Alert issued for to increase awareness of defect (see attachment).\n3) One Point Lesson issued for Cell Techs and Process Techs\n4) Additional component level inspection implemented on AM-171 | 04-Jun-2020 | 29-May-2020 | CFF-V-Short | Part is not completely formed; Missing plastic | Shot Size and Hold Speed | The root cause was determined to be a combination of shot size and hold speed being incorrect. | Accept' 'Reject | 2350.0 27860.0 24830.0 130.0 200.0 | 0.0 | Closure | CLOSED | 02-Sep-2020 | 04-Aug-2020 | 61. | SS | Container
NC-IL081-23042 | 05-Jun-2020 | JACLYN.GRISKO | General | Internal | S7R | RAM-IL081 |  |  | 6492743 | 12ML LUER LOCK SYRINGE, PRINTED | SFF-SA-V-Mixed Product 2 | Incomplete assembly of product , upside rubber tip.\n\nOriginal NCR NC-IL081-22016 was voided as the system update did not allow the NC to move forward in the workflow process after pulling it back for data correction. This NCR is to close out the process for NC-IL081-22016. All paperwork associated with NC-IL081-22016 will reflect NC-IL081-22016 and will not be updated to reflect this NCR number. | 20D07363X | 800.0 | Cleared Line, scrapped box 116, replaced cables on sensors. P. Myerson 04/19/2020 | 19-Apr-2020 | 19-Apr-2020 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Conveyor Speed | In-line vibratory rail not vibrating due to wires from on/off switch not making contact at terminal strip. Wires at quick connect at the vibratory block barely making contact. \nHigh plunger sensor not working properly\nPlunger inspect station not detecting defect | Reject | 800.0 | 0.0 | Closure | CLOSED | 12-Jun-2020 | 08-Jun-2020 | 2. | 7 | Syringe
NC-IL081-23043 | 05-Jun-2020 | PATTI.SWISHER | General | Internal | CME | CME-IL081 |  |  |  |  | Incorrect Document Used | Wrong tip tumbling card FM20009239 Rev. 3 being used. Should be FM20007837 Rev.0 |  |  | NC will cover use of wrong FM document until corrected. No Impact to product as all identification elements are consistent and can be traced up and down stream.\n\nF/G 4502570 lot # 20E04663X was used to open BPCS hold # 20060013. No product was placed on QA hold. | 05-Jun-2020 | 03-Jun-2020 | Incorrect Document Used | Incorrect Document Used | Human Error | The individual printed the incorrect tote tags because they did not check the revision date effectivity. |  |  | 0.0 | Closure | CLOSED | 03-Sep-2020 | 05-Aug-2020 | 60.
NC-IL081-23092 | 09-Jun-2020 | MARIA.ROSALES | General | Internal | S7N | Pad Printer-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-P-V-Alpha Numeric 2 | Illegible Print or missing one or more letter. | 20F01663X | 10000.0 | Stopped the line, cleared hoppers, Tech added 25Lbs of pressure, increased the flow of ink to doctor blade and cliche. Cleaned up dried ink. L. Reed 06/09/2020 | 09-Jun-2020 | 09-Jun-2020 | SFF-P-V-Alpha Numeric 2 | Type 2: Illegible print or missing one or more letter | Printer Malfunction | Printer Malfunction | Accept | 10000.0 | 0.0 | Closure | CLOSED | 19-Jul-2020 | 23-Jun-2020 | 14. | 7 | Syringe
NC-IL081-23217 | 12-Jun-2020 | MARIA.ROSALES | General | Internal | CME | CME-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SPIC-F-UDI Barcode | "GTIN numbers don't match the human readable when barcode is scanned." | 20E11163X' '20E11063X' '20E03663X | 3360.0 172800.0 | Program deleted , and new one with correct information uploaded from Bell Mark. K. Kaecker 06/12/2020 | 11-Jun-2020 | 11-Jun-2020 | SFF-SPIC-F-UDI Barcode | UDI Barcode | Procedure | Root cause: GS10000267 Rev. 204 outlines requirements for Print File Maintenance on label/art files\nmanaged at machines. The procedure requires use of existing machine files for maintenance and does not\ncover the storage/use of any backup files as would have been required when faced with a corrupt file situation. As written, this procedure does not provide enough guidance. The aforementioned procedure does require verification after modification of print files. It was confirmed that no verification was performed following GS10000267 Rev. 204 requirements; therefore, failure to follow procedure is also contributing factor. | Reject' 'Accept | 3360.0 171360.0 171840.0 1440.0 960.0 | 0.0 | Closure | CLOSED | 25-Sep-2020 | 09-Oct-2020 | 119. | 3 | Syringe
NC-IL081-23225 | 12-Jun-2020 | GRACE.AUSTIN | General | Internal | C74 | Press 74-IL081 |  |  | 8985 | 8 GA HINGED LID CHEMO -YELLOW | CFF-V-Labeling - Product & Overshipper 1 | Overshipper Label has incorrect lot number 30C10763, the containers inside the box the lot# 20C10763. found by the Distribution Area. Product was made start on 04/09/20 and EOR on 04/20/20. | 20C10763 | 540.0 | Post production. Distribution Area. Person found the defect. Line Startup Process Correction not applicable. Ken Witse 06/16/20. | 12-Jun-2020 | 12-Jun-2020 | CFF-V-Labeling - Product & Overshipper 1 | Product and Lot Number is incomplete or missing | Human Error | The Manufacturing Associate failed to follow Standard Work FM20008998 and failed to examine the overshipper label for correct product code, GTIN code, Lot # and legible barcode print. | Accept | 540.0 | 0.0 | Closure | CLOSED | 11-Aug-2020 | 16-Jul-2020 | 34. | G8 | Container
NC-IL081-23226 | 12-Jun-2020 | FIDELIA.CANTOR | Inspection | Supplier | S8C | 8C-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Mix product | On 6/11/20 manufacturing reported that the wrong rubber tip (3mL) was found inside the hopper on line 8 in CME. | 20D20-1-12 | 2160000.0 |  | 11-Jun-2020 | 11-Jun-2020 | Visual- Mixed Product | Visual- Mixed Product | PR-Unknown | Unknown-MEDMFGPR | Accept | 2160000.0 | 0.0 | Closure | CLOSED | 31-Jul-2020 | 15-Jul-2020 | 33. | Tips | Syringe
NC-IL081-23231 | 12-Jun-2020 | ANGELICA.AGUILAR | General | Internal | RSN | INCORRECT RESIN-IL081 |  |  | RM165 | Resin, Greenwood blend, Pre - colored Red | Raw Material - Does Not Meet Functional Test Specification | Batch History Record revealed failed impact testing. | 00193003 | 23005.0 | N/A - raw material not consumed into Finished Goods and fully contained. | 12-Jun-2020 | 12-Jun-2020 | Raw Material - Does Not Meet Functional Test Specification | Raw Material - Does Not Meet Functional Test Specification | Machine Failure | Potential intermittent Machine Failure | Reject | 23005.0 | 0.0 | Closure | CLOSED | 31-Jul-2020 | 01-Jul-2020 | 19. | Recy | Container
NC-IL081-23269 | 15-Jun-2020 | ANGELICA.AGUILAR | General | Internal | C31 | Press 31-IL081 |  |  | 89671 | 2 GA HORIZONTAL DROP - T RED | CFF-V-Incorrect Items | Used Incorrect IFU CL63R9876B packed into FG. Per BOM should use CL63R9870B | 20E09563 | 140.0 | "Sort IFU's at the Press and removed incorrect ones per Shop Order. Ken Wiltse 06/16/20" | 15-Jun-2020 | 15-Jun-2020 | CFF-V-Incorrect Items | Incorrect IFU/Absorbent pads/Other Kit items/Other Labels in the Overshipper Box | Human Error | Human Error | Accept | 140.0 | 0.0 | Closure | CLOSED | 31-Jul-2020 | 01-Jul-2020 | 16. | G2 | Container
NC-IL081-23368 | 20-Jun-2020 | MARIA.ROSALES | General | Internal | S2R | RAM-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SA-V-Damaged Components 1 | Damaged product, non-functional. (barrel) | 20E18363X | 6720.0 | Stopped the line, call maintenance to investigate. Technicians disassembled plunger drop station/Barrel drop station, pulled back on seating rod pressure. L. Reed 06/20/2020 | 20-Jun-2020 | 20-Jun-2020 | SFF-SA-V-Damaged Components 1 | Damaged Components Type 1: Damaged product/components (non-functional) | Rod pressure | The Seating rod pressure was too high | Reject | 6720.0 | 0.0 | Closure | CLOSED | 20-Aug-2020 | 22-Jul-2020 | 32. | 3 | Syringe
NC-IL081-23374 | 21-Jun-2020 | LAKEYSHA.SIMPSON | General | Internal | C67 | Press 67-IL081 |  |  | 8934PG2 | SHARPS CONTAINER 12GAL SLIDE- | CFF-V-Label Verification and Barcode Verification 2 | Biohazard label with missing print on UDI and Product description. | 20B23663 | 400.0 | Replaced the Printer. Delsa Iradiel. 06/21/20 | 21-Jun-2020 | 21-Jun-2020 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Printer Malfunction | Printer Malfunction | Accept | 400.0 | 0.0 | Closure | CLOSED | 21-Jul-2020 | 30-Jul-2020 | 39. | G12 | Container
NC-IL081-23473 | 25-Jun-2020 | GRACE.AUSTIN | General | Internal | C56 | Press 56-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | CFF-V-Label Verification and Barcode Verification 2 | Incomplete (Missing print) UDI and lot number on biohazard label. | 20D29563 | 580.0 | Technician retreaded the printer to correct. Andrea Jacklin 06/25/20 | 25-Jun-2020 | 25-Jun-2020 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Incorrect Film Orientation | The ink ribbon was not threaded appropriately. This is a manual task and is influenced by human error. | Accept | 580.0 | 0.0 | Closure | CLOSED | 24-Aug-2020 | 07-Aug-2020 | 42. | G8 | Container
NC-IL081-23496 | 26-Jun-2020 | JACLYN.GRISKO | General | Internal | QA | QA-IL081 |  |  | V8536 | CONTAINER 3 GALLON | Non-Compliant w Procedure | Failure to hold product tied to Event-2020-03014 linked to NC-IL081-23020.\n• The product in question had already been released at the site at the time of the incident and is not representative of actively choosing to release product on hold. The site did, however, apparently fail to properly physically restrict the product, which should have raised a flag when an attempt was made to ship\n• The product was considered potentially impacted, but when tested at the site did not show evidence of defect. This absence of actual defect failed to drive the site QA team to return that product to an unreleased state | 20C16163 | 1.0 | Place all product from this lot on hold. J. Grisko 06/26/2020 | 26-Jun-2020 | 26-Jun-2020 | Non-Compliant w Procedure | Non-Compliant w Procedure | Non-Compliant w Procedure | Non-Compliant w Procedure |  |  | 0.0 | Closure | CLOSED | 31-Jul-2020 | 23-Jul-2020 | 26. | SS | Container
NC-IL081-23539 | 30-Jun-2020 | BENJAMIN.KELLY | General | Internal | S8K | Kahle-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SA-V-Mixed Product 1 | Mixed Product. Pre-fill barrels found mixed with barrels at Kale barrel feeder. | 20F16163X | 57120.0 | Shut down line, cleared out all barrels and sub assemblies. changed lots. P. Myerson 06/29/2020 | 29-Jun-2020 | 29-Jun-2020 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Failed line clearance | Line was not appropriately cleared on the end of run before the execution of the startup on product code 1181200777 was performed | Accept' 'Reject | 55680.0 1440.0 | 0.0 | Closure | CLOSED | 29-Aug-2020 | 22-Dec-2020 | 175. | 8 | Syringe
NC-IL081-23556 | 30-Jun-2020 | BENJAMIN.KELLY | General | Internal | S1R | RAM-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-M-V-Broken/Damaged 2 | Broken/Damaged plunger functional but less than desired | 20F14663X | 4800.0 | "Cleared line of all sub assembly's Scrapped 6cs, techs removed rubber bands from the plunger Adlin that were causing the sensor not to work as intended and replaced the spring. Pat Myerson 06-30-20" | 30-Jun-2020 | 30-Jun-2020 | SFF-M-V-Broken/Damaged 2 | Broken or damaged product but still Functional | Damaged spring | There was a damaged spring on the plunger aidlin | Reject | 4800.0 | 0.0 | Closure | CLOSED | 30-Aug-2020 | 23-Jul-2020 | 22. | 1 | Syringe
NC-IL081-23607 | 06-Jul-2020 | GRACE.AUSTIN | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | CFF-V-Label Verification and Barcode Verification 2 | Incorrect Product Code number printed on the Biohazard Label. Printed WP7464 insted 1522SA. | 20D18963 | 1800.0 | Reprogrammed labeler with the correct product code number on biohazard label. Ken Wiltse. 07/06/20. | 06-Jul-2020 | 06-Jul-2020 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Human Error | The Manufacturing Associate did not realize that there was a supporting shop order (R7464) at the job location and grabbed the shop order and Work Instruction Sheet from Press 43C because it was located next to the printer. The Associate then failed to follow FM20008300 which states to visually examine labels for correct product code, GTIN code, and legible barcode print. | Accept | 1800.0 | 0.0 | Closure | CLOSED | 05-Aug-2020 | 28-Jul-2020 | 21. | G2 | Container
NC-IL081-23612 | 06-Jul-2020 | BERTHA.HERNANDEZ | General | Internal | CME | CME-IL081 |  |  | 8881112985 | 10ML STERILE WATER SYR BULK NS | Procedure Not Followed | Found candy on the floor in the gowning room area. | 20F16863X | 1.0 | "Pickup the candy's from the floor. B. Hernandez 07/06/2020" | 06-Jul-2020 | 06-Jul-2020 | Procedure Not Followed | Procedure Not Followed | Human Error | The most likely root cause is that an individual brought candy into the environment which goes against GS20002295 Rev. 204. |  |  | 0.0 | Closure | CLOSED | 04-Sep-2020 | 05-Aug-2020 | 29. | 7 | Syringe
NC-IL081-23624 | 07-Jul-2020 | MARK.VENTURA | General | Internal | CME | CME-IL081 |  |  | 4006202 | Plunger Tip, 12mL, Smooth, Black, Latex Free | Defective Pkg. Material (Incoming Material) | Extra rubber material found after tumbling of rubber tips/short shot/malformed | A000183934 | 10000.0 | production stopped. Check back performed | 07-Jul-2020 | 06-Jul-2020 |  |  |  |  | Accept |  | 0.0 | Investigation | VOID |  |  |  | Tips | Syringe
NC-IL081-23652 | 08-Jul-2020 | BENJAMIN.KELLY | General | Internal | S2R | RAM-IL081 |  |  | 8881112059 | 12ML SYRINGE REGULAR NS | SFF-SA-V-Mixed Product 2 | Rubber tip not seated on to plunger. | 20F15063X | 10500.0 | Stopped the line and cleared the ram. Made adjustments to the timing of the rubber tip rod and rubber tip sensor for reject parts. Logan Reed 07-07-20 | 07-Jul-2020 | 07-Jul-2020 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Bent cyclinder pin | The side B rubber tip insert cylinder pin was bent causing the rubber tip to be dragged down the inside of the barrel on its side not allowing the plunger to fully assemble due to not being centered. | Reject | 10500.0 | 0.0 | Closure | CLOSED | 07-Sep-2020 | 17-Jul-2020 | 9. | 7 | Syringe
NC-IL081-23831 | 16-Jul-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Damage Gauge | WSV-002, Damaged - Unrepairable. |  |  | Gauge removed from the Calibration system. Damaged unit cannot yield results. | 16-Jul-2020 | 16-Jul-2020 | Calibration Damage Gauge |  | Bad power cord |  |  |  | 0.0 | Closure | CLOSED | 14-Oct-2020 | 09-Sep-2020 | 55.
NC-IL081-23885 | 18-Jul-2020 | ANGELICA.AGUILAR | General | Internal | C77 | Press 77-IL081 |  |  | 85221R | 3GA ALWAYS OPEN INROOM -T RED | CFF-V-Label Verification and Barcode Verification 8 | Biohazard label out of specification on one side of container. 3 11/16" and should be 3 5/16" +/- 3/16" as result of crooked labels. | 20C22463 | 420.0 | Left side labeler was misplacing labels due to broken fitting. Technician replaced air hose fitting on the cylinder so the label pad started moving and it was properly attaching label on the containers. O. Ratinskiy 7/17/20. | 17-Jul-2020 | 17-Jul-2020 | CFF-V-Label Verification and Barcode Verification 8 | BH - Incorrect label placement | Printer Malfunction | Printer Malfunction | Accept | 420.0 | 0.0 | Closure | CLOSED | 15-Sep-2020 | 24-Jul-2020 | 6. | G3 | Container
NC-IL081-23889 | 18-Jul-2020 | BENJAMIN.KELLY | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-PPHP-V-Heat Steak 2 | Missing heat Stake allowing for loss of sterility. | 20F08863X | 1920.0 | Clear hopper and packer of all sub-assemblies. Found a stripped screw on heat stake A side and replaced.\nPat Myerson 07/07/20 | 17-Jul-2020 | 17-Jul-2020 | SFF-PPHP-V-Heat Steak 2 | Type 2,3: One heat stake broken/missing (multiple heat stake products only) | Defective screw | Side A screw for the heat stake was stripped so it moved off to the side causing no seal. | Accept | 1920.0 | 0.0 | Closure | CLOSED | 17-Sep-2020 | 27-Jul-2020 | 9. | 3 | Syringe
NC-IL081-23890 | 18-Jul-2020 | ALEJANDRA.CASTANEDA | General | Internal | C35 | Press 35-IL081 |  |  | WP8935 | 12 GA HEAT STK LID DOOR &RAILS | PR-Parameters out of specification | One of the component of WP8935 form Press 35 (MP-1061) Mold number R7416A PCR was out of Parameter on Back Pressure it was on 300 psi since start-up, Specified Process 200-250 psi. WP8935 Lids was used at Press 54 finished goods 8938. | 20F193157' '20F06663 | 1393.0 3820.0 | Changed to specific tolerance. Minda Sailor, 07/17/20 | 18-Jul-2020 | 17-Jul-2020 | PR-Parameters out of specification | Parameters out of specification-MEDMFGPR | Human Error | Human Error | Accept | 3820.0 1393.0 | 0.0 | Closure | CLOSED | 27-Aug-2020 | 04-Aug-2020 | 17. | G18 | Container
NC-IL081-24026 | 24-Jul-2020 | GRACE.AUSTIN | General | Internal | C42 | Press 42-IL081 |  |  | 8525Y | 3 GA ROTOR/HINGE LID-YELLOW 10 | CFF-V-Packaging 1 | Incorrect packaging per Standard Work FM2008378 Rev. 2, The sealed inner carton with lids should be place on top of the two stacks of containers. | 20D25363 | 660.0 |  | 24-Jul-2020 | 24-Jul-2020 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | The Manufacturing Associate failed to follow Standard Work FM20008378 and failed to place the lids in the finished good case correctly because they did not read the standard work document. | Accept | 660.0 | 0.0 | Closure | CLOSED | 23-Aug-2020 | 29-Jul-2020 | 4. | G3 | Container
NC-IL081-24078 | 28-Jul-2020 | ANGELICA.AGUILAR | General | Internal | C74 | Press 74-IL081 |  |  | 8985S | 8 GA SLIDE LID CHEMO - YELLOW | CFF-V-Quantity per Case | Missing 2 lids into finished product (FG). Found during QA Final case inspection audit. | 20B14563 | 160.0 | Operator was instructed to follow standard work packaging FM20009004. Oleg Ratinskiy, 07/27/2020 | 27-Jul-2020 | 27-Jul-2020 | 1080 | MISSING COMPONENT | Human Error | The Manufacturing Associate failed to follow Standard Work FM20009004 and failed to place the required amount of lids in one finished good case. | Accept | 160.0 | 0.0 | Closure | CLOSED | 26-Aug-2020 | 30-Jul-2020 | 2. | G8 | Container
NC-IL081-24119 | 30-Jul-2020 | BENJAMIN.KELLY | General | Internal | S6A | 6A-IL081 |  |  |  |  | Procedure Not Followed | Failure to clean Mfg, line 6 within a calendar week per. SOP. week of 07/19/20- 07/25/20. BPCS hold number 20070027 issued for tracking/trending purposes. BPCS hold number is linked to production lot 4502570, 20G10563X due to system limitations, no product was actually impacted by this NCR. |  |  | Production was stopped and manufacturing line cleaned. | 29-Jul-2020 | 25-Jul-2020 | Procedure Not Followed |  | Human Error | Human Error |  |  | 0.0 | Closure | CLOSED | 07-Sep-2020 | 04-Aug-2020 | 5.
NC-IL081-24171 | 01-Aug-2020 | ALEJANDRA.CASTANEDA | General | Internal | C31 | Press 31-IL081 |  |  | 89671 | 2 GA HORIZONTAL DROP - T RED | CFF-V-Label Verification and Barcode Verification 8 | Biohazard label missing and misaligned. Label placement out of specification, placed more than 50% on polished window coverage. Per Standard Work FM20007202 Up to 50% of the polished window can be covered horizontally. | 20E09663 | 1420.0 | Adjusted label placement on both side labeler. Minda Sailor 08/03/2020 | 01-Aug-2020 | 01-Aug-2020 | CFF-V-Label Verification and Barcode Verification 8 | BH - Incorrect label placement | Human Error | Human Error | Accept | 1420.0 | 0.0 | Closure | CLOSED | 10-Sep-2020 | 19-Aug-2020 | 18. | G2 | Container
NC-IL081-24198 | 03-Aug-2020 | ANGELICA.AGUILAR | General | Internal | WHS | WAREHOUSE-IL081 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | Non-Compliant w Procedure | Missing pallet 5613857 with 18 cases from NC-IL081-23020. Product Code 8537SA, Lot# 20C17363. BPCS Hold# 20080002 Issued for tracking and trending purposes. | 20C17363 | 180.0 |  | 03-Aug-2020 | 03-Aug-2020 | Procedure Not Followed | Procedure Not Followed | Inadequate Procedure | Inadequate Procedure - procedure was not clear on how to handle Hold product due to capacity constraints. | Accept | 180.0 | 0.0 | Closure | CLOSED | 28-Aug-2020 | 13-Aug-2020 | 9. | SS | Container
NC-IL081-24232 | 04-Aug-2020 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | R8930LA | BIOMAX FOAM IN PLACE LID | CFF-D-Molding 3 | On 08/04/20, the lids failed the incoming inspection for the maximum height inspection (.310"), AQL S-4 4.0 (A-5, R-6). 50 lids were found out of specification per RMPS10022317 Rev. 2 and the LS10000108 Rev. 11. | 184 | 10950.0 | Placed product on hold. | 04-Aug-2020 | 04-Aug-2020 | Dimensional- out of specification | Dimensional- out of specification | Process/Method | This incident appears to be an isolated incident. There have been no previous failures for gasket height for R8930LA. The root cause was determined to be that the gasket material flow ran on the higher end of the specification during this production run. | UseAsIs | 10950.0 | 0.0 | Closure | CLOSED | 30-Nov-2020 | 06-Oct-2020 | 62. | MS | Container
NC-IL081-24245 | 05-Aug-2020 | PATTI.SWISHER | General | Internal | S6P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | PR-Other | Failed Endotoxin specification <=0.3 EU Device | 20G10563X' '20G15763X | 335000.0 | Contain lots and re-pull samples | 05-Aug-2020 | 04-Aug-2020 | Failed Pyrogen Test | Failed Pyrogen Test | Handling/Storage | Handling/Storage | Accept | 335000.0 | 0.0 | Closure | CLOSED | 15-Oct-2020 | 17-Aug-2020 | 12. | Sub | Syringe
NC-IL081-24266 | 05-Aug-2020 | BEATRIZ.GUILLEN | General | Internal | C67 | Press 67-IL081 |  |  | 8935 | CONTAINER BIOHAZARD 12GL SHARP | CFF-V-Label Verification and Barcode Verification 2 | Biohazard label missing print on product code number and UDI. | 20F08463 | 220.0 | Advanced ribbon and cleaned print head. S. Schirmer 08/05/20. | 05-Aug-2020 | 05-Aug-2020 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Human Error | Technician found that the new roll of ribbon was out of position. This was most likely due to the ribbon not being secured properly upon setup. | Accept | 220.0 | 0.0 | Closure | CLOSED | 04-Sep-2020 | 13-Aug-2020 | 7. | G12 | Container
NC-IL081-24301 | 07-Aug-2020 | BENJAMIN.KELLY | General | Internal | S7C | 7C-IL081 | Equipment ID | 58020-03-IL081 | 8881907003 | 10ML ORAL MED SYRINGE (AMBER) | SFF-M-V-Discoloration/Marbling/Color Streaks 3 | Failure of light transmission testing | 20G00763X' '20G00863X' '20G00663X | 144000.0 |  | 06-Aug-2020 | 06-Aug-2020 | SFF-M-V-Discoloration/Marbling/Color Streaks 3 | Amber samples are visually reviewed and were found to be within visual limits but failed the functional test for light. See attached test results | Mixing plate actuator binding | The Mixing plate was stalling in the blender hopper. The cylinder actuator was binding up, this caused material to not mix properly it was just being dumped. | Reject | 144000.0 | 0.0 | Closure | CLOSED | 04-Dec-2020 | 29-Oct-2020 | 83. | 7 | Syringe
NC-IL081-24335 | 07-Aug-2020 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 4006202 | Plunger Tip, 12mL, Smooth, Black, Latex Free | Visual- Cleanliness | On 08/07/20 the incoming inspection failed for Cleanliness of rubber tips on part number 4006202, lot number 00001914320010, rubber plunger tips received from West Pharmaceutical, Kearney. Found 10ea plunger tips containing loose white particles, failing the Incoming AQL (Inspection Level II, AQL 0.25, Normal (A=7, R=8) per GS20003456, Rev. 2. | 00001914320010 | 5600000.0 | Placed affected product on hold | 07-Aug-2020 | 07-Aug-2020 | Visual- Loose Particulate | Visual- Loose Particulate | PR-Undetermined | PR-Undetermined-MEDMFGPR | Reject | 5600000.0 | 0.0 | Closure | CLOSED | 23-Oct-2020 | 13-May-2021 | 278. | Tips | Syringe
NC-IL081-24338 | 08-Aug-2020 | ALEJANDRA.CASTANEDA | General | Internal | C76B | Press 76B-IL081 |  |  | WP7477 | Chemotherapy Hood with Printed Instruction Label, 8 /12 /18 | Documentation Missing | PCR not completed once per shift. PCR setting unchanged from last verified records. | 20E193078 | 1.0 |  | 08-Aug-2020 | 07-Aug-2020 | [6 empty] | 0.0 | Initiation Verification | VOID
NC-IL081-24354 | 10-Aug-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0755 could not be located. The individual to whom this gauge was issued is no longer with the company. |  |  | Gage removed from the Calibration system. | 10-Aug-2020 | 10-Aug-2020 | Calibration Lost Gauge |  | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 08-Nov-2020 | 09-Sep-2020 | 29.
NC-IL081-24361 | 10-Aug-2020 | ANGELICA.AGUILAR | General | Internal | C42 | Press 42-IL081 |  |  | 8527R | 3 GA ROTOR HINGED LID - T RED | CFF-V-Packaging 1 | Incorrect packaging per Standard Work FM20008378 Rev. 2. Lids were packed inside of containers instead on top per S.W. | 20G02063 | 680.0 | Retrain associate on Packaging Standard Work. | 10-Aug-2020 | 10-Aug-2020 | CFF-V-Packaging 1 | Product is not packaged as per packaging standards | Human Error | Human Error-The Manufacturing Associate failed to follow Standard Work FM20008378 and failed to place the lids in each finished good case in the correct orientation. | Accept | 680.0 | 0.0 | Closure | CLOSED | 10-Sep-2020 | 20-Aug-2020 | 9. | G3 | Container
NC-IL081-24417 | 12-Aug-2020 | BENJAMIN.KELLY | General | Internal | S7P | Packer-IL081 |  |  | 6492743 | 12ML LUER LOCK SYRINGE, PRINTED | SFF-SA-V-Damaged Components 1 | Damaged Components rendering part nonfunctional. Tip broken off of plunger. | 20H04563X | 7200.0 | Cleaned Line of all sub-assemblies, Scrapped 7 cases, Tech/Cell-Tech determined we had a jam in plunger Aidlin Starwheel. \nP. Myerson 08/11/20 | 11-Aug-2020 | 11-Aug-2020 | SFF-SA-V-Damaged Components 1 | Damaged Components Type 1: Damaged product/components (non-functional) | Jam | There was a jam in the plunger aidlin starwheel, one of the broken parts transported to the ram and was packed into a finished case. | Reject | 7200.0 | 0.0 | Closure | CLOSED | 09-Oct-2020 | 14-Aug-2020 | 2. | 7 | Syringe
NC-IL081-24463 | 13-Aug-2020 | BENJAMIN.KELLY | General | Internal | S8A | 8A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Broken/Damaged 1 | Broken or damaged product rendering it non-functional. Broken tip on barrel cavity # 8D. | 20H04863X | 2000.0 | Cleared line of all barrels, Cavity blocked and work order CL-323548 written for repair. \nP. Myerson 08/12/20 | 12-Aug-2020 | 12-Aug-2020 | SFF-M-V-Broken/Damaged 1 | Broken or damaged product rendering it Non Functional | Pin came out of core | Due to normal running of the mold, the pin from cavity 8D had come loose and pulled out of the core. | Reject | 2000.0 | 0.0 | Closure | CLOSED | 12-Oct-2020 | 19-Aug-2020 | 6. | Sub | Syringe
NC-IL081-24559 | 17-Aug-2020 | JACLYN.GRISKO | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Over Due | QA-335; Calibration overdue, due to staff hours reduction.\n\nOriginal NCR NC-IL081-18173 was pulled back to add information; however, due to a system issues is unable to be moved forward. This NC will be used to capture all outstanding items from the original NC. |  |  |  | 20-Nov-2019 | 07-Nov-2019 | Calibration Over Due | Calibration Over Due | Overdue |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 0.
NC-IL081-24560 | 17-Aug-2020 | JACLYN.GRISKO | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Over Due | QA-0275, Calibration Overdue, due to staff hours reduction.\n\nOriginal NCR NC-IL081-17963 was pulled back in the system to add information; however due to a system glitch, it can not be moved forward. This NCR will be used to capture all outstanding items from the original. |  |  |  | 13-Nov-2019 | 07-Nov-2019 | Calibration Over Due | Calibration Over Due | Overdue |  |  |  | 0.0 | Closure | CLOSED | 18-Sep-2020 | 17-Aug-2020 | 0.
NC-IL081-24579 | 17-Aug-2020 | BENJAMIN.KELLY | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-F-Spraymation | SprayMation test results above 6.8 average | 20H03363X | 132000.0 | Changed head 4,7,8 & 9 on time from 20 to 10. Retested with passing results.\nP. Myerson 08/17/20 | 17-Aug-2020 | 17-Aug-2020 | SFF-SA-F-Spraymation | Spraymation | Excessive silicone | No visual issues with the spraymation heads. Drip tray and dials found to have some excessive silicone. Possibly samples mis-handled. | Accept | 132000.0 | 0.0 | Closure | CLOSED | 30-Nov-2020 | 06-Nov-2020 | 80. | Sub | Syringe
NC-IL081-24588 | 18-Aug-2020 | BENJAMIN.KELLY | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Scratch / drag mark on the non fluid pathway that can be felt with a finger. Cav. #44B | 20H03363X | 81000.0 | Cleared line of all barrels and sub-assemblies, Cavity blocked.\nP. Myerson 08/17/20 | 17-Aug-2020 | 17-Aug-2020 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Damaged gate | Damaged gate on cavity 44B | Accept | 81000.0 | 0.0 | Closure | CLOSED | 01-Dec-2020 | 07-Dec-2020 | 111. | Sub | Syringe
NC-IL081-24640 | 20-Aug-2020 | GRACE.AUSTIN | General | Internal | RSN | INCORRECT RESIN-IL081 |  |  | RM165 | Resin, Greenwood blend, Pre - colored Red | PR-- PR--oduct testing fails | Missing to pull samples to perform inspection on resin batch 08/19/20. | 193757 | 19755.0 |  | 20-Aug-2020 | 20-Aug-2020 | PR-Process Step Missed | Process Step Missed-MEDMFGPR | Man | Human error. | Accept | 19755.0 | 0.0 | Closure | CLOSED | 04-Sep-2020 | 28-Aug-2020 | 8. | Recy | Container
NC-IL081-24660 | 20-Aug-2020 | BENJAMIN.KELLY | General | Internal | S8F | Form, Fill Sealer-IL081 |  |  | 8881540122 | 12ML SMARTIP MED-PREP CANNULA | SFF-PPSP-V-Seal Integrity 1 | Soft pack - Seal Integrity Cav. #16 | 20H10163X | 17280.0 |  | 20-Aug-2020 | 20-Aug-2020 | SFF-PPSP-V-Seal Integrity 1 | Void allows air into sealed package | Damaged gasket | The gasket seal that was on the multi-vac was damaged, causing a void on package 16. The damaged happened either by a molded component or plastic material entered the sealing process and caused the damage to the gasket. | Accept | 17280.0 | 0.0 | Closure | CLOSED | 20-Oct-2020 | 17-Sep-2020 | 27. | 8 | Syringe
NC-IL081-24772 | 26-Aug-2020 | BENJAMIN.KELLY | General | Internal | S1B | 1B-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-M-V-Broken/Damaged 1 | Broken or damaged product rendering it non-functional. Cav. 6 Broken luer skirt | 20H03063X | 30800.0 | Cleared Line Of barrels and sub-assemblies, Cavity 6 was blocked. Changed lots. P. Myerson 08/25/20 | 25-Aug-2020 | 25-Aug-2020 | SFF-M-V-Broken/Damaged 1 | Broken or damaged product rendering it Non Functional | Normal machine wear | The un-winder was in set position and not winding causing the piece to break. Normal machine wear. | Reject' 'Accept | 30400.0 400.0 | 0.0 | Closure | CLOSED | 01-Dec-2020 | 25-Sep-2020 | 30. | 1 | Syringe
NC-IL081-24820 | 27-Aug-2020 | BEATRIZ.GUILLEN | General | Internal | C82 | Press 82-IL081 |  |  | 31144010 | 5 QT TP SHARPS CONT -TRANS RED | Documentation Incorrect/Unclear | Press 82 (MP-1093 was using incorrect PCR at the press from Press 42 (MP-1104). | 20C21563 | 1.0 | Verified all parameters from PCR MP-1104 to PCR MP-1093. All parameters were within the range. Transcribed all entry readings from DC00087929 to the correct document form PS20003043. M. Sailor 08/27/20. | 27-Aug-2020 | 27-Aug-2020 | Documentation Incorrect/Unclear | Documentation Incorrect/Unclear | People/Human Error/ | People/Human Error/ | Accept | 1.0 | 0.0 | Closure | CLOSED | 26-Sep-2020 | 04-Sep-2020 | 7. | G5 | Container
NC-IL081-24853 | 29-Aug-2020 | LAKEYSHA.SIMPSON | General | Internal | C31 | Press 31-IL081 |  |  | 047-006 | SUT,STERNOTOMY,#7,MONOFILAMENT18IN,3/8 CIR NDL,MYO/WIRE,NS | CFF-V-Label | Missing lot number on bio hazard label. | 20F10863 | 520.0 | Printer was turned off and turned back on and was reprogrammed. Print on the label was according to specification. O. Ratinskiy 08/28/20. | 28-Aug-2020 | 28-Aug-2020 | Missing Print - Lot Number | Missing Print - Lot Number | IT/Computer Systems | IT/Computer Systems | Accept' 'Reject | 500.0 20.0 | 0.0 | Closure | CLOSED | 27-Sep-2020 | 09-Sep-2020 | 11.
NC-IL081-24889 | 31-Aug-2020 | BENJAMIN.KELLY | General | Internal | CME | CME-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | Procedure Not Followed | Failure to perform august 2020 viable testing | 20H19763X | 172800.0 |  | 31-Aug-2020 | 31-Aug-2020 | PR-- Missing Part | Inability to acquire agar plates for testing. | Component/Parts | Component/Parts |  |  | 0.0 | Closure | CLOSED | 15-Oct-2020 | 29-Sep-2020 | 28. | 3 | Syringe
NC-IL081-25068 | 09-Sep-2020 | GRACE.AUSTIN | Inspection | Supplier | INC | Incoming-INC-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Visual- Loose Particulate | During Incoming inspection on 09/09/20 white particles were found on part 4001204, rubber plunger tips, received from Kokoku Rubber Inc. (12ea) box # 70, 98, 143, 179, 215, 249, 317 and 392, failing Incoming inspection. AQL is T-II, 0.25 A/R = 5/6. | 20J04 -1 - 12 | 4320000.0 | Place all affected product on Hold | 09-Sep-2020 | 09-Sep-2020 | Visual- Loose Particulate | Visual- Loose Particulate | PR-Undetermined | PR-Undetermined-MEDMFGPR---Investigation documented in SCAR CAPA-IL081-01715 | Reject | 4320000.0 | 0.0 | Closure | CLOSED | 24-Oct-2020 | 16-Sep-2020 | 6. | Tips | Syringe
NC-IL081-25117 | 11-Sep-2020 | ANGELICA.AGUILAR | General | Internal | 171 | Cell 171-IL081 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-F-Assembly 3 | Non-functional lids (warped Lids R7461E) packed into finished product. | 20F08063 | 280.0 | Blew out the mold to get any blockage out. Guillermo Cruz, 19/10/20 | 10-Sep-2020 | 10-Sep-2020 | CFF-F-Assembly 3 | Manual Function Test Failure | Blockage | Blockage in the mold caused by sitting water within the manifold that builds up with rust internally, causing the mold to get hot and parts to warp upon ejection. | Reject | 280.0 | 0.0 | Closure | CLOSED | 09-Nov-2020 | 25-Sep-2020 | 14. | SS | Container
NC-IL081-25126 | 11-Sep-2020 | BEATRIZ.GUILLEN | General | Internal | C30 | Press 30-IL081 |  |  | 8534SA | 2 GA SHARPSTAR - TRANS. RED | CFF-V-Blisters, Bubbles and Voids | Splay Marks and bubbles out of specification on bottom of container. | 20E16763 | 2700.0 | Screw evaluated. Work Order CL-3243028. G. Cruz 09/11/20. | 11-Sep-2020 | 11-Sep-2020 | CFF-V-Blisters, Bubbles and Voids | Single bubble greater than or equal to 0.080”; Up to 3 bubbles between 0.060 and 0.080”; Up to 4 bubbles less than or equal to 0.060”; Up to 8 bubbles less than 0.030”; Living Hinge = Reject All | Machine Failure | Machine Failure | Accept | 2700.0 | 0.0 | Closure | CLOSED | 11-Oct-2020 | 09-Oct-2020 | 28. | SS | Container
NC-IL081-25160 | 14-Sep-2020 | BENJAMIN.KELLY | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | Out of Spec Product Functional | Inner cartons failed UDI Barcode Scan. | 20J00663X' '20J00763X | 420.0 4.0 | Cleaned print Head.\nPat Myerson 09/14/20 | 14-Sep-2020 | 14-Sep-2020 | Out of Spec Product Functional | Out of Spec Product Functional | Packer up/down |  | Accept | 420.0 4.0 | 0.0 | Closure | CLOSED | 14-Oct-2020 | 21-Sep-2020 | 6. | 3 | Syringe
NC-IL081-25165 | 15-Sep-2020 | BERTHA.HERNANDEZ | General | Internal |  |  |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SPBP-V-Labeling/Print 6 | Failed UDI grade for inner cartons (grade F) | 20J00763X | 6.0 | Stopped and clean the ink cartage. Started check back. | 15-Sep-2020 | 15-Sep-2020 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | 3 | Syringe
NC-IL081-25255 | 18-Sep-2020 | BENJAMIN.KELLY | General | Internal | CME | CME-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | PR-Environmental | Spider found in carton during rework operation. | 19K15863X | 1000.0 | Case 331 was scrapped out. Rework table was cleaned with isopropyl alcohol. Pat Myerson 09/17/20 | 17-Sep-2020 | 17-Sep-2020 | Contamination - insect | Contamination - insect | No Problems Found | No Problems Found | Reject | 1000.0 | 0.0 | Closure | CLOSED | 12-Oct-2020 | 21-Oct-2020 | 33. | Sub | Syringe
NC-IL081-25267 | 18-Sep-2020 | BERTHA.HERNANDEZ | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Outside Validated Parameters | Spraymation head temperature was out of specification. found at 170 spec is 175-195. | 20J05363X' '20J05263X' '20J05163X | 335000.0 291000.0 | Parameter brought back into spec & restarted the line.\nKim Kaecker 09/18/20 | 18-Sep-2020 | 15-Sep-2020 | Outside Validated Parameters | Outside Validated Parameters | Human Error | The Cell Tech failed to follow GS10004646 Rev 208 Section 6.3 which requires that the PCR readings are recorded and verified once per shift. When the cell tech took their PCR reading, they recorded the temperature at 170. The range is 175 to 195. The cell tech failed to recognize that the fill time was out of parameters and continued to run. | Accept | 335000.0 291000.0 | 0.0 | Closure | CLOSED | 18-Oct-2020 | 08-Oct-2020 | 20. | Sub | Syringe
NC-IL081-25280 | 18-Sep-2020 | LAKEYSHA.SIMPSON | General | Internal |  |  |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-Drag Marks /Gouges/Nicks Scratches 2 | Scratch outside of barrel that can be felt with finger. | 20J05463X | 38000.0 | Cleared line of all barrels and subassemblies. Blocked cavity 11. Changed lots. \nPat Myerson 09/18/20 | 18-Sep-2020 | 18-Sep-2020 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Damaged gate | Cavity 11 has a damaged gate | Reject' 'Accept | 37000.0 1000.0 | 0.0 | Closure | CLOSED | 18-Oct-2020 | 25-Sep-2020 | 6. | Sub | Syringe
NC-IL081-25356 | 22-Sep-2020 | MARIA.ROSALES | General | Internal | C90 | Press 90-IL081 |  |  | WP7478 | CHEMO CAP DISC ASSM | Documentation Missing | WP7478/ 20H193894 Failed to perform Start-up checklist after mold R7051A and R7478B were pulled for PM. | 20H193894 | 1.0 | Start-Up checklist was filled out and First Articles pulled. S. Schirmer 09/22/2020. | 22-Sep-2020 | 22-Sep-2020 | Documentation Missing | Documentation Missing | Human Error | The Lead on shift was not aware that the tool had been pulled for PM which would require a start-up checklist before starting per GS10000122. There is no trend identified for this failure type and it is likely to be a one-off incident. | Accept | 1.0 | 0.0 | Closure | CLOSED | 22-Oct-2020 | 02-Oct-2020 | 10. | G8 | Container
NC-IL081-25386 | 23-Sep-2020 | LAKEYSHA.SIMPSON | General | Internal |  |  |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF--PPSP-F-Labeling/Print7 | Failed inner carton UDI grade barcode scan due to missing print on label. | 20J09463X | 250.0 | Printer switched out and reprogrammed. K. Kaecker 09/23/2020 | 23-Sep-2020 | 23-Sep-2020 | SFF-SPIC-F-UDI Barcode | UDI Barcode | Bad Printer | The labeler needed to be switched out for it to get repaired. Work order CL-3249190 opened and printer red tagged. | Accept | 250.0 | 0.0 | Closure | CLOSED | 23-Oct-2020 | 30-Sep-2020 | 6. | 412 | Syringe
NC-IL081-25451 | 25-Sep-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-634, Scale is "Out of Tolerance - As Found". |  |  | Scale removed from service 09/16/2020. | 25-Sep-2020 | 16-Sep-2020 | Calibration Out of Tolerance |  | Wear and Tear | The scale was used consistently and was near the end of its life cycle. |  |  | 0.0 | Closure | CLOSED | 25-Oct-2020 | 07-Oct-2020 | 12.
NC-IL081-25533 | 30-Sep-2020 | BENJAMIN.KELLY | General | Internal |  |  |  |  | R7642B | Barrel, 12 mL | SFF-M-V-Flash/Gate 3 | Flash on non-functional area. Loose flash on the shoulder O/D. Cav. A3 | 20J14963 | 9000.0 | Cleared Line. Process tech fixed issue by Lowering shot size and injection speed. Changed lots.\nPat Myerson 09-29-20 | 29-Sep-2020 | 29-Sep-2020 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | excessive plastic | When the process tech closed a cavity earlier in the day he did not adjust the shot size enough causing excessive plastic going into the mold and causing flash on barrel A3. | Reject | 9000.0 | 0.0 | Closure | CLOSED | 29-Oct-2020 | 08-Oct-2020 | 8. | 8 | Syringe
NC-IL081-25597 | 02-Oct-2020 | BERTHA.HERNANDEZ | General | Internal | S2N | Hot Stamp Printer-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-P-V-Numbering 1 | Missing one or more numerals on barrel print. | 20J14763X | 7200.0 | Cleared line of all printed barrels and sub-assemblies. Scrapped 9cs. Changed all the blankets and the cliché on #1.\nPatricia Myerson 10/01/20 | 02-Oct-2020 | 01-Oct-2020 | SFF-P-V-Numbering 1 | Type 1: Missing 1 or more numeral(s) | Loose screw | A screw that holds the blankets in place was loose and protruding out. | Reject | 7200.0 | 0.0 | Closure | CLOSED | 02-Dec-2020 | 07-Oct-2020 | 5. | 7 | Syringe
NC-IL081-25650 | 06-Oct-2020 | ALEJANDRA.CASTANEDA | General | Internal | C42 | Press 42-IL081 |  |  | 8525Y | 3 GA ROTOR/HINGE LID-YELLOW 10 | Out of Parameters | PCR out of parameters. Hold Pressure reading documented 8000, PCR specification 8217 -9082 psi. | 20J10363 | 1.0 | Affected Product was removed. Process Technician changed Hold PSI to 8250 to be within process parameters. Berinice Villatoro 10/05/20 | 05-Oct-2020 | 05-Oct-2020 | Out of Parameters |  | Human Error | Human Error |  |  | 0.0 | Closure | CLOSED | 05-Nov-2020 | 27-Oct-2020 | 21. | G3 | Container
NC-IL081-25652 | 06-Oct-2020 | MARK.VENTURA | General | Internal | S2B | 2B-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-M-V-Shorts 1 | Shorts in fluid pathway located at barrel tip cavity 23B | 8881512852 | 23040.0 | Closed cavity 23B and scrapped all affected product. F Bahena 10/06/2020 | 06-Oct-2020 | 06-Oct-2020 | SFF-M-V-Shorts 1 | Shorts in fluid pathway or in area where Sterility is required | Mold Wear or Damage | Mold Wear or Damage | Reject | 23040.0 | 0.0 | Closure | CLOSED | 05-Nov-2020 | 19-Oct-2020 | 13. | 3 | Syringe
NC-IL081-25697 | 07-Oct-2020 | BERTHA.HERNANDEZ | General | Internal | S7P | Packer-IL081 |  |  | 8881907003 | 10ML ORAL MED SYRINGE (AMBER) | SFF-SPIC-V-Labeling/Print 1 | Incorrect Expiration date on inner carton labels, found during QA final inspection labels with EXP 2025-10-07 correct date is 2025-10-02. | 20J21663X | 1085.0 | Reprogrammed inner carton label printer with correct information. K. Kaecker 10/07/2020 | 07-Oct-2020 | 06-Oct-2020 | SFF-SPIC-V-Labeling/Print 1 | Lot Number and Expiration Date (where applicable) is incomplete or missing | Human Error | Human Error | Accept | 1085.0 | 0.0 | Closure | CLOSED | 13-Nov-2020 | 26-Oct-2020 | 19. | 7 | Syringe
NC-IL081-25730 | 08-Oct-2020 | MARK.VENTURA | General | Internal | S1B | 1B-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-M-V-Deformation 1 | Pulled ring stop, cavity 47 | 20J20263X' '20J14363X | 345600.0 20000.0 | Blocked cavity 47 and cleared the line, scrapped all affected product and from molding. F. Bahena 10/08/2020 | 08-Oct-2020 | 08-Oct-2020 | SFF-M-V-Deformation 1 | The ring stop has been pulled from the inside dimension of the barrel possibly rendering the barrel unusable | un-balanced mold cavitation | Mold cavitation reduced causing variation during the filling process | Accept' 'Reject | 20000.0 5620.0 339980.0 | 0.0 | Closure | CLOSED | 31-Dec-2020 | 22-Jan-2021 | 106. | 1 | Syringe
NC-IL081-25745 | 08-Oct-2020 | BERTHA.HERNANDEZ | General | Internal | S7N | Pad Printer-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-P-V-Alpha Numeric 2 | Illegible print and missing one or more letters on word Monoject on barrel print. | 20J21163X' '20J21663X | 2000.0 7000.0 | Line stopped and cleared, print pad cleaned, due to being dry. K. Kaecker 10/08/2020 | 08-Oct-2020 | 08-Oct-2020 | SFF-P-V-Alpha Numeric 2 | Type 2: Illegible print or missing one or more letter | Dried ink |  | Reject | 2000.0 7000.0 | 0.0 | Closure | CLOSED | 07-Nov-2020 | 21-Oct-2020 | 13. | 7 | Syringe
NC-IL081-25784 | 09-Oct-2020 | GRACE.AUSTIN | General | Internal | Shipping | Shipping-MYS03 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | PR-Process Violation | QA Crew C, unable to locate the 6 pallets of 8537SA. it was moved to Distribution (F-12) without QA Green Released Stickers. | 20F17463 | 1.0 |  | 09-Oct-2020 | 09-Oct-2020 | PR-Process Violation | Process Violation-MEDMFGPR | Human Error | Human Error | Accept | 1.0 | 0.0 | Closure | CLOSED | 08-Nov-2020 | 23-Oct-2020 | 14. | SS | Container
NC-IL081-25794 | 09-Oct-2020 | LAKEYSHA.SIMPSON | General | Internal | S1R | RAM-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-SA-V-Mixed Product 2 | Incorrect assembly of product.(2)Missing rubber tip found in F/G case during checkback | 20J020263X | 105600.0 |  | 09-Oct-2020 | 09-Oct-2020 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Incorrect Setup | Two fiber optics were crossed between amplifiers which gave false readings and increased the risk of missing rubber tips making its way into finished goods. | Accept | 105600.0 | 0.0 | Closure | CLOSED | 08-Mar-2021 | 21-Dec-2020 | 72. | 1 | Syringe
NC-IL081-25912 | 13-Oct-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0909; Reported Lost by Veronica. MD-0907 issued to PR-54 as a replacement. |  |  |  | 13-Oct-2020 | 13-Oct-2020 | Calibration Lost Gauge | Calibration Lost Gauge | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 12-Nov-2020 | 27-Oct-2020 | 14.
NC-IL081-25941 | 14-Oct-2020 | MARK.VENTURA | General | Internal | S1P | Packer-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | Visual- Identification not per spec | Corrugated part number 5500149 and Manufacturer certificate not printed on overshipper. | 20J20463X' '20K04663X | 17.0 95.0 | Removed unprinted boxes from line. D. Breceda 10/14/2020 | 14-Oct-2020 | 14-Oct-2020 | Packaging Artwork - Poor / Missing Print | Packaging Artwork - Poor / Missing Print | Packaging & Labeling Processes | Packaging & Labeling Processes | Accept | 17.0 95.0 | 0.0 | Closure | CLOSED | 13-Nov-2020 | 30-Oct-2020 | 16. | 1 | Syringe
NC-IL081-25942 | 14-Oct-2020 | MARK.VENTURA | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 5500149 | Corrugated Case, Unprinted, 16.75 Inch x 14.125 Inch x 11.12 | Packaging Artwork - Poor / Missing Print | On 10/13/20 in the Syringe Focus Factory, it was discovered that on line 1 and line 412, corrugate received from Packaging Corporation of America (PCA), part number 5500149, did not have either the part number nor Manufacturer certification printed on over shipper as required per RM10008013, Rev. 7. | 525485 | 1003.0 | Stopped manufacturing and isolated unused corrugate for SCAR. | 13-Oct-2020 | 13-Oct-2020 | Visual- Identification not per spec | Visual- Identification not per spec | Process/Method | Process/Method-- The change in print registration after the machine jammed up caused this print issue. The operator also did not check 1 carton/unit as required with the quality check process | Reject | 1003.0 | 0.0 | Closure | CLOSED | 28-Feb-2021 | 13-Dec-2020 | 60.
NC-IL081-26009 | 15-Oct-2020 | PATTI.SWISHER | General | Internal | CME | CME-IL081 |  |  |  |  | Environmental Monitoring | Non-Viable testing failures for the month of October, North CME locations 45 and 46, Packaging Locations 16,18,19,20,21,23,24,27 and 28 | 20H08863X | 1.0 |  | 15-Oct-2020 | 13-Oct-2020 | Environmental Particulate (OOS) | Environmental Particulate (OOS) | Environmental monitoring parameters exceeded limits | Environmental monitoring parameters exceeded limits |  |  | 0.0 | Closure | CLOSED | 11-Dec-2020 | 01-Dec-2020 | 47.
NC-IL081-26130 | 20-Oct-2020 | BENJAMIN.KELLY | General | Internal | S1R | RAM-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-SA-V-Damaged Components 3 | Damaged product/Component functional but less than desired. Damaged thumb pad. | 20K04063X | 7300.0 | Cleared ram to packer of all plungers and sub-assys. Replaced the rubber bands on the plunger drop station & tightened loose tube on the side 2 plunger adlin.\nP. Myerson 10/19/20 | 19-Oct-2020 | 19-Oct-2020 | SFF-SA-V-Damaged Components 3 | Damaged Components Type 3: Damaged product/component (functional but less than desired); Scratches/drag marks on the barrel | Sensor not working | Worn rubber bands on plunger drop station at the ram and Side 1 plunger aidlin low level sensor not working properly. | Reject | 7300.0 | 0.0 | Closure | CLOSED | 18-Dec-2020 | 27-Oct-2020 | 7. | 1 | Syringe
NC-IL081-26189 | 21-Oct-2020 | ANGELICA.AGUILAR | General | Internal | C56 | Press 56-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | Unapproved Material | Un-tested resin RM165 from Silo 1 was released into approved resin in Silo 5. | 20G14463 | 7380.0 | Resin transfer from silo to production was stopped. G. Cruz Jr. 10/21/20 | 21-Oct-2020 | 21-Oct-2020 | Unapproved Material | Unapproved Material | Inadequate Procedure | Inadequate Procedure | Accept | 7380.0 | 0.0 | Closure | CLOSED | 11-Nov-2020 | 28-Oct-2020 | 7. | G8 | Container
NC-IL081-26212 | 22-Oct-2020 | BENJAMIN.KELLY | General | Internal | S7N | Pad Printer-IL081 |  |  | 8881112059 | 12ML SYRINGE REGULAR NS | SFF-P-V-Ink Rings 1 | Print on syringe - ink-rings on the outside of syringe barrel. | 19G24963X | 1.0 |  | 22-Oct-2020 | 22-Oct-2020 | SFF-P-V-Ink Rings 1 | Ink ring present that interferes with graduation line | Procedure/Method | Procedure/Method | Accept | 1.0 | 0.0 | Closure | CLOSED | 06-Dec-2020 | 19-Nov-2020 | 28. | 7 | Syringe
NC-IL081-26213 | 22-Oct-2020 | LAKEYSHA.SIMPSON | General | Internal | S6C | 6C-IL081 |  |  | 4502575 | 12 CC SYR BLACK PT SFF | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag marks in Non-Fluid Pathway that can be felt with bare finger. | 19K16863X | 264000.0 |  | 22-Oct-2020 | 30-Oct-2019 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Material blocked gate | An assessment of the checkback results showed that the defect was inconsistent in nature which would be more representative of a material blocked gate versus a damaged gate as a damaged gate would be a permanent defect type. | Reject | 264000.0 | 0.0 | Closure | CLOSED | 06-Dec-2020 | 01-Dec-2020 | 39. | Sub | Syringe
NC-IL081-26364 | 27-Oct-2020 | ANGELICA.AGUILAR | General | Internal | C82 | Press 82-IL081 |  |  | 31143699 | 1 GA SHARPS CONTAINER | Out of Spec Visual Inspection (Visual Defect) | Biohazard label with missing print on product code and lot number. | 20F11363 | 1792.0 | Ribbon roll was set-up properly, printer was restarted and print on the labels was within the specs. O. Ratinskiy 10/26/20 | 26-Oct-2020 | 26-Oct-2020 | Missing Print - Lot Number | Missing Print - Lot Number | Printer Malfunction | Printer Malfunction | Accept | 1792.0 | 0.0 | Closure | CLOSED | 25-Nov-2020 | 12-Nov-2020 | 16. | G1 | Container
NC-IL081-26424 | 28-Oct-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-557 lost; 6" ruler; Daniel Martinez states that the ruler is lost; discovered during normal calibration rounds on 10/28/2020. |  |  |  | 28-Oct-2020 | 28-Oct-2020 | Calibration Lost Gauge | Calibration Lost Gauge | Lost Gauge |  |  |  | 0.0 | Closure | CLOSED | 27-Nov-2020 | 06-Nov-2020 | 8.
NC-IL081-26456 | 29-Oct-2020 | ANGELICA.AGUILAR | General | Internal | C82 | Press 82-IL081 | Equipment ID | R7103A-IL081 | 31143699 | 1 GA SHARPS CONTAINER | Failure to follow Procedure | Missing Start Up Checklist after the Tool was removed for repair and set-up back for production on 10/21/2020. | 20F11363 | 15616.0 | End of run already performed. No further action required. | 29-Oct-2020 | 21-Oct-2020 | Failure to follow Procedure |  | Human Error | Human Error | Accept | 15616.0 | 0.0 | Closure | CLOSED | 08-Nov-2020 | 29-Oct-2020 | 0. | G1 | Container
NC-IL081-26485 | 30-Oct-2020 | PATTI.SWISHER | General | Internal | D1 | Defib-1-MA028 |  |  | 3150450 | Barrel, 12 mL Luer Lock with Finger Flange | Documentation Missing | DC00073712 PCR for P/N 3150450 missing | 20H193802 | 1.0 |  | 30-Oct-2020 | 17-Aug-2020 | [6 empty] | 0.0 | Investigation | VOID
NC-IL081-26495 | 30-Oct-2020 | BERTHA.HERNANDEZ | General | Internal | S8D | 8D-IL081 |  |  | 1181200777T | 12ML SYRINGE LUER LOCK TIP | SFF-M-V-Flash/Gate 3 | Flash in non-functional area exceeds 1/16" on plunger ; cavity 22E, mold # 58121 | 20K11963X | 2880.0 | Cleared line of all plungers and sub-assemblies. Scrapped case 131 - 136. Found a part stuck in the mold, It was removed by the cell tech.\nP. Myerson 10-30-20 | 30-Oct-2020 | 30-Oct-2020 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | "Part didn't eject" | Defect caused by the mold closing on a part that failed to eject resulting in a gap along the parting line and flashed parts | Reject | 2880.0 | 0.0 | Closure | CLOSED | 30-Dec-2020 | 04-Nov-2020 | 4.
NC-IL081-26549 | 02-Nov-2020 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | CL63R1776 | Corrugate Case, Printed, 25.25 Inch x 11.00 Inch x 21.50 Inc | Packaging Artwork - Poor / Missing Print | On 10/31/20 in the containers Focus Factory, it was discovered that on AM-131 cell, corrugate received from American Eagle Packaging Corp. part number CL63R1776, did not have the part number printed on over shipper as required per GS00109707 (Master specification for corrugated fiberboard boxes). | 139292 | 2826.0 | All affected product Placed on hold. | 02-Nov-2020 | 02-Nov-2020 | Visual- Identification not per spec | Visual- Identification not per spec | Process/Method | Process/Method- The replacement was a temporary banded part# that was taped to the mount and fell off during the run. | Reject | 2826.0 | 0.0 | Closure | CLOSED | 30-Nov-2020 | 19-Nov-2020 | 16.
NC-IL081-26585 | 03-Nov-2020 | MARK.VENTURA | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Scratch in non-fluid pathway, can be felt with bare finger | 20K16763X | 44000.0 | Barrel entry dial to far pushed in. K. Kaecker 11/03/2020 | 03-Nov-2020 | 03-Nov-2020 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Part Out of Pocket | Part Out of Pocket | Reject | 44000.0 | 0.0 | Closure | CLOSED | 03-Dec-2020 | 05-Nov-2020 | 2. | Sub | Syringe
NC-IL081-26630 | 04-Nov-2020 | PATTI.SWISHER | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Grease/Lint/Particulate 1 | particulate inside of the fluid pathway, | 19K20663X | 8.0 | Scar-20-014 | 04-Nov-2020 | 04-Sep-2020 | SFF-SA-V-Particulate 1 | Particulate Type 1: Extraneous matter in fluid pathway | Human Error | Based on the investigation a root cause could not be confirmed, however the most probable cause is failure to clean following maintenance activities during the production lot. | Reject | 8.0 | 0.0 | Closure | CLOSED | 24-Dec-2020 | 04-Dec-2020 | 29. | Sub | Syringe
NC-IL081-26672 | 05-Nov-2020 | LAKEYSHA.SIMPSON | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks in Non-Fluid Pathway. | 20K17063X | 87000.0 | Cleared line of all barrels and sub-assemblies. Process tech the cavity by cleaning the gate. Lot was cut short and ended.\n P. Myerson 11/05/20 | 05-Nov-2020 | 05-Nov-2020 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Plastic caught in gate | Plastic had gotten caught in the gate on cavity 13, it was sticking out enough to cause the drag mark on the outside of the barrel. | Accept | 87000.0 | 0.0 | Closure | CLOSED | 05-Jan-2021 | 01-Feb-2021 | 88. | Sub | Syringe
NC-IL081-26700 | 06-Nov-2020 | LAKEYSHA.SIMPSON | General | Internal | S4A | 412A-IL081 |  |  | R7642B | Barrel, 12 mL | SFF-M-V-Flash/Gate 1 | Flash in fluid pathway exceeds 1/64" cavity A3 | 20J14963 | 9000.0 |  | 06-Nov-2020 | 06-Nov-2020 | SFF-M-V-Flash/Gate 1 | Flash Type 1: Flash in Fluid Path (Luer Tip) exceeds 1/64” (.016”). Does not meets drawing specification | Shot Size | A cavity was closed earlier in the run and shot size was not adjusted enough causing excessive plastic going through the mold resulting in flash. | Reject | 9000.0 | 0.0 | Closure | CLOSED | 06-Dec-2020 | 20-Nov-2020 | 14. | 8 | Syringe
NC-IL081-26764 | 09-Nov-2020 | PATTI.SWISHER | General | Internal | CME | CME-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | PR-Environmental | Bug found in CME rework area, inside bag, pallet case # 11-251 and 11-255, Lot # 19K17763X, worm/caterpillar found in case of Lot# 19K17863X | 19K17763X' '19K17863X | 335000.0 | MFG removed and scrapped affected cases. | 09-Nov-2020 | 09-Nov-2020 | Contamination - Insect/Rodent | Contamination - Insect/Rodent | Dirt/Foreign Matter found in/on the product | Dirt/Foreign Matter found in/on the product | Accept | 335000.0 | 0.0 | Closure | CLOSED | 07-Jan-2021 | 15-Dec-2020 | 36. | Sub | Syringe
NC-IL081-26791 | 10-Nov-2020 | LAKEYSHA.SIMPSON | General | Internal | S6C | 6C-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 1 | Flash in fluid pathway (luer tip) that exceeds 1/64" . mold 58114 | 20K21363X | 53000.0 | Cleared line and blocked cavity 38B W.O.# CL-3259471 K. Kaecker 10/11/2020 | 10-Nov-2020 | 10-Nov-2020 | SFF-M-V-Flash/Gate 1 | Flash Type 1: Flash in Fluid Path (Luer Tip) exceeds 1/64” (.016”). Does not meets drawing specification | Tool Damage | Tool damage in the tip caused excess plastic which resulted in unacceptable tip flash. | Reject | 53000.0 | 0.0 | Closure | CLOSED | 10-Dec-2020 | 18-Nov-2020 | 7. | Sub | Syringe
NC-IL081-26879 | 12-Nov-2020 | BENJAMIN.KELLY | General | Internal | S6B | 6B-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | PR-Parameters out of specification | Parameter out of spec. Hold Time 3.44 Spec. is 3.70 - 4.20 | 20K21463X' '20K21363X | 335000.0 60000.0 |  | 12-Nov-2020 | 11-Nov-2020 | PR-Parameters out of specification | Parameters out of specification-MEDMFGPR | Human Error | Changes to parameters are manual in nature and without the technician or Cell Tech documenting the change on the PCR. This error could not be driven down to the individual. There is no trend for this type of issue identified. | Accept | 335000.0 60000.0 | 0.0 | Closure | CLOSED | 12-Dec-2020 | 04-Dec-2020 | 21. | Sub | Syringe
NC-IL081-26900 | 13-Nov-2020 | MARK.VENTURA | General | Internal | S3N | Hot Stamp Printer-IL081 |  |  | 8881512852 | 12ML SYRINGE REGULAR LUER RP | SFF-P-V-Numbering 1 | Missing 1 or more numbers (barrel print) | 20L02163X | 5760.0 | Changed out one cliché, one pad, 6 mandrels, 6 white tips, adjust ink roller angle, adjusted pressure from cliché to pad, and recalibrated all pads. P. Myerson 11/13/2020 | 13-Nov-2020 | 13-Nov-2020 | SFF-P-V-Numbering 1 | Type 1: Missing 1 or more numeral(s) | Incorrect Pressure | Incorrect Pressure | Reject | 5760.0 | 0.0 | Closure | CLOSED | 13-Dec-2020 | 25-Nov-2020 | 12. | 3 | Syringe
NC-IL081-26915 | 13-Nov-2020 | BERTHA.HERNANDEZ | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Mixed Product 2 | Incomplete assembly of product, rubber tip is not blocking enough of the nozzle to prevent the silicone from drying out or particulate from getting in. | 20K21663X | 16000.0 | Cleared line of all sub assemblies. Tech removed a plunger that was jammed under the wheel that pushes the plunger/tip assembly into the barrel. P. Myerson 11/13/2020 | 13-Nov-2020 | 13-Nov-2020 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Misaligned Gate | The gate for part guide was misaligned and out of place. | Reject | 16000.0 | 0.0 | Closure | CLOSED | 13-Dec-2020 | 25-Nov-2020 | 11. | Sub | Syringe
NC-IL081-27025 | 17-Nov-2020 | LAKEYSHA.SIMPSON | General | Internal | S7N | Pad Printer-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-P-V-ALPHA NUMERIC 2 | Illegible print on barrel. | 20J22163X | 23000.0 | Clear line of all barrels and sub assemblies, cleaned the ink pad and tray. Due to caked up ink, changed lots. P. Myerson 11/17/2020 | 17-Nov-2020 | 17-Nov-2020 | SFF-P-V-Alpha Numeric 2 | Type 2: Illegible print or missing one or more letter | Dried up ink on equipment | The ink pad and tray had caked up dried ink which was causing inconsistent barrel print. | Reject | 23000.0 | 0.0 | Closure | CLOSED | 29-Jan-2021 | 21-Dec-2020 | 33. | 7 | Syringe
NC-IL081-27157 | 21-Nov-2020 | ANGELICA.AGUILAR | General | Internal | C70 | Press 70-IL081 |  |  | 31325410 | FISHER 1 QT PHLEBOTOMY | Label information incorrect | QA Technician found during her scheduled inspection a double and different time printed on over shipper label. Traceability issue. | 20H12763 | 16830.0 |  | 20-Nov-2020 | 20-Nov-2020 | Label information incorrect | Label information incorrect | Label Template Design | Label Template Design created double time printed on the overshipper label. | Accept | 16830.0 | 0.0 | Closure | CLOSED | 31-Dec-2020 | 22-Dec-2020 | 31. | G1 | Container
NC-IL081-27218 | 23-Nov-2020 | BERTHA.HERNANDEZ | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SPIC-F-UDI Barcode | Failed UDI barcode. | 20L11863X | 504.0 | Cell Tech replaced inner carton ink cartridges and cleaned them before restarting.\nR. Hamilton 11/23/20 | 23-Nov-2020 | 21-Nov-2020 | SFF-SPIC-F-UDI Barcode | UDI Barcode | Inconsistent Feed |  | Accept | 504.0 | 0.0 | Closure | CLOSED | 23-Dec-2020 | 02-Dec-2020 | 8. | 3 | Syringe
NC-IL081-27231 | 24-Nov-2020 | BENJAMIN.KELLY | General | Internal | S1P | Packer-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-SA-V-Mixed Product 2 | Incomplete assembly of product; missing rubber tip. | 20L06663X | 2400.0 | Cell tech cleaned silicone from rubber tip rail so rubber tips would advance properly & checked to make sure each sensor was functioning properly.\nR. Hamilton 11/23/20 | 23-Nov-2020 | 23-Nov-2020 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Dirty sensor | The sensor on side 1A was not working properly due to a build up of debris and silicone. Sensor was unable to detect when a rubber tip was missing. | Reject | 2400.0 | 0.0 | Closure | CLOSED | 22-Jan-2021 | 02-Dec-2020 | 8. | 1 | Syringe
NC-IL081-27435 | 02-Dec-2020 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 4101801 | TIP 12 CC RAYCO WG2714 RUBBER TIP | Visual- Embedded Particulate (not metal) | Sister site, KenMex reported that their Customer, SCA, reported that there was an insect embedded in the syringe plunger tip good part 8881512258. No patient injury reported. Lot # for 4101801 rubber tips would be from YL050619-22, YL090619-011 or YL130619-017 (plunger tip lots used to assemble the finished good affected in this NCR). PO # 819855. | YL050619-22' 'YL130619-017' 'YL090619-011 | 1.0 | Defect was found in 1 part by the Customer (SCA). Customer performs 100% inspection. No other defects were reported. No patient injury reported. | 02-Dec-2020 | 05-Nov-2020 | Visual- Embedded Particulate (not metal) | Visual- Embedded Particulate (not metal) | Procedure/Method | Procedure/Method\n\nRayco determined that the current pest control action may be not strong enough in the mosquito season, which cause the insect have the possibility to contaminate the preform film, was molded in the part | Accept | 1.0 | 0.0 | Closure | CLOSED | 28-Feb-2021 | 12-Jan-2021 | 40. | Tips | Syringe
NC-IL081-27480 | 04-Dec-2020 | BENJAMIN.KELLY | General | Internal | Rework | Rework |  |  |  |  | Procedure Not Followed | During the assessment of product count verification for NC-IL081-24588, it was found that product code 4502570, Lot# 20H03363X was being reworked for excessive silicone under NC-IL081-24579. This lot was previously on hold under NC-IL081-24579 for a failed spraymation determination by weight. \n\nNC-IL081-24579 concluded that there was no impact to product as the extraction test method confirmed that the product met both design and ISO specifications. This NCR was previously closed on 11/06/2020 with no product correction needed.\n\nAt the time of finding, the product was being 100% sorted for excessive silicone. \n\nBPCS hold number 20120006 has been issued for tracking and trending purposes of this failure. The hold number is tied to Product Code 4502570, Lot# 20H03363X |  |  | Rework of product was immediately stopped with a line clearance performed. | 04-Dec-2020 | 30-Nov-2020 | Procedure Not Followed | Procedure Not Followed | Human Error | The Production Team Lead did not verify the disposition in SmartSolve prior to reworking the product. |  |  | 0.0 | Closure | CLOSED | 03-Jan-2021 | 10-Dec-2020 | 5.
NC-IL081-27486 | 05-Dec-2020 | FIDELIA.CANTOR | General | Internal | C37 | Press 37-IL081 |  |  | 8836SA | 3GA PHARMASTAR - TRANS WHITE | CFF-V-Short | On 12/05/20 the start up failed the locking tab inspection for PT00096177 lids due to Shorts on locking tab on cav.2 and also found shorts on previous job 8836SA/ WP8836P lids (R7461E tool).\n\nThe PT00096177 lids will be reflected in this NC as the WP8836P lids due to a system glitch. | 20H18863' '20J194374 | 170.0 1340.0 | Increased shot size and velocity. Justin Pixius 12/07/2020 | 05-Dec-2020 | 05-Dec-2020 | CFF-V-Short | Part is not completely formed; Missing plastic | Shot Size |  | Accept' 'Reject | 120.0 734.0 50.0 606.0 | 0.0 | Closure | CLOSED | 04-Jan-2021 | 30-Dec-2020 | 24. | SS | Container
NC-IL081-27510 | 07-Dec-2020 | LAKEYSHA.SIMPSON | General | Internal | S8F | Form, Fill Sealer-IL081 |  |  | 8881540122 | 12ML SMARTIP MED-PREP CANNULA | SFF-PPSP-V-Seal Integrity 2 | Seal intact but less than 1/8inch width seal. | 20M02163X | 17760.0 | Cleared line of all sub-assembly packages,scrapped 37cs. Removed rotary cutter blades and reinstalled on correct side of the fixed blades. Par Myerson 12/07/20 | 07-Dec-2020 | 07-Dec-2020 | SFF-PPSP-V-Seal Integrity 2 | Seal is intact but less than 3 mm (1/8 inch) width seal. For 60 mL and 35 mL products, less than 5 mm width. | Rotary cutter blades installed on wrong side of fixed blade | Due to high scrap the rotary cutter blades were adjusted but when they installed them they were put on the incorrect side of the fixed blades. | Reject | 17760.0 | 0.0 | Closure | CLOSED | 05-Feb-2021 | 16-Dec-2020 | 8. | 8 | Syringe
NC-IL081-27534 | 08-Dec-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | QA-335, Pressure gage was out of tolerance as found. Gage was repaired (adjusted) and calibrated. Calibration passed. |  |  |  | 08-Dec-2020 | 01-Dec-2020 | Calibration Out of Tolerance | Calibration Out of Tolerance | Damaged measurement instrument | Damaged instrument causing measurement variation |  |  | 0.0 | Closure | CLOSED | 29-Jan-2021 | 17-Dec-2020 | 9.
NC-IL081-27589 | 10-Dec-2020 | BERTHA.HERNANDEZ | General | Internal | S7R | RAM-IL081 |  |  | 8881112985 | 10ML STERILE WATER SYR BULK NS | SFF-SA-V-Mixed Product 2 | Incomplete assembly; rubber tip was not properly seated | 20M01963X | 6000.0 | Line stopped and initiated checkback process | 10-Dec-2020 | 10-Dec-2020 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Push rods too high | Root Cause: It is possible that a slight variation in part length from the previous production run using a different barrel type caused the push down rods to not fully engage the rubber tips. | Reject | 6000.0 | 0.0 | Closure | CLOSED | 10-Feb-2021 | 16-Dec-2020 | 5. | 7 | Syringe
NC-IL081-27618 | 11-Dec-2020 | GRACE.AUSTIN | General | Internal | RSN | INCORRECT RESIN-IL081 |  |  | RM165 | Resin, Greenwood blend, Pre - colored Red | PR-- PR--oduct testing fails | Results on Batch History Record shows that Impact Testing failed. | 195493' '195180' '195184' '195284 | 22940.0 23100.0 23065.0 22860.0 | All affected resin has been quarantined. | 11-Dec-2020 | 11-Dec-2020 | PR-- PR--oduct testing fails | Product testing fails-MEDMFGPR | Incorrect mix of resin | Mix of resins may impact the properties of a component | Reject | 22940.0 23100.0 23065.0 22860.0 | 0.0 | Closure | CLOSED | 10-Apr-2021 | 17-Mar-2021 | 96. | Recy | Container
NC-IL081-27680 | 14-Dec-2020 | BENJAMIN.KELLY | General | Internal | S7R | RAM-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-SA-V-Mixed Product 2 | Customer received product not correctly assembled. Rubber tip is missing for the plunger. Complaint #: REJ # 21043 | 20E07963X | 1.0 |  | 14-Dec-2020 | 03-Dec-2020 | SFF-SA-V-Mixed Product 2 | Incomplete or incorrect assembly of product | Machine Failure | Machine Failure | Accept | 1.0 | 0.0 | Closure | CLOSED | 11-Feb-2021 | 12-Jan-2021 | 28. | 7 | Syringe
NC-IL081-27720 | 15-Dec-2020 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 4001204 | Smooth Rubber Tip with Trim, 12 mL, Latex Free, Kokoku | Procedural- Suspect Material | On 05/20/2020 and 06/22/2020, lots 20E15-1-12, 20F19-7-12, respectively, Kokoku plunger tips component code 4001204 were received at the Crystal Lake plant. Senior Incoming Quality Technician, FC, performed the required incoming quality testing and assessment for these tips. Upon this examination, FC observed an atypical visual appearance of the tips, similar to the spots remaining when water dries on a surface. The manufacturing process for these tips includes a wash/rinse in water by the tip manufacturer, however, this event identification occurred proximally to the introduction of a surfactant to the manufacturing process. As the subject lot of tips was not to have contained the surfactant, further investigation was conducted. It has been confirmed that no actual nonconformance exists. This SmartSolve NC record exists to document the investigation. | 20E15-1-12' '20F19-7-12 | 4320000.0 2160000.0 | All affected product placed on hold. | 15-Dec-2020 | 20-May-2020 | Procedural- Suspect Material | Procedural- Suspect Material | PR-Not Applicable | Not Applicable-MEDMFGPR\n\nIt has been confirmed that no actual nonconformance exists. SmartSolve NCR record exists to document investigation | Accept | 4320000.0 2160000.0 | 0.0 | Closure | CLOSED | 31-Jan-2021 | 16-Dec-2020 | 0. | Tips | Syringe
NC-IL081-27829 | 18-Dec-2020 | PATTI.SWISHER | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | Out of Spec Visual Inspection (Visual Defect) | Damaged product/component non functional | 20K21763X | 54.0 | Customer returns scar-20-015 | 18-Dec-2020 | 11-Dec-2020 | SFF-SA-V-Damaged Components 1 | Damaged Components Type 1: Damaged product/components (non-functional) | Machine Failure | Though a root cause could not be confirmed, the most probable cause resides with the rubber stopper and sensor positioning. | Reject | 54.0 | 0.0 | Closure | CLOSED | 15-Feb-2021 | 22-Jan-2021 | 34. | Sub | Syringe
NC-IL081-27836 | 18-Dec-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0499; OOT-AF; Adjusted back into tolerance. |  |  |  | 18-Dec-2020 | 18-Dec-2020 | Calibration Out of Tolerance | Calibration Out of Tolerance | Component/Parts | Component/Parts - The ultrasonic welders on the AM-171 assembly cell are obsolete and no longer supported by the manufacturer (Herrmann Ultrasonics). Replacement parts are no longer available. Excessive wear decreases the ability to maintain a calibrated state for the 6 month interval between calibrations. |  |  | 0.0 | Closure | CLOSED | 12-Feb-2021 | 14-Jan-2021 | 27.
NC-IL081-27837 | 18-Dec-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0500; OOT-AF; Adjusted back into tolerance. |  |  |  | 18-Dec-2020 | 18-Dec-2020 | Calibration Out of Tolerance | Calibration Out of Tolerance | Component/Parts | Component/Parts |  |  | 0.0 | Closure | CLOSED | 12-Feb-2021 | 14-Jan-2021 | 27.
NC-IL081-27838 | 18-Dec-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0501; OOT-AF; Unit required the replacement of the proportioning valve in order to be able to adjust back into tolerance; valve replaced by maintenance on W.O. CL-3268347; adjusted back into tolerance. |  |  |  | 18-Dec-2020 | 18-Dec-2020 | Calibration Out of Tolerance | Calibration Out of Tolerance | Component/Parts | Component/Parts - The Ultrasonic welders on the AM-171 assembly cell are obsolete and no longer supported by the manufacturer (Herrmann Ultrasonics). Replacement parts are no longer available. Excessive wear decreases the ability to maintain a calibrated state for the 6 month interval between calibrations. |  |  | 0.0 | Closure | CLOSED | 12-Feb-2021 | 14-Jan-2021 | 27.
NC-IL081-27839 | 18-Dec-2020 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0502; OOT-AF; Adjusted back into tolerance. |  |  |  | 18-Dec-2020 | 18-Dec-2020 | Calibration Out of Tolerance | Calibration Out of Tolerance | Component/Parts | Component/Parts - The ultrasonic welders on the AM-171 assembly cell are obsolete and no longer supported by the manufacturer (Herrmann Ultrasonics). Replacement parts are no longer available. Excessive wear decreases the ability to maintain a calibrated state for the 6 month interval between calibrations. |  |  | 0.0 | Closure | CLOSED | 12-Feb-2021 | 14-Jan-2021 | 27.
NC-IL081-27905 | 21-Dec-2020 | GRACE.AUSTIN | General | Internal | RSN | INCORRECT RESIN-IL081 |  |  | RM165 | Resin, Greenwood blend, Pre - colored Red | PR-- PR--oduct testing fails | Impact Testing Failed on Batch History Record. | 195184 | 23065.0 | Contained all affected product. | 21-Dec-2020 | 21-Dec-2020 | PR-- PR--oduct testing fails | Product testing fails-MEDMFGPR | Cause Unknown | Cause Unknown | Reject | 23065.0 | 0.0 | Closure | CLOSED | 31-Jan-2021 | 30-Dec-2020 | 9. | Recy | Container
NC-IL081-27906 | 21-Dec-2020 | FIDELIA.CANTOR | Receipt | Supplier |  |  |  |  | PT00007239 | Colorant, Red, Concentrate | Outside Validated Parameters | On 12/18/20 the supplier (CHROMA corporation) provided the COA (certificate of Analysis) with PALLET PER GRAM out of specification of ( 42) spec is 60-90 per RMPS (DC00007907 Rev. 0). | M1120486 | 9466.0 | Placed affected product on hold. | 21-Dec-2020 | 18-Dec-2020 | Procedural- Incorrect Information on COA/COC | Procedural- Incorrect Information on COA/COC | Process/Method | 1,2, & 5- Process/Method. Chroma is in the process of implementing a new ERP System. When the CIMPRO/BC Data transfer was done the pellet count did not get put on the BC data worksheet for distribution. 3- Pellet count was out of specification, as mentioned, the product had shipped already. 4- Customer service asked for a deviation since pellet count was out of specification, this was not accepted by the customer, CAH. | Reject | 9466.0 | 0.0 | Closure | CLOSED | 09-Apr-2021 | 29-Jan-2021 | 39.
NC-IL081-27934 | 22-Dec-2020 | GRACE.AUSTIN | General | Internal | C31 | Press 31-IL081 |  |  | 8990SA | 2GA HINGED LID SHARP CONT -RED | Molding Defect - Incorrect Resin | Incorrect Resin and PCR used for the R7028M Tool. RM093 Resin and PCR for R7028J Tool used for R7028M Tool from start-up. | 20L20663 | 2400.0 | Contain all affected product. | 22-Dec-2020 | 22-Dec-2020 | Molding Defect - Incorrect Resin | Molding Defect - Incorrect Resin | Human Error | Human Error | Reject | 2400.0 | 0.0 | Closure | CLOSED | 21-Jan-2021 | 07-Jan-2021 | 16. | G2 | Container
NC-IL081-28111 | 04-Jan-2021 | BEATRIZ.GUILLEN | General | Internal | RSN | INCORRECT RESIN-IL081 |  |  | RM165 | Resin, Greenwood blend, Pre - colored Red | PR-- PR--oduct testing fails | Impact testing failed on batch history record. | 00195284 | 22860.0 | Contained all affected Resin | 04-Jan-2021 | 04-Jan-2021 | PR-- PR--oduct testing fails | Product testing fails-MEDMFGPR | Cause Unknown | Cause Unknown; NC-IL081-28111, NC-IL081-27905, and NC-IL081-27618 will both be investigated and dispositioned under NC-IL081-27618. | Reject | 22860.0 | 0.0 | Closure | CLOSED | 03-Feb-2021 | 15-Jan-2021 | 10. | Recy | Container
NC-IL081-28182 | 06-Jan-2021 | BERTHA.HERNANDEZ | General | Internal | S2P | Packer-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | Documentation Incorrect | Incorrect component documentation, used component (3154020) produced in MP-1131 mold 58115 not listed on FS (FS200003060) parts were from Kanban. | 20M18763X | 30400.0 | All parts (barrels) were used, line down. | 06-Jan-2021 | 06-Jan-2021 | PR-Other | Other-MEDMFGPR | Other | There has been no process violation; however since the flexibility to use MP-1131 with tool 58115 is now needed, this NC is being used to capture the incident and add approved MSG/PCR PS20002614 to FS20003060. |  |  | 0.0 | Closure | CLOSED | 05-Feb-2021 | 14-Jan-2021 | 8. | 7 | Syringe
NC-IL081-28186 | 06-Jan-2021 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | CL63R1779 | Corrugated Case, Unprinted, 22.00 Inch x 14.00 Inch x 20.5 I | Visual- Assembly | On 01/06/21 manufacturing rejected 2 skids of product code CL63R1779 lot # 140134 PO # 824072 due to box had a gap on the top, causing the box not closing properly. | 140134 | 2410.0 | Placed affected product on hold. | 06-Jan-2021 | 06-Jan-2021 | Dimensional- out of specification | Dimensional- out of specification | Process/Method | Process/Method\nWhile in a stack, it makes it hard for the converting equipment operators to see the defect. After running the sheet through the machine, the end carton had a rough edge. | Reject | 2410.0 | 0.0 | Closure | CLOSED | 31-Mar-2021 | 10-Feb-2021 | 34.
NC-IL081-28217 | 07-Jan-2021 | BERTHA.HERNANDEZ | General | Internal | S2P | Packer-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-SA-V-Mixed Product 1 | Incorrect or mixed product, used rubber tips part#4006202, correct part#4101801 | 20M18763X | 26400.0 | Stopped and cleared the line. | 07-Jan-2021 | 07-Jan-2021 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Human Error | The employee in the tip tumbler did not place the rubber tips inside the correct color tote by accident because they were working on autopilot. | Accept | 26400.0 | 0.0 | Closure | CLOSED | 06-Feb-2021 | 08-Jan-2021 | 1. | 7 | Syringe
NC-IL081-28232 | 08-Jan-2021 | BENJAMIN.KELLY | General | Internal | S2B | 2B-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Nick / Gouge in the non-Fluid Pathway cavity 44B mold 58115 MP-1131 | 20M18763X | 6400.0 | stop line, perform checkback, placed product on Q A Hold | 07-Jan-2021 | 07-Jan-2021 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Machine Failure | Machine Failure | Accept | 6400.0 | 0.0 | Closure | CLOSED | 06-Feb-2021 | 11-Jan-2021 | 3. | 7 | Syringe
NC-IL081-28406 | 13-Jan-2021 | GRACE.AUSTIN | General | Internal | C82 | Press 82-IL081 |  |  | 31144010 | 5 QT TP SHARPS CONT -TRANS RED | CFF-V-Label Verification and Barcode Verification 2 | Product code number incomplete and missing print on Biohazard labels. | 20K07363 | 4050.0 | All affected product contained. | 13-Jan-2021 | 13-Jan-2021 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Incorrect Roll Alignment | Incorrect Roll Alignment- Improper threading of ribbon causes wrinkles which does not allow for consistent ink transfer to label. | Accept | 4050.0 | 0.0 | Closure | CLOSED | 12-Feb-2021 | 15-Jan-2021 | 1. | G5 | Container
NC-IL081-28498 | 16-Jan-2021 | ALEJANDRA.CASTANEDA | General | Internal | C46 | Press 46-IL081 |  |  | 8536SA | 3 GA SHARPSTAR - CLEAR | CFF-F-WELDING | Manufacturing reported bad welding on the front of the lids, failed weld strength. Defect is visually identifiable, functional testing confirmed failure. | 20H25163' '20K15163 | 2090.0 4320.0 | Line went down and QA initiated the check back to quarantine the affected product.\n\nThe following actions were taken to return the line to manufacturing:\n1) Lift station alignment pins on back left corner and front right corner were replaced under Maximo work order CL-3276402. \n2) Startup AQL was conducted post repairs with passing results. See attachment titled, "Restart AQL"\n3) Production changed jobs and tightened inspection was implemented with no further failures identified after running for 24 hours at the tightened inspection level. | 16-Jan-2021 | 15-Jan-2021 | CFF-F-Welding | Weld Strength | Procedures: UnClear |  | Accept | 2090.0 4320.0 | 0.0 | Closure | CLOSED | 15-Feb-2021 | 05-Feb-2021 | 20. | SS | Container
NC-IL081-28591 | 19-Jan-2021 | PATTI.SWISHER | General | Internal | S4A | 412A-IL081 |  |  | R7642B | Barrel, 12 mL | SFF-M-V-Flash/Gate 1 | Flash in the fluid pathway cavity A8 | 21A05063 | 2000.0 | Line was stopped, checkback performed, product placed on QA Hold | 19-Jan-2021 | 15-Jan-2021 | SFF-M-V-Flash/Gate 1 | Flash Type 1: Flash in Fluid Path (Luer Tip) exceeds 1/64” (.016”). Does not meets drawing specification | Tooling/ fixture wear | Wear on tooling causes variation | Reject | 2000.0 | 0.0 | Closure | CLOSED | 18-Feb-2021 | 22-Jan-2021 | 3. | 8 | Syringe
NC-IL081-28653 | 20-Jan-2021 | JACLYN.GRISKO | General | Internal | RSN | INCORRECT RESIN-IL081 |  |  | 0301009 | Resin, Clarified Polypropylene, Grade P5M6K-080 | V-28 Incorrect resin mix | During receipt of RM093, the BOL was mislabeled and was dumped on top of the P5M6K in silo 10. | Mixed - Surge' 'Mixed - Production Lines' 'Mixed - Silo' '2101104' '20M18463X' '20M12563X' '21A04763X | 2247.0 110.0 152380.0 46740.0 246000.0 108480.0 192000.0 | Silo 10 (P5M6K - p/n 03010009) feeds production in the Syringe Focus Factory. Production on the lines running P5M6K material was immediately stopped. | 20-Jan-2021 | 19-Jan-2021 | V-28 Incorrect resin mix | V-28 Incorrect resin mix | Procedures: UnClear | Procedures: UnClear | Reject' 'Accept | 110.0 152380.0 2247.0 4800.0 246000.0 108480.0 46740.0 187200.0 | 0.0 | Closure | CLOSED | 06-Mar-2021 | 25-Feb-2021 | 35.
NC-IL081-28827 | 26-Jan-2021 | LAKEYSHA.SIMPSON | General | Internal | S1P | Packer-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-SA-V-Mixed Product 1 | Incorrect or mixed product. Used rubber tips part\n # 4101801, correct part # 4006202 | 21A09063X | 30400.0 | Line cleared and lot was cut short. | 25-Jan-2021 | 25-Jan-2021 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Tips in incorrect totes | 4006202 part is the correct rubber tip for product code 8881112599. Quality observed mixed tips with 4101801 in the finish cases. No other totes of 4006202 were observed on the floor or in the tip room to be mixed with the incorrect part. Most probable root cause would be someone had changed over to part number 4101801 in the tip room but failed to switch over to the correct colored tote. | Reject | 30400.0 | 0.0 | Closure | CLOSED | 25-Mar-2021 | 17-Feb-2021 | 22. | 1 | Syringe
NC-IL081-28933 | 28-Jan-2021 | JACLYN.GRISKO | General | Internal | QA | QA-IL081 |  |  |  |  | PR-Procedures: Not Clear | After the calendar change from 2020 to 2021, the PGII certification year was not update on the label template for the following product codes: 8602RC, 8602U, 8985PG2Y, 8985PG2, 8997PG2, 8607RC, 8607U, 8608RC, 8608U, 8997SPG2, 8985SPG2, 8931PG2, 8932PG2, 8611RC, 8611U, 8931PG2Y, 8934PG2, 8936PG2, 8612RC, 8612U, 8998PG2, 8989PG2, 8617RC, 8617U, 8918GSTC, 8871, 8871NB, 8998SPG2, 8939PG2, 8618RC, V8618, 8618U, and 8930SA.\n\nBPCS hold number 21010026 opened under 8602RC, Lot# 20J19363 for trending purposes only. |  |  | "Immediate containment of product is not required because although the label template was not updated, the site's PGII certification had an expiration date of 01/31/21 which encompassed the dates of production from 01/04/2021 (first day on site after holiday) till 01/31/21." | 04-Jan-2021 | 04-Jan-2021 | PR-Procedures: Not Correct | Procedures: Not Correct-MEDMFGPR | PR-Documentation | Documentation-MEDMFGPR |  |  | 0.0 | Closure | CLOSED | 27-Feb-2021 | 18-Feb-2021 | 20.
NC-IL081-28949 | 29-Jan-2021 | BENJAMIN.KELLY | General | Internal | S7N | Pad Printer-IL081 |  |  | 4501820 | 12mL Syringe Oral Purple Subs | SFF-P-V-Alpha Numeric 4 | Graduation line and alpha character printing on Syringe smeared to the point of being illegible. | 21A16063X | 2800.0 | Stopped and cleared line of all affected product. | 28-Jan-2021 | 28-Jan-2021 | SFF-P-V-Alpha Numeric 4 | Type 3: Print is smeared to the point of being illegible | Worn area on pad | There was a worn spot on the pad which was causing print to smear. | Accept | 2800.0 | 0.0 | Closure | CLOSED | 26-Mar-2021 | 01-Feb-2021 | 3. | 7 | Syringe
NC-IL081-28963 | 29-Jan-2021 | ANGELICA.AGUILAR | General | Internal | CME | CME-IL081 |  |  |  |  | Environmental issue | 0.5um, non-viable environmental data collected between June 2020 and January 2021 has demonstrated a trend for out of specification conditions at locations North 45, North 46, Packaging 16, 19, 20, 21, 23, 24 and 27. |  |  | There is no potential product impacted by this non-conformance as the monitoring of product through in process functional checks and quarterly dose audit checks have all demonstrated acceptable results for each product manufactured and each respective product family. | 29-Jan-2021 | 01-Jun-2020 | Environmental Particulate (OOS) | Environmental Particulate (OOS) | Environmental monitoring parameters exceeded limits | Environmental monitoring parameters exceeded limits |  |  | 0.0 | Closure | CLOSED | 17-Mar-2021 | 22-Feb-2021 | 23.
NC-IL081-28977 | 30-Jan-2021 | BENJAMIN.KELLY | General | Internal | S6A | 6A-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Deformation 1 | Molding Deformation = Pulled ring stop Cav. #44 | 21A09463X | 53000.0 | Line cleared of all impacted product. Cav. # 44 closed. | 29-Jan-2021 | 29-Jan-2021 | SFF-M-V-Deformation 1 | The ring stop has been pulled from the inside dimension of the barrel possibly rendering the barrel unusable | Machine Failure | Machine Failure- Root cause: The most probable root cause was not getting enough cooling to the mold cavity due to a blocked bubbler.\n\nCorrection: The cavity was blocked and work order #CL-3282907 written to regain cavitation. | Reject | 53000.0 | 0.0 | Closure | CLOSED | 26-Feb-2021 | 23-Feb-2021 | 24. | Sub | Syringe
NC-IL081-29007 | 01-Feb-2021 | LAKEYSHA.SIMPSON | General | Internal | S1R | RAM-IL081 |  |  | 8881112599 | 12ML LLK SMOOTH TIP BULK NS | SFF-SA-V-MIXED PRODUCT 2 | Incomplete or incorrect assembly of product. Missing rubber tip. | 21A15363X | 19200.0 | stop line, perform checkback, place product on QA Hold. | 01-Feb-2021 | 30-Jan-2021 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Process Not Robust | Line 1 RAM was reviewed for potential causes. This RAM consists of two sets of spraymation heads within an enclosed space. The additional spraymation heads compared to the other RAMS creates additional buildup within the system. The process for silicone application was reviewed and found to not be optimized for residual buildup. The residual buildup will decrease the sensitivity of the sensors if not cleaned. | Accept | 19200.0 | 0.0 | Closure | CLOSED | 02-Apr-2021 | 11-Mar-2021 | 38. | 1 | Syringe
NC-IL081-29044 | 02-Feb-2021 | MARIA.ROSALES | General | Internal | C42 | Press 42-IL081 |  |  | 8525Y | 3 GA ROTOR/HINGE LID-YELLOW 10 | Approved Documents Incorrect | Used Incorrect PCR and MSG for trans yellow container Doc.# 130720-2712 should be Doc.# 130814-2803 | 21A22763 | 3440.0 | Performed EOR. | 02-Feb-2021 | 02-Feb-2021 | Approved Documents Incorrect | Approved Documents Incorrect | Human Error | The database for the specified processes lists multiple options and the clerk did not catch the error. | Accept | 3440.0 | 0.0 | Closure | CLOSED | 04-Mar-2021 | 18-Mar-2021 | 44. | G3 | Container
NC-IL081-29184 | 05-Feb-2021 | BERTHA.HERNANDEZ | General | Internal | S1B | 1B-IL081 |  |  | 6492743 | 12ML LUER LOCK SYRINGE, PRINTED | SFF-M-V-Deformation 1 | Deformation pulled ring stop on cav#46 | 20M08163X' '21A22963X | 44000.0 52000.0 | Stopped and cleared the line, placed product on hold. | 05-Feb-2021 | 05-Feb-2021 | SFF-M-V-Deformation 1 | The ring stop has been pulled from the inside dimension of the barrel possibly rendering the barrel unusable | Tooling/ fixture damaged | The most probable root cause was not getting enough cooling to the mold cavity due to a blocked bubbler. | Accept | 44000.0 52000.0 | 0.0 | Closure | CLOSED | 04-Aug-2021 | 09-Jul-2021 | 154. | 7 | Syringe
NC-IL081-29207 | 06-Feb-2021 | FIDELIA.CANTOR | General | Internal | C83 | Press 83-IL081 |  |  | 8950SA | 5 QT ROUND ROTOR LID - RED | Packaging - Incorrect Count | Missing lids and incorrect packaging, not packed per standard work DC00077004 rev. 1 . | 20L14763 | 800.0 | Placed affected product on hold. | 06-Feb-2021 | 06-Feb-2021 | Packaging - Incorrect Count | Packaging - Incorrect Count | Human Error | Human Error | Accept' 'Reject | 760.0 40.0 | 0.0 | Closure | CLOSED | 08-Mar-2021 | 10-Feb-2021 | 4. | G5 | Container
NC-IL081-29292 | 10-Feb-2021 | AUGUSTINA.MASH | General | Internal | S6K | Kahle-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Particulate 3 | Extraneous matter in fluid pathway | 21B05563X' '21A15963X | 198000.0 41000.0 | Line stopped and cleared of all impacted product | 09-Feb-2021 | 09-Feb-2021 | SFF-SA-V-Particulate 3 | Particulate Type 3: ‘Enteral Oral’ Syringes ONLY: Extraneous matter in fluid pathway. (Does not apply to Enfit Enteral Syringe with external luer threads | Conveyor belt rubbing | The belt on the conveyor going into the barrel bowl was not staying in the track causing it to move side to side causing rubbing and degrading. Small bits of the material were getting into the syringes. Lack of PM. | Reject | 198000.0 41000.0 | 0.0 | Closure | CLOSED | 30-Apr-2021 | 12-Mar-2021 | 30. | Sub | Syringe
NC-IL081-29350 | 11-Feb-2021 | BERTHA.HERNANDEZ | General | Internal | CME | CME-IL081 |  |  | 5551731218 | PLNGR 12CC CLEAR | SFF-SPBP-V-Labeling/Print 6 | Failed barcode verification unable to grade overshipper labels. | 21B10563X | 63.0 | Placed product on hold. | 11-Feb-2021 | 10-Feb-2021 | SFF-SPBP-V-Labeling/Print 6 | Fails UDI Grade Test | Illegible Labeling | Illegible Labeling | Accept | 63.0 | 0.0 | Closure | CLOSED | 13-Mar-2021 | 23-Feb-2021 | 11. | 7 | Syringe
NC-IL081-29371 | 11-Feb-2021 | JACLYN.GRISKO | General | Internal | S4N | Hot Stamp Printer-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | Supplier Certification | The Pick and Place for the Line 412 printer was bent beyond repair from a crash/jam. This requires replacement of the unit. The supplier for the replacement is not currently on the GASL (Global Approved Supplier List). This NCR will be used to capture the purchase of the replacement while the supplier is added to the GASL.\n\nShop Order 195769 (8881412012, 21A23363X) is the first scheduled shop order for line 412 which would require print. | 21A23363X | 120000.0 | Production line went immediately down at time of event. Suppliers were sought and the vendor that could supply the replacement was found to not be on the GASL. GASL paperwork to start the approval process to add them to the GASL was started. | 11-Feb-2021 | 11-Feb-2021 | Supplier Certification | Supplier Certification | Procedure/Method | Procedure/Method | Accept | 120000.0 | 0.0 | Closure | CLOSED | 26-Mar-2021 | 23-Mar-2021 | 39. | 412 | Syringe
NC-IL081-29393 | 12-Feb-2021 | BERTHA.HERNANDEZ | General | Internal | S8B | 8B-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | Out of Validated Parameter | Charge recovery time documented out of spec, QA confirmed parameter out of spec. 3.77 spec is 3.8-8.2 | 21B06763X | 45600.0 | Press was put back to validated parameters, affected product was placed on hold. | 12-Feb-2021 | 11-Feb-2021 | Out of Validated Parameter | Out of Validated Parameter | Equipment variation requiring maintenance adjustments | MP-1133 (8B) machine log was reviewed and showed that the charge recovery started fluctuating going in and out of spec starting on the shift when a Process Technician closed a cavity. Adjustments were made to screw speed for the blocking of the cavity, but the technician did not look at the recovery time when they were done. If the screw recovery time wasn’t adjusted for this then this could cause the screw recovery time to drop slightly. | Accept | 45600.0 | 0.0 | Closure | CLOSED | 14-Mar-2021 | 23-Mar-2021 | 39. | 8 | Syringe
NC-IL081-29483 | 15-Feb-2021 | FIDELIA.CANTOR | Receipt | Supplier |  |  |  |  | 4101801 | TIP 12 CC RAYCO WG2714 RUBBER TIP | Visual- Water in package/on tip | During incoming inspection on 02/15/21 was discovered that 2cs(2 bags) of lot numbers YL211120-032 box # 3 and YL251120-022 box # 2 had water(condensation) in the bags, AQL for presence of water in packaging is N-II .065 A-0 R-1 per GS20003456. | YL211120-032-YL251120-022 | 189000.0 | Placed product on hold. | 15-Feb-2021 | 15-Feb-2021 | [6 empty] | 0.0 | Initiation Verification | VOID |  |  |  | Tips | Syringe
NC-IL081-29488 | 16-Feb-2021 | MONA.ABOKAUID | General | Internal | S8N | Pad Printer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SPIC-V-Labeling/Print 1 | Incorrect Lot Number on soft pack \nwrong # 20B06863X CORRECT # 21B06863X | 21B06863X | 46080.0 | Stop line, place product on QA Hold | 16-Feb-2021 | 15-Feb-2021 | SFF-SPIC-V-Labeling/Print 1 | Lot Number and Expiration Date (where applicable) is incomplete or missing | Human Error | Human Error | Reject | 46080.0 | 0.0 | Closure | CLOSED | 18-Mar-2021 | 22-Feb-2021 | 6. | 8 | Syringe
NC-IL081-29496 | 16-Feb-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | RC-012 Calibration out of tolerance |  |  |  | 16-Feb-2021 | 16-Feb-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Facilities/Equipment | Facilities/Equipment\nThe investigation into the cause of the load cell failures leads back to an issue with the chiller system that shares the same electrical feed as the extruders. Work order CL-3268736 was opened to document that one compressor had shorted to ground in late December 2020 causing back fed amperage that damaged the two load cells on RC-1 causing them to fail. |  |  | 0.0 | Closure | CLOSED | 31-Aug-2021 | 07-Sep-2021 | 202.
NC-IL081-29506 | 17-Feb-2021 | MONA.ABOKAUID | General | Internal | S4A | 412A-IL081 |  |  |  |  | Documentation Missing | Failure to document PCR parameters once per shift found during the line audit from R7642B, Lot# 21A09163. The process was reviewed at time of finding and found to be within parameters. The Sr. Process technician reviewed the change history log on the press and confirmed that no changes were made to the process during the shift that was missing the PCR record.\n\nBPCS hold number 21020024 initiated for trending purposes only. |  |  | Process parameters documented on PCR. | 17-Feb-2021 | 16-Feb-2021 | Documentation Missing |  | Human Error | Human Error |  |  | 0.0 | Closure | CLOSED | 19-Mar-2021 | 22-Feb-2021 | 5.
NC-IL081-29515 | 17-Feb-2021 | BERTHA.HERNANDEZ | General | Internal | S7P | Packer-IL081 |  |  | 8881112059 | 12ML SYRINGE REGULAR NS | SFF-SA-V-Mixed Product 1 | Incorrect or mixed product, traceability sheet shows a record of 3154020 (LK barrel) being used, correct part# PT00050783 (SR barrel). | 21B06363X | 144000.0 | Placed product on hold. | 17-Feb-2021 | 17-Feb-2021 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Procedures: UnClear | Traceability procedure GS10000105 was reviewed and a gap was identified. The procedure did not outline that the tote tickets needed to be removed upon consumption of components from brutes. | Reject' 'Accept | 500.0 143500.0 | 0.0 | Closure | CLOSED | 18-May-2021 | 21-Jun-2021 | 124. | 7 | Syringe
NC-IL081-29652 | 23-Feb-2021 | BERTHA.HERNANDEZ | General | Internal | S8A | 8A-IL081 |  |  |  |  | Procedure Not Followed | Decoupled process from press 8A directly fed into Finished Good shop order# 196203 unable to meet section 5.4 of GS1000122 revision 210. \n\nBPCS hold number 21020035 issued towards 4502570, 21B18463X for trending purposes. |  |  | Finished the lot and performed an EOR. Part number and PCR of the components from 8A reviewed against the BOM/Design Specification for 4502570. The components produced from 8A are an approved component and process per the BOM/Design Specification. | 23-Feb-2021 | 23-Feb-2021 | PR-Procedure Not Followed | The most likely root cause is the commodity change which may have caused the individual to assume that the parts could be directly fed. The in-line conveyor was not physically barricaded to prevent the in-line feed which could have caused the individual to pause to confirm if parts could be fed. | Procedure not followed | Procedure not followed |  |  | 0.0 | Closure | CLOSED | 25-Mar-2021 | 24-Feb-2021 | 1.
NC-IL081-29754 | 27-Feb-2021 | BENJAMIN.KELLY | General | Internal |  |  |  |  | 8881907199 | 10ML ORAL DOSE SYR UPT STER | Incorrect Raw Marterial Used / Issued | Syringe Header bag equipment was found with 16-8700Q (ink) & 16-8705Q (make up fluid) loaded in the machine. 8881907199 only lists EXP6026 ink on BOM. No production was running at the time of identification. | 19H23563 | 1.0 | Ink and make up fluid removed from machine. | 26-Feb-2021 | 26-Feb-2021 | Incorrect Raw Marterial Used / Issued | Incorrect Raw Marterial Used / Issued | Process/Method | Root Cause (EXP6021): Since EC018171 released 8/31/15 which was after the corrective action plan for ACT000003-AUD-005930 was submitted, the items were missed. | Accept | 1.0 | 0.0 | Closure | CLOSED | 26-Jun-2021 | 18-Jun-2021 | 111. | 7 | Syringe
NC-IL081-29829 | 02-Mar-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MM-041, pressure gage. Calibration out of tolerance as found; unit could not be adjusted. |  |  | Unit removed from service. | 02-Mar-2021 | 02-Feb-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Handling/Storage | Handling/Storage |  |  | 0.0 | Closure | CLOSED | 27-Aug-2021 | 18-Aug-2021 | 169.
NC-IL081-29924 | 06-Mar-2021 | BENJAMIN.KELLY | General | Internal | S6A | 6A-IL081 | Equipment ID | 58110-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 1 | Flash in fluid pathway exceeding 1/64 Cav #17 | 21B36063X | 4000.0 | Line stopped and cleared. Check back procedure began. | 05-Mar-2021 | 05-Mar-2021 | SFF-M-V-Flash/Gate 1 | Flash Type 1: Flash in Fluid Path (Luer Tip) exceeds 1/64” (.016”). Does not meets drawing specification | Tool damage | Plastic behind the insert cavity #17 | Reject | 4000.0 | 0.0 | Closure | CLOSED | 05-May-2021 | 12-Mar-2021 | 6. | Sub | Syringe
NC-IL081-29966 | 08-Mar-2021 | GRACE.AUSTIN | General | Internal | C19 | Press 19-IL081 |  |  | 8507SA | CONTAINER SHARPS 5QT RED | CFF-V-Short | Short on the rim of container, cavity #1, found during QA final case inspection. | 20M13263 | 6060.0 | Contain all affected product on Hold | 08-Mar-2021 | 08-Mar-2021 | CFF-V-Short | Part is not completely formed; Missing plastic | Mold Wear or Damage | The most likely root cause is that there is damage to the mold which was causing flash. In order to eliminate the flash, the Process Technician made adjustments to the fill pressure, resulting in shorts. | Reject' 'Accept | 180.0 5880.0 | 0.0 | Closure | CLOSED | 07-Apr-2021 | 29-Mar-2021 | 21. | SS | Container
NC-IL081-29983 | 08-Mar-2021 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | RM175 | Resin, Linear Low Density Polyethylene LL3001 | Packaging Damaged | While unloading the gaylords of RM175 resin, lot # 062012045A, from Channel Prime, on 3/08/21 the driver discovered that one gaylord was leaking resin due to two holes in the gaylord. | 062012045A | 24000.0 | Placed the resin on hold. | 08-Mar-2021 | 08-Mar-2021 | Visual- Shipping Damage | Visual- Shipping Damage-- While unloading the gaylords of RM175 resin, lot # 062012045A, from Channel Prime, on 3/08/21 the driver discovered that one gaylord was leaking resin due to two holes in the gaylord. | PR-Unknown | Unknown-MEDMFGPR. Supplier could not determine root cause with the 3rd party carrier. The shipping warehouse confirmed the material was in good shipping condition with no reported damage at the time of loading the material onto the delivery carriers truck. The Ravago Logistics teamed confirmed the driver did not handle (load or unload) the material in question. The driver delivered with no incidences reported. \n\nThe Driver did sign the BOL (attached) documenting the damage to the gaylord upon receipt at CAH Crystal Lake. | Reject | 24000.0 | 0.0 | Closure | CLOSED | 30-Apr-2021 | 26-Mar-2021 | 17.
NC-IL081-30058 | 10-Mar-2021 | MARIA.ROSALES | General | Internal | C81 | Press 81-IL081 |  |  | 89671 | 2 GA HORIZONTAL DROP - T RED | CFF-D-Molding 1 | Manufacturing found during scheduled inspection, wall thickness out specification on both cavities of containers. \nCavity 1-0.052" and cavity 2-0.053" . Minimum wall thickness specification should be 0.055". | 21A08363' '21C21163 | 22800.0 14100.0 | Technician shut the press down to make adjustment and all affected product has been quarantined. | 10-Mar-2021 | 10-Mar-2021 | CFF-D-Molding 1 | Wall Thickness | Mold Wear or Damage | Mold Wear or Damage | Reject' 'Accept | 420.0 22380.0 13680.0 | 0.0 | Closure | CLOSED | 09-May-2021 | 05-Apr-2021 | 26. | G2 | Container
NC-IL081-30059 | 10-Mar-2021 | MONA.ABOKAUID | General | Internal | S4A | 412A-IL081 |  |  | 5551771405 | 412 BBL STR UPT CLR | SFF-M-V-Deformation 2 | Barrel shoulder from cavity# 13B is deformed. Sunken shoulder. | 21B30063X | 4000.0 | Stopped line and closed the cavity. | 10-Mar-2021 | 10-Mar-2021 | SFF-M-V-Deformation 2 | The dimension of the barrel end is different than specification | High Mold Temp | The thermolator was accidently shut off which caused low water in the DME resulting in failure to cool parts prior to ejection. | Reject | 4000.0 | 0.0 | Closure | CLOSED | 09-Apr-2021 | 12-Mar-2021 | 2. | 412 | Syringe
NC-IL081-30082 | 11-Mar-2021 | AUGUSTINA.MASH | General | Internal | S6A | 6A-IL081 | Equipment ID | 58110-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 2 | Flash in functional area (luer skirt) exceeds 1/32 (.031") Major C/0.40. Mold 58110, Cavities 7, 15, 23 ,31 & 39 | 21C03163X | 40000.0 | Line was stopped and cleared; Tool pulled for evaluation. | 10-Mar-2021 | 10-Mar-2021 | Out of Spec Visual Inspection (Visual Defect) | Out of Spec Visual Inspection (Visual Defect) | Bolt came out | The bolt on the under side of the rack for the core pull came out, causing rack to not be able to go down completely. | Reject | 40000.0 | 0.0 | Closure | CLOSED | 10-May-2021 | 12-Mar-2021 | 1. | Sub | Syringe
NC-IL081-30194 | 15-Mar-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | FIPC-05 (8 Gallon F.I.P. Height Gage) could not be found for calibration. Assigned to Containers Production. |  |  |  | 15-Mar-2021 | 11-Mar-2021 | Calibration Lost Gauge |  | Calibration/Equip.Verification |  |  |  | 0.0 | Closure | CLOSED | 19-Apr-2021 | 13-Apr-2021 | 29.
NC-IL081-30195 | 15-Mar-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0512 (.125 GAGE BLOCK) Lost. Assigned to Containers Production. |  |  |  | 15-Mar-2021 | 11-Mar-2021 | [6 empty] | 0.0 | Investigation | VOID
NC-IL081-30224 | 16-Mar-2021 | BERTHA.HERNANDEZ | General | Internal | S8N | Pad Printer-IL081 | Equipment ID | 58119-IL081 | 8881112015 | 12ML SYRINGE PURPLE EN-FIT TIP | SFF-M-V-Deformation 1 | Deformation, the ring stop has been pulled from the inside dimension of the barrel possible rendering the barrel unusable, cavities A2,A7 and A8 | 21C23963' '21A09163 | 4000.0 140000.0 | Stopped and cleared the line, placed product on hold. | 16-Mar-2021 | 11-Feb-2021 | SFF-M-V-Deformation 1 | The ring stop has been pulled from the inside dimension of the barrel possibly rendering the barrel unusable | un-balanced mold cavitation | The most probable root cause is that there were too many cavities blocked on an 8 cavity mold affecting cooling. | Reject' 'Accept | 4000.0 112000.0 28000.0 | 0.0 | Closure | CLOSED | 30-Apr-2021 | 23-Mar-2021 | 7. | 8 | Syringe
NC-IL081-30247 | 17-Mar-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Past Due | QA-0050; CALIBRATION PAST DUE; CALIBRATION WAS DUE 03/01/2021 AND OCCURRED 03/02/2021. AN NCR WAS GENERATED FOR THIS EVENT PER GS10000123; OUT-OF-TOLERANCE, AS-FOUND; AND OOT-AF WITHIN LAST TWO PREVIOUS PERIODS. |  |  |  | 17-Mar-2021 | 02-Mar-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Tooling/ fixture wear | Wear on tooling causes variation |  |  | 0.0 | Closure | CLOSED | 16-Apr-2021 | 20-Apr-2021 | 34.
NC-IL081-30301 | 19-Mar-2021 | ALEJANDRA.CASTANEDA | General | Internal | C70 | Press 70-IL081 |  |  | 8900SA | 1 QT PHLEBOTOMY CONT | CFF-V-Label Verification and Barcode Verification 2 | Incorrect lot number printed on biohazard label: 21A29063\nCorrect lot number should be: 21A22063 | 21A22063 | 5000.0 | Contain effected product. | 19-Mar-2021 | 19-Mar-2021 | CFF-V-Label Verification and Barcode Verification 2 | BH - Bad/Missing label print | Human Error | Human Error | Reject' 'Accept | 3400.0 1600.0 | 0.0 | Closure | CLOSED | 18-Apr-2021 | 08-Apr-2021 | 20. | G1 | Container
NC-IL081-30316 | 19-Mar-2021 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | 5170588 | Poly Bag, Gusseted, 22.5 Inch x 16.25 Inch x 42.00 Inch | Visual- Contamination (not for rubber tips) | On 3/19/21 warehouse personnel discovered that 3cs of 5170588 lot # UP90021517 PO # 824318 were received damaged (open), and bags have black specs of a foreign matter on them. AQL is S-4, 0.065, A-0 R-1. | UP90021517 | 675.0 | Placed affected product on hold. | 19-Mar-2021 | 19-Mar-2021 | Visual- Contamination (not for rubber tips) | Visual- Contamination (not for rubber tips) | PR-Unknown | Unknown-MEDMFGPR. Root cause documented in CAPA-IL081-02219 | Reject | 675.0 | 0.0 | Closure | CLOSED | 04-Jun-2021 | 26-Apr-2021 | 38. | IQA | Incoming
NC-IL081-30352 | 22-Mar-2021 | BEATRIZ.GUILLEN | General | Internal | C52 | Press 52-IL081 |  |  | 8989 | 18 GA HINGED CHEMO - YELLOW | CFF-V-Blisters, Bubbles and Voids | Manufacturing found during unscheduled inspection, blister out of spec. on both sides of container "A" surface. | 21A03163 | 220.0 | Performed End of Run and placed product on hold. | 22-Mar-2021 | 22-Mar-2021 | CFF-V-Blisters, Bubbles and Voids | Single bubble greater than or equal to 0.080”; Up to 3 bubbles between 0.060 and 0.080”; Up to 4 bubbles less than or equal to 0.060”; Up to 8 bubbles less than 0.030”; Living Hinge = Reject All | incorrect tooling/ fixture | The air was being trapped in certain positions as the material had uneven flow due to out of spec gate inserts. | Reject | 220.0 | 0.0 | Closure | CLOSED | 26-Apr-2021 | 19-Apr-2021 | 28. | G18 | Container
NC-IL081-30358 | 22-Mar-2021 | LAKEYSHA.SIMPSON | General | Internal | [6 empty] | I | Spider found on the floor. | 19K20963X | 1.0 |  | 22-Mar-2021 | 22-Mar-2021 | [6 empty] | 0.0 | Initiation Verification | VOID
NC-IL081-30385 | 23-Mar-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0693; Lost gage; could not be located during calibration rounds on 03/19/2021. |  |  |  | 23-Mar-2021 | 19-Mar-2021 | Calibration Lost Gauge | Calibration Lost Gauge | Calibration/Equip.Verification | Calibration/Equip.Verification |  |  | 0.0 | Closure | CLOSED | 22-Apr-2021 | 12-Apr-2021 | 19.
NC-IL081-30393 | 23-Mar-2021 | GRACE.AUSTIN | General | Internal | 171 | Cell 171-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0499, Welder 1 out of tolerance as found. | 1557405 | 1.0 | Welder adjusted to return to tolerance. | 23-Mar-2021 | 23-Mar-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Root Cause will be further defined as part of CAPA-IL081-02251. |  |  | 0.0 | Closure | CLOSED | 27-Apr-2021 | 21-Apr-2021 | 28.
NC-IL081-30394 | 23-Mar-2021 | ANGELICA.AGUILAR | General | Internal | 171 | Cell 171-IL081 |  |  | NA - NOT PRODUCT SPECIFIC | NA- Not Product Specific | Calibration Expired | Gauge was found out of tolerance in AM171 Cell. | System Serial SM7.274/05 | 1.0 | Gauge was adjusted into tolerance. | 23-Mar-2021 | 09-Mar-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 27-Apr-2021 | 21-Apr-2021 | 28.
NC-IL081-30395 | 23-Mar-2021 | ANGELICA.AGUILAR | General | Internal | 171 | Cell 171-IL081 |  |  | NA - NOT PRODUCT SPECIFIC | NA- Not Product Specific | Calibration Expired | MD-501, Welder #3 was found out of calibration in AM171 Cell. | SM6713/04 | 1.0 | Gauge was adjusted into tolerance. | 23-Mar-2021 | 09-Mar-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 27-Apr-2021 | 21-Apr-2021 | 28.
NC-IL081-30396 | 23-Mar-2021 | GRACE.AUSTIN | General | Internal | 171 | Cell 171-IL081 |  |  | NA - NOT PRODUCT SPECIFIC | NA- Not Product Specific | Calibration Out of Tolerance | MD-0502; Welder was found out of tolerance in AM171 Cell. | 1985781 | 1.0 | Welder adjusted to return to tolerance. | 23-Mar-2021 | 23-Mar-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 27-Apr-2021 | 21-Apr-2021 | 28.
NC-IL081-30415 | 24-Mar-2021 | MONA.ABOKAUID | General | Internal | S6C | 6C-IL081 | Equipment ID | 58114-IL081 | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Scratch /Drag mark in Non-Fluid pathway ,can be felt with a bare finger . cavity 23B MP-1136 mold 58114 | 21C29763X | 16000.0 | stop line, clear , perform checkback, place product on QA Hold | 23-Mar-2021 | 23-Mar-2021 | SFF-M-V-Drag Marks/Gouges/Nicks Scratches 2 | Drag Marks/Gouges/Nicks/Scratches Type 2/3: Scratch/Drag Marks in Non-Fluid Pathway - Material removed is measurable. If the scratch/scuff/drag mark can be felt with a bare finger, it will be considered too deep. | Bad gate | Gate damaged on cavity 23B plastic caught in the gate causing scratch. | Reject | 16000.0 | 0.0 | Closure | CLOSED | 21-May-2021 | 12-Apr-2021 | 19. | Sub | Syringe
NC-IL081-30459 | 24-Mar-2021 | ANGELICA.AGUILAR | General | Internal | 131 | Cell 131-IL081 |  |  | 8506SA | CONTAINER BIOHAZARD 5 QT | CFF-F-Welding | Lids with bad weld and short was found during QA scheduled inspection. | 21C08863 | 440.0 | Cell went down to make adjustments and the End of Run was performed. All affected product was placed on hold. | 24-Mar-2021 | 24-Mar-2021 | CFF-F-Welding | Weld Strength | Misaligned assembly transfer | Misaligned assembly transfer | Accept | 440.0 | 0.0 | Closure | CLOSED | 23-Apr-2021 | 26-Mar-2021 | 1. | SS | Container
NC-IL081-30524 | 26-Mar-2021 | JEREMY.BROWNE | General | Internal | C25 | Press 25-IL081 |  |  | 8507SA | CONTAINER SHARPS 5QT RED | CFF-V-Flash (Functional Area) | Flash on the Rim on Cavity #1 out of Specification and pulling on the rim due to the container stocker machine. | 21C09463 | 51740.0 | Stopped the press and performed End of Run. | 26-Mar-2021 | 26-Mar-2021 | CFF-V-Flash (Functional Area) | Sharp Flash; Hinge Pins and Holes; All other areas: Less than or equal to 0.060" and no more than 3 locations per part. | Machine Failure | Machine Failure | Reject' 'Accept | 780.0 50960.0 | 0.0 | Closure | CLOSED | 25-Apr-2021 | 02-Apr-2021 | 7. | SS | Container
NC-IL081-30546 | 27-Mar-2021 | BENJAMIN.KELLY | General | Internal | S8P | Packer-IL081 |  |  | 1181200777 | 12ML SYRINGE LUER LOCK TIP | SFF-SPBP-V-Labeling/Print 3 | Over shipper label does not match drawing 21V30763X (21C30763X is correct) | 21C30763X | 93.0 | Stop & Clear line 8 packer. | 26-Mar-2021 | 26-Mar-2021 | SFF-SPBP-V-Labeling/Print 3 | Print does not match drawing | Programed incorrectly | Upon re-starting of the packer, operator had to re-program the labels. He incorrectly typed in the lot number. | Accept | 93.0 | 0.0 | Closure | CLOSED | 26-May-2021 | 21-Apr-2021 | 25. | 8 | Syringe
NC-IL081-30554 | 27-Mar-2021 | FIDELIA.CANTOR | General | Internal | C56 | Press 56-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | CFF-D-Molding 1 | Manufacturing found container with wall thickness out of specification during scheduled inspection. 0.058" and should be 0.060" min. per GS20002953 Rev. 5. | 21B00263 | 400.0 | Placed affected product on hold. | 27-Mar-2021 | 27-Mar-2021 | CFF-D-Molding 1 | Wall Thickness | Tooling/ fixture damaged | Damaged tooling/ fixture causing process variation | Reject | 400.0 | 0.0 | Closure | CLOSED | 26-Apr-2021 | 16-Apr-2021 | 19. | G8 | Container
NC-IL081-30579 | 29-Mar-2021 | LAKEYSHA.SIMPSON | General | Internal | Workstation | Workstation-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-SPIC-V-Labeling/Print 1 | Incorrect lot number on the inner carton (21C05363X) Should be 21B05363X. | 21B05363X | 220.0 | Stop and clear line, perform checkback and place product on QA Hold | 29-Mar-2021 | 29-Mar-2021 | SFF-SPIC-V-Labeling/Print 1 | Lot Number and Expiration Date (where applicable) is incomplete or missing | Inadequate Procedure | Procedural Gap identified in GS20003063 for lack of verification guidance when printers lose their programming. Printer programming in the Syringe Focus Factory is a manual process and susceptible to human error. | Accept | 220.0 | 0.0 | Closure | CLOSED | 28-May-2021 | 21-Apr-2021 | 23. | 412 | Syringe
NC-IL081-30616 | 30-Mar-2021 | BERTHA.HERNANDEZ | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-PPHP-V-Lot Number 5 | Incorrect expiration date on cap print side B, found 2036-03-27, correct date 2026-03-27 | 21B33763X | 172800.0 | Packer stopped and cleared, re-programmed side B with the correct expiration date. | 30-Mar-2021 | 29-Mar-2021 | SFF-PPHP-V-Lot Number 5 | Printing is in the incorrect position and illegible | Inadequate Procedure | Procedural Gap identified in GS20003063 for lack of verification guidance when printers lose their programming. Printer programming in the Syringe Focus Factory is a manual process and susceptible to human error. | Reject' 'Accept | 7040.0 165760.0 | 0.0 | Closure | CLOSED | 30-Jun-2021 | 14-Jul-2021 | 105. | 3 | Syringe
NC-IL081-30623 | 30-Mar-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-684 WAS OOT-AF DURING VENDOR CALIBRATION. VENDOR ADJUSTED INTO TOLERANCE; CALIBRATION PASSED; NCR INITIATED PER GS10000123; ONE EVENT OCCURRED DURING THE PREVIOUS TWO CALIBRATION PERIODS. |  |  |  | 30-Mar-2021 | 15-Mar-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Calibration/Equip.Verification | Calibration/Equip.Verification |  |  | 0.0 | Closure | CLOSED | 29-Apr-2021 | 19-Apr-2021 | 20.
NC-IL081-30637 | 31-Mar-2021 | ANGELICA.AGUILAR | General | Internal | C74 | Press 74-IL081 |  |  | 8980S | 8 GA RENEWABLES SLIDE LID | CFF-V-Quantity per Case | QA found Finished Good product missing components (Lids) during a scheduled inspection. | 21A10063 | 180.0 | Performed End of Run Line Clearance and all affected product was placed on hold. | 30-Mar-2021 | 30-Mar-2021 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | Human Error | Accept | 180.0 | 0.0 | Closure | CLOSED | 29-Apr-2021 | 08-Apr-2021 | 8. | G8 | Container
NC-IL081-30653 | 31-Mar-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | SR-1111; OUT OF TOLERANCE AS FOUND; VENDOR ADJUSTED BACK INTO TOLERANCE; CALIBRATION PASSED. |  |  | None required. Force Gage not released for use. | 31-Mar-2021 | 24-Mar-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment variation requiring maintenance adjustments | Equipment variation requiring maintenance adjustments |  |  | 0.0 | Closure | CLOSED | 29-Jun-2021 | 07-Jun-2021 | 67.
NC-IL081-30680 | 01-Apr-2021 | MONA.ABOKAUID | General | Internal | S6B | 6B-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Broken/Damaged 2 | Damaged / broken component functional but less than desired. Cavity 18H, mold 58205, MP-1137 | 21C36463X | 211000.0 | Stop and clear line, scrap WIP, perform checkback and place product on QA Hold. | 01-Apr-2021 | 31-Mar-2021 | SFF-M-V-Broken/Damaged 2 | Broken or damaged product but still Functional | Machine Failure | Machine Failure | Reject | 211000.0 | 0.0 | Closure | CLOSED | 02-Jul-2021 | 24-Jun-2021 | 84. | Sub | Syringe
NC-IL081-30783 | 06-Apr-2021 | BERTHA.HERNANDEZ | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SPIC-V-Labeling/Print 4 | Failed UDI grade test for inner carton. | 21C47463X | 330.0 | Stopped and cleared the line, Performed checkback, placed product on hold. | 06-Apr-2021 | 05-Apr-2021 | SFF-SPIC-V-Labeling/Print 4 | Fails UDI Grade Test | Ink Viscosity | Ink Viscosity | Reject' 'Accept | 18.0 312.0 | 0.0 | Closure | CLOSED | 26-Apr-2021 | 12-Apr-2021 | 6. | 3 | Syringe
NC-IL081-30797 | 06-Apr-2021 | BENJAMIN.KELLY | General | Internal | S6P | Packer-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Plugs/Broken Pins 2 | Missing pin created a plugged opening Cav# 17. Mold 58110, MP-1134 | 21C48063X | 6000.0 | Stopped line and cleared of all in-process parts, performed checkback and placed product on QA Hold. | 06-Apr-2021 | 06-Apr-2021 | SFF-M-V-Plugs/Broken Pins 2 | Missing pins created a plugged opening | Pin came out of the core | The pin came loose from the core causing tip to fill in with material. | Accept | 6000.0 | 0.0 | Closure | CLOSED | 04-Jun-2021 | 20-Apr-2021 | 13. | Sub | Syringe
NC-IL081-30856 | 09-Apr-2021 | ANGELICA.AGUILAR | General | Internal | C67 | Press 67-IL081 |  |  | 31143665 | 10 GA SHARPS CONTAINER | CFF-V-Quantity per Case | During scheduled inspection, QA found packed 10 lids into the FG box instead 6 lids per Standard Work FM20008927 Rev. 0 | 21C38663 | 3240.0 | All affected product was placed on hold and manufacturing performed the End of Run Line Clearance. | 08-Apr-2021 | 08-Apr-2021 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | Human Error | Accept | 3240.0 | 0.0 | Closure | CLOSED | 08-May-2021 | 20-Apr-2021 | 11. | G10 | Container
NC-IL081-30857 | 09-Apr-2021 | AUGUSTINA.MASH | General | Internal | S2N | Hot Stamp Printer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-P-V-Numbering 1 | Print missing for 1 or more numeral(s) | 21C47563X | 104160.0 |  | 08-Apr-2021 | 08-Apr-2021 | SFF-P-V-Numbering 1 | Type 1: Missing 1 or more numeral(s) | Bad cliche | The cliché on head 1 was worn causing print to not transfer onto the barrel. | Accept' 'Reject | 95506.0 8654.0 | 0.0 | Closure | CLOSED | 08-Jun-2021 | 29-Jul-2021 | 111. | 3 | Syringe
NC-IL081-30866 | 09-Apr-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | UW-01 OOT-AF; ADJUSTED INTO TOLERANCE; CALIBRATION PASSED. |  |  |  | 09-Apr-2021 | 08-Apr-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 13-May-2021 | 21-Apr-2021 | 12.
NC-IL081-30867 | 09-Apr-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | UW-02; OOT-AF; ADJUSTED INTO TOLERANCE; CALIBRATION PASSED. |  |  |  | 09-Apr-2021 | 08-Apr-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 13-May-2021 | 21-Apr-2021 | 12.
NC-IL081-30868 | 09-Apr-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | UW-03; OOT-AF; ADJUSTED INTO TOLERANCE; CALIBRATION PASSED. |  |  |  | 09-Apr-2021 | 08-Apr-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 13-May-2021 | 21-Apr-2021 | 12.
NC-IL081-30870 | 09-Apr-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0698; LOST; UNABLE TO LOCATE FRIDAY 04/09/2021. |  |  |  | 09-Apr-2021 | 09-Apr-2021 | Calibration Lost Guage |  | Calibration/Equip.Verification | Calibration/Equip. Verification |  |  | 0.0 | Closure | CLOSED | 13-May-2021 | 19-Apr-2021 | 10.
NC-IL081-30919 | 12-Apr-2021 | BENJAMIN.KELLY | General | Internal | S7P | Packer-IL081 |  |  |  |  | Procedure Not Followed | Failure to follow procedure. Lot found to have insufficient inner carton inspections post release. \n238 performed, 315 required. Found on DHR for product code 8881907102, Lot# 21C03963X. BPCS hold number 21040027 initiated for trending purposes only. |  |  | No immediate actions required. Initial assessment of start and end labels show no evidence of incorrect lot information nor legibility issues. | 12-Apr-2021 | 08-Apr-2021 | Procedure Not Followed | Procedure Not Followed | PR-Non Compliant with Procedure | The employee didn’t realize that more inner cartons needed to be sampled at the end of run and failed to verify that the correct number of inner cartons were sampled. |  |  | 0.0 | Closure | CLOSED | 12-May-2021 | 28-Apr-2021 | 15.
NC-IL081-31025 | 17-Apr-2021 | BERTHA.HERNANDEZ | General | Internal | S4A | 412A-IL081 | Equipment ID | 58203-IL081 | 8881412012 | 412 CURVED TIP SYRINGE | SFF-M-V-Shorts 1 | Short in fluid pathway (barrel tip) cavity 13B. | 21D13163X | 17500.0 | Stopped and cleared the line, placed product on hold. | 17-Apr-2021 | 17-Apr-2021 | SFF-M-V-Shorts 1 | Shorts in fluid pathway or in area where Sterility is required | Blocked Vents | The design of the tool has small vents as the tip of the syringe in very intricate. The small vents can become blocked reducing the ability for the air to escape. The vents becoming blocked are an inherent part of the molding process. | Accept' 'Reject | 17499.0 1.0 | 0.0 | Closure | CLOSED | 16-Jun-2021 | 14-Jun-2021 | 58. | 412 | Syringe
NC-IL081-31046 | 19-Apr-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | MD-0851 (6" ruler) could not be located during normal calibration rounds on 04/14/2021. Mass email sent asking for help in locating ruler, also on 04/14/2021. Calibration was due 04/17/2021 |  |  |  | 19-Apr-2021 | 17-Apr-2021 | Calibration Lost Gauge | Calibrations Lost Gauge | Calibration/Equip.Verification | Calibration/Equip.Verification |  |  | 0.0 | Closure | CLOSED | 19-May-2021 | 21-Apr-2021 | 2.
NC-IL081-31065 | 19-Apr-2021 | ANGELICA.AGUILAR | General | Internal | C56 | Press 56-IL081 |  |  | 8980S | 8 GA RENEWABLES SLIDE LID | CFF-V-Quantity per Case | QA found during scheduled inspection, Finished Product missing component (IFU). | 21A20963 | 8620.0 | Machined went down to perform the End of Run Line Clearance. All affected was placed on hold. | 19-Apr-2021 | 19-Apr-2021 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | Human Error | Accept | 8620.0 | 0.0 | Closure | CLOSED | 19-May-2021 | 23-Apr-2021 | 3. | G8 | Container
NC-IL081-31158 | 22-Apr-2021 | GRACE.AUSTIN | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | CFF-V-Label Verification and Barcode Verification 10 | Missing Biohazard Label on both side of the container and one container missing one biohazard label on one side. Found during QA scheduled inspection. | 21B21063 | 2280.0 | Check back perform and contain affected product. | 22-Apr-2021 | 22-Apr-2021 | CFF-V-Label Verification and Barcode Verification 10 | BH - Missing label | Human Error | The operator failed to follow the standard work FM20008300 Rev: 1 to label both sides straight and wrinkle free because they forgot to jog the labels forward after performing an inspection on the labels. | Accept | 2280.0 | 0.0 | Closure | CLOSED | 22-May-2021 | 05-May-2021 | 12. | G2 | Container
NC-IL081-31181 | 23-Apr-2021 | GRACE.AUSTIN | General | Internal | C42 | Press 42-IL081 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-D-Molding 1 | Documented Low Wall Thickness Cav. #2, 0.053"s, on 4/21/21 Crew B. Min. Spec. 0.055". \nQA found during DHR final review on EOR. | 21A21763' '21A07563' '21D15763 | 47100.0 30220.0 20880.0 | All affected product has been contained. | 23-Apr-2021 | 23-Apr-2021 | CFF-D-Molding 1 | Wall Thickness | un-balanced mold cavitation | Mold cavitation reduced causing variation during the filling process | UseAsIs | 47100.0 30220.0 20880.0 | 0.0 | Closure | CLOSED | 23-May-2021 | 17-May-2021 | 24. | SS | Container
NC-IL081-31184 | 23-Apr-2021 | BERTHA.HERNANDEZ | General | Internal | S4R | RAM-IL081 |  |  | 8881412012 | 412 CURVED TIP SYRINGE | SFF-SA-V-Mixed Product 2 | Incomplete assembly of product (missing plunger). | 21D18963X | 1000.0 | Stopped and cleared the packer, Performed checkback, placed product on hold. | 23-Apr-2021 | 23-Apr-2021 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Tooling/ fixture wear | The cylinder for assembly in the RAM was not working properly by not pushing the plungers fully down into the tip. This caused the plungers to dislodge from the assembly during transfer. | Accept | 1000.0 | 0.0 | Closure | CLOSED | 23-May-2021 | 30-Apr-2021 | 7. | 412 | Syringe
NC-IL081-31186 | 23-Apr-2021 | MARIA.ROSALES | General | Internal | C81 | Press 81-IL081 |  |  | 89671 | 2 GA HORIZONTAL DROP - T RED | CFF-V-Warping, Deformation and Pulling | Deformation on rim area for cavity 1 and 2. Found during QA scheduled inspection. | 21C18863 | 23080.0 | Press went down and performed the End of Run Line Clearance. All affected product was contained. | 23-Apr-2021 | 23-Apr-2021 | CFF-V-Warping, Deformation and Pulling | Distortion affecting functionality | Tooling/ fixture wear | Upon investigation, it was found that the bolts on the end of arm tooling were not fully tightened so that end of arm tooling was loose. This caused the end of arm tooling to pick the containers off of the mold improperly, resulting in deformation on the rim of the container. | Accept | 23080.0 | 0.0 | Closure | CLOSED | 22-Jun-2021 | 04-Jun-2021 | 42. | G2 | Container
NC-IL081-31371 | 01-May-2021 | MONA.ABOKAUID | General | Internal | S8A | 8A-IL081 |  |  |  |  | Documentation Missing | failure to document PCR once per shift. (04/30/21 Crew A PCR reading missing) | 21D197299 | 35000.0 |  | 30-Apr-2021 | 30-Apr-2021 | Documentation Missing | Documentation Missing | Human Error | Human Error |  |  | 0.0 | Closure | CLOSED | 31-May-2021 | 21-May-2021 | 20.
NC-IL081-31411 | 04-May-2021 | MARIA.ROSALES | General | Internal | C52 | Press 52-IL081 |  |  | 8870- | 18GA PHARMSAFETY WASTE HINGED | CFF-V-Contamination | Grease and oil contamination on containers was found during QA scheduled inspection. | 21C00763 | 365.0 | Press went down to perform correction for the non-conformance and begin the End of Run Line Clearance. | 04-May-2021 | 04-May-2021 | CFF-V-Contamination | Carbon or White Specks greater than 0.080" (Surfaces A, B, C, & D) | Machine Failure | "The machine's clamp cylinder was found to have worn seals that were leaking oil. This component was nearing the end of its expected life." | Reject' 'Accept | 65.0 300.0 | 0.0 | Closure | CLOSED | 03-Jul-2021 | 17-Jun-2021 | 44. | G18 | Container
NC-IL081-31414 | 04-May-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | RC-013; RECYCLER LOAD CELL; OUT OF TOLERANCE |  |  |  | 04-May-2021 | 04-May-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Facilities/Equipment | Facilities/Equipment - Power surge causing load cell overload |  |  | 0.0 | Closure | CLOSED | 03-Jul-2021 | 30-Jun-2021 | 56.
NC-IL081-31427 | 04-May-2021 | MARIA.ROSALES | General | Internal | C84 | Press 84-IL081 |  |  | WP7605 | ASSMBLY 5QT CLR TORTUOUS LID | Procedure Not Followed | Failed to perform Start-up checklist when Tool 7606/7568 were pulled and set back to the press. | 21A195740 | 1.0 | Document the non-conformance and manufacturing performed the Start-up Checklist required for the procedure. Not product was impacted. | 04-May-2021 | 30-Mar-2021 | Procedure Not Followed | Procedure Not Followed | People/Human Error/ | The completed mold removal and mold set-up checklists were not filed in the Device History Record (DHR) as outlined in GS10000122, step 5.8.2 and GS10004611, step 5.14.2. The Lead on shift was not aware that the tool had been pulled for PM, which would require a start-up checklist before starting per GS10000122. |  |  | 0.0 | Closure | CLOSED | 03-Jun-2021 | 17-May-2021 | 12. | G5 | Container
NC-IL081-31502 | 06-May-2021 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | CM285160000 | Unit Box 12ml Soft Pack printed - Covidien | Packaging Artwork - Poor / Missing Print | On 5/05/21 manufacturing discovered that inner cartons, part number CM285160000, are missing print on one side of the inner carton (quantity per box), lot # 529254 PO # 823991. Not acceptable per artwork (AW20021997 Rev. 1). | 529254 | 800.0 | Placed affected product on hold | 05-May-2021 | 05-May-2021 | Packaging - Missing Unit Box Print | Packaging - Missing Unit Box Print | Process/Method | Process/Method. Not enough anilox roll pressure on our die cutting/printing machine caused the print to fade on a few cartons.\n\nDuring the inspection, the machine set up was found to be completed properly. | Reject | 800.0 | 0.0 | Closure | CLOSED | 31-Aug-2021 | 15-Jun-2021 | 40.
NC-IL081-31503 | 06-May-2021 | AUGUSTINA.MASH | General | Internal | S7N | Pad Printer-IL081 |  |  | 8881112059 | 12ML SYRINGE REGULAR NS | SFF-P-V-Alpha Numeric 1 | illegible print or missing one unit of measure | 21D14963X | 3000.0 | Stopped and cleared line | 06-May-2021 | 06-May-2021 | SFF-P-V-Alpha Numeric 1 | Type 1: Illegible print or missing one unit of measure | Barrel jams | Barrels were jamming under the print pad causing pieces of the material to break off. | Accept | 3000.0 | 0.0 | Closure | CLOSED | 07-Jul-2021 | 12-May-2021 | 5. | 7 | Syringe
NC-IL081-31527 | 07-May-2021 | BERTHA.HERNANDEZ | General | Internal | S7N | Pad Printer-IL081 |  |  | 8881112059 | 12ML SYRINGE REGULAR NS | SFF-P-V-Numbering 1 | Missing one or more numerals | 21D14963X | 4000.0 | Stopped and cleared the line, performed check back and placed product on hold. | 07-May-2021 | 07-May-2021 | SFF-P-V-Numbering 1 | Type 1: Missing 1 or more numeral(s) | Printer Malfunction | AM-224 was found to be producing barrels with missing print due to print pad being damaged from the barrels jamming under print pad which was coming in from infeed off of the star wheel. Additionally, the set screw kept working its way loose. | Accept | 4000.0 | 0.0 | Closure | CLOSED | 06-Jun-2021 | 17-May-2021 | 10. | 7 | Syringe
NC-IL081-31641 | 13-May-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0502; OOTAF; AM-171, WELDER 4.\nTHIS CALIBRATION PERFORMED EARLY DUE TO MAINTENANCE REPAIR ACTIVITY ON THE WELDER STROKE. LINEAR ENCODER AND DOWNSPEED WERE OOTAF. BOTH WERE ADJUSTED BACK INTO TOLERANCE. CALIBRATION PASSED AFTER ADJUSTMENT. |  |  |  | 13-May-2021 | 12-May-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 17-Jun-2021 | 07-Jun-2021 | 24.
NC-IL081-31777 | 19-May-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0499, MD-0500, MD-0501, MD-0502: All out of tolerance as found for linear encoder. All adjusted back into tolerance.\nMD-0501 and MD-0502: Both out of tolerance as found for force. MD-0501 adjusted back into tolerance. MD-0502 required repair on W.O. CL-3308252, and then was adjusted into tolerance. |  |  | MD-0501 adjusted back into tolerance. MD-0502 required repair on W.O. CL-3308252, and then was adjusted into tolerance. | 19-May-2021 | 19-May-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 30-Jun-2021 | 23-Jun-2021 | 34.
NC-IL081-31800 | 20-May-2021 | BEATRIZ.GUILLEN | General | Internal | S8D | 8D-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-SA-V-Mixed Product 1 | Incorrect or mixed product, used plunger Part# 3200804 should be X9900. | 21E09263X | 127000.0 | Plunger Press stopped and line cleared. Performed check back. | 20-May-2021 | 20-May-2021 | SFF-SA-V-Mixed Product 1 | Mixed Product Type 1: Incorrect or mixed product | Procedure/Method | The Start-Up checklist was followed as designed but did not cover the Cavity Location Sheet or the Tool Room Defect Sheet. This resulted in missing the 58121 tool was converted to 58121-02 and causing softpack plungers to be fed to line 6 or mixed product. | Reject | 127000.0 | 0.0 | Closure | CLOSED | 19-Jul-2021 | 22-Jun-2021 | 33. | Sub | Syringe
NC-IL081-31827 | 21-May-2021 | BERTHA.HERNANDEZ | General | Internal | S2P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-PPHP-V-Lot Number 1 | Smeared print on cap not legible. | 21D20963X | 130080.0 | Replaced cartridge, inspected pallet 1 case 1-5 no defects found. placed product on hold. | 21-May-2021 | 20-May-2021 | SFF-PPHP-V-Lot Number 1 | Missing Print on the Cap | Dirty Printer | Side A was down, and prior to restarting the cartridge was not cleaned per the standard work to ensure it was not dried out from being down. | Reject' 'Accept | 13760.0 116320.0 | 0.0 | Closure | CLOSED | 18-Sep-2021 | 08-Jul-2021 | 47. | 3 | Syringe
NC-IL081-31841 | 21-May-2021 | ANGELICA.AGUILAR | General | Internal | C81 | Press 81-IL081 |  |  | 89671 | 2 GA HORIZONTAL DROP - T RED | CFF-V-Blisters, Bubbles and Voids | Bubble out specification on the gate area of the container cavity 1 was found during scheduled QA final inspection. | 21C35163 | 15680.0 | Press went down and manufacturing started the End of Run Line Clearance. Al affected product has been contained. | 21-May-2021 | 21-May-2021 | CFF-V-Blisters, Bubbles and Voids | Single bubble greater than or equal to 0.080”; Up to 3 bubbles between 0.060 and 0.080”; Up to 4 bubbles less than or equal to 0.060”; Up to 8 bubbles less than 0.030”; Living Hinge = Reject All | Machine Failure | Machine Failure | Reject' 'Accept | 6620.0 9060.0 | 0.0 | Closure | CLOSED | 20-Jun-2021 | 17-Jun-2021 | 26. | G2 | Container
NC-IL081-31874 | 24-May-2021 | GRACE.AUSTIN | General | Internal | C89 | Press 89-IL081 |  |  | 8605RC | 5 QT RCRA HAZARDOUS WASTE | CFF-V-Short | Containers with shorts were packed into FG and found during QA scheduled inspection. | 21C26963 | 448.0 | Press went down and manufacturing proceeded to End of Run Line Clearance to perform lot change. | 24-May-2021 | 24-May-2021 | CFF-V-Short | Part is not completely formed; Missing plastic | Complex machine set-up | The press went down multiple times due to a mold protection alarm. During restart, the technician raised the transfer position from 24.35 to 34.35 for at least 2 shots and then dropped it back down to 24.35 to continue running. Raising the transfer position creates shorts in the part and is done to prevent flashing the cavity on restart of the machine. | Accept' 'Reject | 420.0 28.0 | 0.0 | Closure | CLOSED | 23-Jun-2021 | 01-Jun-2021 | 8. | G5 | Container
NC-IL081-31897 | 25-May-2021 | BERTHA.HERNANDEZ | General | Internal | S8B | 8B-IL081 |  |  | 4502570 | 12CC ANTI-REFLUX SYRINGE | SFF-M-V-Flash/Gate 3 | Flash in non-functional area exceeds 1/16" on barrel 58119 cavity 7E | 21E16363X' '21E16463X | 335000.0 106000.0 | Line 6 was stopped and cleared, start check back. | 25-May-2021 | 25-May-2021 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Tooling/ fixture wear | Mold 58119 cavity 7E had gate string that was unable to be processed out resulting in NCR. The most probable root cause of this is the gate sticking on the cavity which resulted in string flash. | Reject' 'Accept | 335000.0 106000.0 | 0.0 | Closure | CLOSED | 23-Aug-2021 | 25-Aug-2021 | 91. | Sub | Syringe
NC-IL081-31909 | 25-May-2021 | AUGUSTINA.MASH | General | Internal | S2D | 2D-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-M-V-Flash/Gate 3 | Flash in non functional area. Cap | 21E08463X | 21120.0 | Line was stopped. Hoppers cleared. All affected product was contained. | 25-May-2021 | 25-May-2021 | SFF-M-V-Flash/Gate 3 | Flash Type 3: Flash in NON-Functional Areas (including gates) exceeds 1/16” (.063”) | Probe temp too high | Injection molding press was down when defect was found at the line, per process technician the most probable cause would have been the probe temperatures too high. | Reject' 'Accept | 2246.0 18874.0 | 0.0 | Closure | CLOSED | 23-Jul-2021 | 14-Jul-2021 | 49. | 3 | Syringe
NC-IL081-31912 | 26-May-2021 | MONA.ABOKAUID | General | Internal | S2R | RAM-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-SA-V-Damaged Components 3 | damaged product functional but less than desired . Lure skirt damage caused by printer. | 21E08463X | 87840.0 | Stop line, clear, perform checkback, place product on QA Hold | 26-May-2021 | 25-May-2021 | SFF-SA-V-Damaged Components 3 | Damaged Components Type 3: Damaged product/component (functional but less than desired); Scratches/drag marks on the barrel | Misaligned assembly transfer | Misaligned assembly transfer | Reject | 87840.0 | 0.0 | Closure | CLOSED | 25-Jun-2021 | 04-Aug-2021 | 70. | 3 | Syringe
NC-IL081-31933 | 27-May-2021 | AUGUSTINA.MASH | General | Internal | S7N | Pad Printer-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-P-F-Tolerance on Graduated Capacity | Skewed print on barrel-Tolerance on graduated capacity | 21D26963X | 28500.0 | Stop line, clear, perform checkback, place product on QA Hold. | 26-May-2021 | 26-May-2021 | SFF-P-F-Tolerance on Graduated Capacity | Tolerance on Graduated Capacity | Rail not set evenly | Back rail front end not set evenly with the back end of back rail.\n\nTransfer sprockets were pushing and lifting barrel too early\n\nGap too wide on the escapement jam trip gate | Reject' 'Accept | 17000.0 11500.0 | 0.0 | Closure | CLOSED | 26-Jul-2021 | 22-Jun-2021 | 26. | 7 | Syringe
NC-IL081-31977 | 28-May-2021 | MARQUES.CALIMEE | General | Internal | C23 | Press 23-IL081 |  |  | WP8520N | 85221R LID | CFF-F-Welding | Weld Strength test failed during production scheduled inspection. | 21D196962 | 70.0 | Press went down. All affected product was contained. | 28-May-2021 | 28-May-2021 | CFF-F-Welding | Weld Strength | Equipment variation requiring maintenance adjustments | The process was not changed from the prior linear encoder which was nearing its end of life. After calibrating the welder and replacing the linear encoder, adjustments need to be made to the Energy (J), Down Speed, and Air Pressure. | Reject | 70.0 | 0.0 | Closure | CLOSED | 27-Jul-2021 | 12-Jul-2021 | 44. | G3 | Container
NC-IL081-32088 | 03-Jun-2021 | BEATRIZ.GUILLEN | General | Internal | S3P | Packer-IL081 |  |  | 8881512878 | 12ML SYRINGE LUER LOCK TIP RP | SFF-PPHP-V-Lot Number 1 | Incorrect cap print on side B reads SR should be SLK per FS10011723 . | 21E23163X' '21E23063X | 4320.0 84960.0 | Packer stopped and cleared the line, \nside B was reprogramed, performed Check Back. | 03-Jun-2021 | 02-Jun-2021 | SFF-PPHP-V-Lot Number 1 | Missing Print on the Cap | Incorrect Setup | The packer lost its programmed lot information for the cap printer due to being shut down. The Cell Tech reprogrammed the packer and the incorrect description was selected. Upon restart the information was not verified against lot packet and a verification was not performed as required in GS20003063. | Reject' 'Accept | 4320.0 28070.0 56890.0 | 0.0 | Closure | CLOSED | 01-Sep-2021 | 22-Jun-2021 | 19. | 3 | Syringe
NC-IL081-32090 | 03-Jun-2021 | GRACE.AUSTIN | General | Internal | C42 | Press 42-IL081 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-V-Label Verification and Barcode Verification 10 | 4 containers with missing biohazard label on both sides and found short on the handle on one container cavity # 1. | 21D01163 | 100.0 | Stop press and all affected product was contained. | 03-Jun-2021 | 03-Jun-2021 | CFF-V-Label Verification and Barcode Verification 10 | BH - Missing label | Incorrect Lot Number | The printer queue was not cleared from the computer that was used to program the labeler. The operator loaded the correct program into the computer, however the incorrect program that was in the queue made its way to the labels. This was detected by the operator and the labeler was shut off, labels with incorrect product code on them removed, and labeler was programmed with correct information. However, during this time, the press was still running until it was eventually stopped by the operator, and a few containers made its way through the production stream without labels. Upon restart, the first shot that produced parts with a short in it was not removed from the production stream. | Accept | 100.0 | 0.0 | Closure | CLOSED | 03-Jul-2021 | 18-Jun-2021 | 15. | SS | Container
NC-IL081-32093 | 03-Jun-2021 | JACLYN.GRISKO | General | Internal | S8B | 8B-IL081 |  |  | 4502500 | SYR 12ml WFF Magellan Sub | Internal Threads issue | During further manufacturing of product code 4502500, lot# 20D02863X at the Norfolk facility, it was determined that there was a thread orientation issue with cavities 34E, 41E, and 43E which are produced off of MP-1133. | 20D02863X | 315000.0 | All affected product is contained at our sister facility and will be returned to the Crystal Lake site for product correction. No further action will be required from a containment perspective as:\n1 - Issue with thread start orientation was corrected in 05/2021 and confirmed with mold repair.\n2 - This defect is limited to the product code 4502500 and this was only lot produced during the suspect timeframe of 04/08/2021 - 04/28/2021. | 03-Jun-2021 | 08-Apr-2020 | Dimensional- out of specification | Dimensional- out of specification | Inadequate Procedure | Inadequate Procedure | Reject | 315000.0 | 0.0 | Closure | CLOSED | 03-Jul-2021 | 30-Aug-2021 | 88. | 8 | Syringe
NC-IL081-32106 | 03-Jun-2021 | BENJAMIN.KELLY | General | Internal | S7P | Packer-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-P-V-Numbering 1 | Missing one or more numerals. | 21E09563X | 1000.0 | Stopped and cleared the line, performed check back and placed product on hold. | 03-Jun-2021 | 03-Jun-2021 | SFF-P-V-Numbering 1 | Type 1: Missing 1 or more numeral(s) | Gap on escapement jam trip gate too wide | AM-224 was found to be producing barrels with twisted print due to parts jumping off chain to early. Found back rail front end not set evenly with the back end of back rail. Pulled apart and re-aligned. Reset micrometer for grad line adjustment to allow room for adjustment to either side. Adjusted grad line and adjusted both transfer sprockets as they were pushing/lifting barrel a bit too early causing twisted print. After adjusted transfer sprockets, adjustments to harmonic collar to center print had to be done several times with every adjustment of transfer sprocket and rail adjustment. ITW came in to work on the printer they adjusted pad height, synchronized the barrel chain in feed to the pad, and adjusted the infeed rail. Printer still having issues ITW returned and found the gap on the escapement jam trip gate was too wide. Adjustments made to narrow the gap to eliminate the turning and jamming into the star wheel. | Reject | 1000.0 | 0.0 | Closure | CLOSED | 03-Aug-2021 | 22-Jun-2021 | 18. | 7 | Syringe
NC-IL081-32213 | 08-Jun-2021 | BEATRIZ.GUILLEN | General | Internal | S7R | RAM-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-SA-V-Mixed Product 2 | Stop line and perform checkback, no WIP was affected | 21E09763X | 500.0 |  | 08-Jun-2021 | 08-Jun-2021 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Station out of alignment | The plunger push station most likely became not centered due to someone accidently bumping it during weekly cleaning activities. As such, the rubber tips were not fully seated down. | Accept | 500.0 | 0.0 | Closure | CLOSED | 08-Jul-2021 | 10-Jun-2021 | 2. | 7 | Syringe
NC-IL081-32244 | 09-Jun-2021 | FIDELIA.CANTOR | Receipt | Supplier | INC | Incoming-INC-IL081 |  |  | PT00006657 | Folding Carton, Printed, MONOJECT 10 mL Oral Syringe, CE Mar | Out of Spec Product Functional | On 6/07/21 was discovered on ( line 7 CME) that the inner cartons of product code PT00006657 LOT # 351783, were falling apart at assembly/not glued at the joint. | 351783 | 3300.0 | Placed affected product on hold. | 09-Jun-2021 | 07-Jun-2021 | Out of Spec Product Functional | Out of Spec Product Functional | PR-Undetermined | PR-Undetermined-MEDMFGPR. | Reject | 3300.0 | 0.0 | Closure | CLOSED | 10-Sep-2021 | 11-Aug-2021 | 62.
NC-IL081-32281 | 10-Jun-2021 | BEATRIZ.GUILLEN | General | Internal | S7R | RAM-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-SA-V-Mixed Product 2 | Incomplete assembly of product, Plunger not seated in shoulder. | 21E09963X | 15500.0 | Line stopped and performed Check Back. Product placed on Hold. | 10-Jun-2021 | 10-Jun-2021 | SFF-SA-V-Mixed Product 2 | Mixed Product Type 2/3: Incomplete or incorrect assembly of product | Air valve malfunction | The air pressure at the completed assembly transfer station was too high and the excess pressure caused the air to enter the barrel and move the plunger back slightly during transfer. One of the air nozzles were also found to be defective. | Accept | 15500.0 | 0.0 | Closure | CLOSED | 10-Jul-2021 | 26-Jun-2021 | 15. | 7 | Syringe
NC-IL081-32403 | 15-Jun-2021 | ANGELICA.AGUILAR | General | Internal | C61 | Press 61-IL081 |  |  | WP7417P | 8618RC LID | CFF-F-Assembly 2 | During manufacturing scheduled inspection, the Gasket Height and Leak Test failed. | 21D197226' '21C45063' '21C24263 | 21000.0 18690.0 6860.0 | FIPC-2 went down to process it and manufacturing started the End of Run Line Clearance. All affected product has been quarantined. | 15-Jun-2021 | 15-Jun-2021 | Inadequate amount of material flowing to head | Incorrect ratio during lid switch will create low gasket. | Ratio adjusted out of spec. | The Polyol was off by 1 gram. The Polyol was set at 4.7g and was supposed to be at 5.7g. This is still within the validated parameters but at the lower end and with material buildup on the head causes the material not to flow at optimal rate causing height issues. | Accept' 'Reject | 6860.0 18690.0 950.0 20050.0 | 0.0 | Closure | CLOSED | 15-Jul-2021 | 26-Jun-2021 | 10. | G18 | Container
NC-IL081-32531 | 22-Jun-2021 | BEATRIZ.GUILLEN | General | Internal | C89 | Press 89-IL081 |  |  | 31323358 | SAG4835TY GATORGUARD | CFF-V-Short | Short on the handle of containers. | 21D06463' '21E21463 | 384.0 112.0 | Stopped the line, Performed Check Back and placed all affected product on Hold. | 22-Jun-2021 | 22-Jun-2021 | CFF-V-Short | Part is not completely formed; Missing plastic | Equipment variation requiring maintenance adjustments | Press 89 is an older piece of equipment with a slow screw recovery when it is started up. When running with liquid colorant if the liquid color is turned on prior to the resin entering the screw, the screw will become coated and can cause slippage until the materials blend more homogeneously. This typically resolves itself during startup activities prior to product entering the finished good value stream. | Accept' 'Reject | 384.0 14.0 98.0 | 0.0 | Closure | CLOSED | 22-Jul-2021 | 15-Jul-2021 | 23. | GG | Container
NC-IL081-32658 | 28-Jun-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | SR-1175; Calibration out of tolerance as found; Gage repaired and passed calibration. |  |  |  | 28-Jun-2021 | 22-Jun-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Calibration/Equip.Verification | Calibration/Equip.Verification |  |  | 0.0 | Closure | CLOSED | 11-Aug-2021 | 05-Aug-2021 | 37.
NC-IL081-32765 | 03-Jul-2021 | MARK.VENTURA | General | Internal | S7C | 7C-IL081 |  |  |  |  | Documentation Missing | Failure to document PCR once per shift (07/02/21 Crew A, PCR readings missing). Line 7, Press MP-1113 | 21F05763X | 54000.0 | placed product on hold | 03-Jul-2021 | 02-Jul-2021 | [6 empty] | 0.0 | Investigation | VOID
NC-IL081-32837 | 07-Jul-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0500; OUT OF TOLERANCE AS FOUND; ADJUSTED INTO TOLERANCE. |  |  |  | 07-Jul-2021 | 01-Jul-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 12-Aug-2021 | 03-Aug-2021 | 26.
NC-IL081-32838 | 07-Jul-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0502; OUTOF TOLERANCE; ADJUSTED INTO TOLERANCE. |  |  |  | 07-Jul-2021 | 01-Jul-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 12-Aug-2021 | 03-Aug-2021 | 26.
NC-IL081-32855 | 08-Jul-2021 | ALEJANDRA.CASTANEDA | General | Internal | C52 | Press 52-IL081 |  |  | 8989 | 18 GA HINGED CHEMO - YELLOW | CFF-D-Molding 1 | Press down on C shift. Floater checked wall thickness on cold parts and discovered a low wall thickness from previous shift, 0.057" documented on Inspection Form and should be 0.060" min. per W.I.S. (GS20003049 Rev. 3). | 21D17763 | 4260.0 | Product affected has been contained. | 08-Jul-2021 | 08-Jul-2021 | CFF-D-Molding 1 | Wall Thickness | Mold Wear or Damage | Mold Wear or Damage |  |  | 0.0 | Implementation | INWORKS | 07-Aug-2021 |  |  | G18 | Container
NC-IL081-32892 | 10-Jul-2021 | BENJAMIN.KELLY | General | Internal | S7R | RAM-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-M-V-Broken/Damaged 1 | damage product component functional but less than desired Broken plunger tip end. | 21F22563X | 500.0 | Stop and clear line, scrap WIP plungers, complete checkback and place product on Hold. | 09-Jul-2021 | 09-Jul-2021 | SFF-M-V-Broken/Damaged 1 | Broken or damaged product rendering it Non Functional | Air nozzle | The air nozzle was not in the correct position and the pressure was too high | Accept | 500.0 | 0.0 | Closure | CLOSED | 09-Sep-2021 | 21-Jul-2021 | 11. | 7 | Syringe
NC-IL081-32925 | 13-Jul-2021 | BENJAMIN.KELLY | General | Internal | S7R | RAM-IL081 |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-PPHP-V-Particulate | Customer complaint # CASE-2021-00151654-1 Particulate in syringes ,\n pistons (rubber tips?) too loose or too tight & \nmarkings on syringes are smudged. | 19M11563X | 307200.0 |  | 12-Jul-2021 | 12-Jul-2021 | 1050 | CONTAMINATION - FOREIGN OBJECT/SUBSTANCE | PR-Particulate Matter | Particulate Matter-MEDMFGPR | Accept | 307200.0 | 0.0 | Closure | CLOSED | 10-Aug-2021 | 06-Aug-2021 | 24. | 7 | Syringe
NC-IL081-32926 | 13-Jul-2021 | BENJAMIN.KELLY | General | Internal |  |  |  |  | 6492739 | 12ML SYRINGE, LUER LOCK, P5M6K-080, | SFF-SA-V-Damaged Components 1 | Plungers Sticking & Particulate. See SCAR: CASE-2021-00151974-1 | 19M11663X | 307200.0 |  | 12-Jul-2021 | 12-Jul-2021 | Assembly Defect - Incomplete Assembly | Assembly Defect - Incomplete Assembly | Conveyor Speed | Conveyor Speed | Accept | 307200.0 | 0.0 | Closure | CLOSED | 11-Aug-2021 | 06-Aug-2021 | 24. | 7 | Syringe
NC-IL081-33067 | 17-Jul-2021 | MARQUES.CALIMEE | General | Internal | C67 | Press 67-IL081 |  |  | 31143665 | 10 GA SHARPS CONTAINER | CFF-D-Molding 1 | Manufacturing found during scheduled inspection wall thickness out of specification, where the reading was .058" and wall thickness specification should be .060". | 21D11563 | 1.0 | Press was shut down, and product contained. | 17-Jul-2021 | 17-Jul-2021 | CFF-D-Molding 1 | Wall Thickness | Data Cable | Loose connection of MagnaMike wire. | Accept | 1.0 | 0.0 | Closure | CLOSED | 21-Aug-2021 | 18-Aug-2021 | 31. | G10 | Container
NC-IL081-33098 | 19-Jul-2021 | ANGELICA.AGUILAR | General | Internal | C67 | Press 67-IL081 |  |  | 31143665 | 10 GA SHARPS CONTAINER | Out of Validated Parameter - Equipment Failure | Process technician found the press running out parameters in the fill time. Press running to 074sec. and should be 0.76-0.86sec the validated parameters. | 21E00963 | 228.0 | Press went down to prepare lot change and all affected product has been quarantined. | 19-Jul-2021 | 19-Jul-2021 | PR-- Operator error | Operator error-MEDMFGPR | Fill Time set below limits | Fill Time was documented running out of validated parameters. | Accept | 228.0 | 0.0 | Closure | CLOSED | 19-Aug-2021 | 26-Jul-2021 | 6. | G10 | Container
NC-IL081-33099 | 19-Jul-2021 | AUGUSTINA.MASH | General | Internal | S7C | 7C-IL081 |  |  | 5551771207 | 12CC BARREL REG WOS CLR | SFF-M-V-Flash/Gate 1 | Flash in luer tip exceeds 1/64". Cavity 59 Tool #58020 | 21G06763X | 18400.0 | Notified Manufacturing, Stopped and cleared line. | 19-Jul-2021 | 19-Jul-2021 | SFF-M-V-Flash/Gate 1 | Flash Type 1: Flash in Fluid Path (Luer Tip) exceeds 1/64” (.016”). Does not meets drawing specification | Tooling/ fixture wear | Wear on tooling causes variation | Accept | 18400.0 | 0.0 | Closure | CLOSED | 17-Sep-2021 | 17-Aug-2021 | 28. | 7 | Syringe
NC-IL081-33223 | 23-Jul-2021 | MARIA.ROSALES | General | Internal | C74 | Press 74-IL081 |  |  | 8608RC | 8 GA RCRA HAZARDOUS WASTE | Documentation Missing | Missing Auxiliary PCR inspection. on IRS001 Infra Red Staker . Actual reading was missing on 7/22/21 between 11:00pm - 3:00am. All Lids (WP7449) consume was used in Finish Good 8608RC. | 21c28363 | 1630.0 | Fallowing Shift performed Actual reading and manufacturing End of Run the product. | 23-Jul-2021 | 23-Jul-2021 | Documentation Missing | Documentation Missing | Human Error | The Cell Tech failed to perform the PCR reading on the welder because they thought that another person performed it due to line coverage to support training. | Accept | 1630.0 | 0.0 | Closure | CLOSED | 22-Aug-2021 | 06-Aug-2021 | 14. | G8 | Container
NC-IL081-33274 | 26-Jul-2021 | GRACE.AUSTIN | General | Internal | C32 | Press 32-IL081 |  |  | 85121 | 5 QT MAILBOX INROOM - CLEAR | CFF-V-Quantity per Case | Missing component (IFU) in 3 cases inspected for QA during scheduled final case inspection. | 21F32363 | 380.0 | Press was down and all affected product has been contained. | 26-Jul-2021 | 26-Jul-2021 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | Human error. The associate failed to follow Standard Work packing procedure, which states to place 1 IFU in the containers for this Finished Good product. | Accept | 380.0 | 0.0 | Closure | CLOSED | 25-Aug-2021 | 04-Aug-2021 | 8. | G5 | Container
NC-IL081-33354 | 28-Jul-2021 | JACLYN.GRISKO | General | Internal | CME | CME-IL081 |  |  |  |  | Documentation Incorrect | Failure to replace the obsolete Form (FM20009675) revision |  |  | Revision 0 of FM20009675, Refrigerator / Incubator Temperature Record was replaced with revision 1, the effective revision. | 28-Jul-2021 | 08-Apr-2021 | PR-Procedure Not Followed | Procedure Not Followed-MEDMFGPR | Procedure not followed | Procedure not followed |  |  | 0.0 | Implementation | INWORKS | 07-Oct-2021
NC-IL081-33445 | 30-Jul-2021 | JACLYN.GRISKO | General | Internal | QA | QA-IL081 |  |  |  |  | Documentation Incorrect | Basic skills paperwork was submitted to the newly transitioned training group and found use of an obsolete revisions of Training Attendance Sheet (GS10005711_PR Attachment 1). Revision 208 was being utilized instead of the current revision 213. Failure to follow section 5.4.4 of GS10000099 Rev 208 |  |  | See attached memo. No further containment action required. | 30-Jul-2021 | 30-Jul-2021 | Documentation Incorrect | Documentation Incorrect | Documentation | HR was removed from the site and access to document control system was not transitioned to the trainers fulfilling the role. |  |  | 0.0 | Closure | CLOSED | 29-Aug-2021 | 24-Aug-2021 | 25.
NC-IL081-33470 | 02-Aug-2021 | GRACE.AUSTIN | General | Internal | C31 | Press 31-IL081 |  |  | 31156916 | FISHER 2 GAL CONTAINER, RED | CFF-V-Fill-ins, Smear, Overprint, Ink Gap/Pad, Torn Label | Biohazard label print smeared and illegible was found during QA scheduled inspection. | 21G17263 | 1820.0 | Manufacturing shut down the press and fix the Label print. All affected product was contained. | 02-Aug-2021 | 02-Aug-2021 | CFF-V-Fill-ins, Smear, Overprint, Ink Gap/Pad, Torn Label | Excessive ink/print greater than 0.080" | Low Cooling Time | The parts were not cooled enough prior to stacking inside the finished good case because conveyor speed was too fast. | Reject' 'Accept | 280.0 1540.0 | 0.0 | Closure | CLOSED | 01-Sep-2021 | 18-Aug-2021 | 16. | G2 | Container
NC-IL081-33491 | 03-Aug-2021 | LAKEYSHA.SIMPSON | General | Internal | C19 | Press 19-IL081 |  |  | 8506SA | CONTAINER BIOHAZARD 5 QT | CFF-V-Flash | Flash in locking tab area more than .060" out of specification per GS10000913 | 21D23563' '21D23563 (WP8506) | 12260.0 320.0 | Fixed the Flash and change over product. contain affected product. | 03-Aug-2021 | 03-Aug-2021 | CFF-V-Flash (Functional Area) | Sharp Flash; Hinge Pins and Holes; All other areas: Less than or equal to 0.060" and no more than 3 locations per part. | Machine Failure | Machine Failure | Accept | 12260.0 320.0 | 0.0 | Closure | CLOSED | 02-Sep-2021 | 31-Aug-2021 | 28. | SS | Container
NC-IL081-33574 | 05-Aug-2021 | MARIA.ROSALES | General | Internal | C70 | Press 70-IL081 |  |  | 8906 | 1QT N/S CONTAINER YELLOW 100 | CFF-V-Labeling - Product & Overshipper 1 | Incorrect lot number on Biohazard Label 21F197970 should be 21D25963. | 21D25963 | 8500.0 | Stop production and Contain product. | 05-Aug-2021 | 05-Aug-2021 | CFF-V-Labeling - Product & Overshipper 1 | Product and Lot Number is incomplete or missing | Operator loads Incorrect product | Operator loads Incorrect product | Reject' 'Accept | 1700.0 6800.0 | 0.0 | Closure | CLOSED | 04-Sep-2021 | 16-Aug-2021 | 10. | G1 | Container
NC-IL081-33589 | 06-Aug-2021 | GRACE.AUSTIN | General | Internal | C67 | Press 67-IL081 |  |  | 8933 | CONTAINER, RED, 12GL, | CFF-V-Labeling - Product & Overshipper 1 | Manufacturing reported incomplete and missing print on lot number of the biohazard label during unscheduled inspection. | 21G04763 | 360.0 | Stop the production and all affected product was placed on hold. | 06-Aug-2021 | 06-Aug-2021 | CFF-V-Labeling - Product & Overshipper 1 | Product and Lot Number is incomplete or missing | Incorrect Setup | The ribbon was found to be slightly out of place, appearing to have shifted during operation. It is possible that the last person who changed the ribbon did not fully seat the ribbon in position which would cause it to shift over time. | Accept | 200.0 160.0 | 0.0 | Closure | CLOSED | 05-Sep-2021 | 20-Aug-2021 | 13. | G12 | Container
NC-IL081-33675 | 10-Aug-2021 | ANGELICA.AGUILAR | General | Internal | C61 | Press 61-IL081 |  |  | WP7417P | 8618RC LID | V-33 Flash out of spec (non functional area) | Manufacturing reported flash out of specification during unscheduled inspection. | 21F197872 | 1425.0 | Press was down to pull the tool and perform the Line Clearance End of Run. All affected product has been contained. | 10-Aug-2021 | 10-Aug-2021 | CFF-V-Flash (Non-Functional Area), Gate | Sharp Flash; Surfaces A & B: greater than or equal to 0.080”; Surface C: greater than or equal to 0.125”; All Surfaces: more than 3 locations in total | Tool Damage | Broken lifter on tool causing the flash. | Accept | 1425.0 | 0.0 | Closure | CLOSED | 09-Sep-2021 | 27-Aug-2021 | 17. | G18 | Container
NC-IL081-33682 | 11-Aug-2021 | ANGELICA.AGUILAR | General | Internal | C56 | Press 56-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | P-09 Calibration date overdue | During DHR End of Run review, was found out that the scale used during the start up of the job, the calibration was expired. | 21E07963 | 25380.0 | The product produced with this lot has been quarantined. | 10-Aug-2021 | 10-Aug-2021 | Calibration Expired | Calibration Expired | Human Error | Human Error, new associate. | Accept | 25380.0 | 0.0 | Closure | CLOSED | 10-Sep-2021 | 17-Aug-2021 | 6. | G8 | Container
NC-IL081-33792 | 13-Aug-2021 | GRACE.AUSTIN | General | Internal | C77 | Press 77-IL081 |  |  | 8541SA | 4 GA SHARPSTAR IV - TRANS RED | CFF-D-Molding 1 | Wall Thickness out of spec. 0.051". minimum 0.055" | 21B20463 | 21670.0 | Contain affected product and production performed End of Run. | 13-Aug-2021 | 13-Aug-2021 | CFF-D-Molding 1 | Wall Thickness | Machine Failure | The most probable cause is 2 of the knockout bar bolts broke which allowed the mold to shift under its own weight. |  |  | 0.0 | Investigation | INWORKS | 12-Oct-2021 |  |  | SS | Container
NC-IL081-33852 | 16-Aug-2021 | GRACE.AUSTIN | General | Internal | C19 | Press 19-IL081 |  |  | 8507SA | CONTAINER SHARPS 5QT RED | Procedure Not Followed | Missing Rim Height Inspection after mold was repair or set. | 21F07863 | 49440.0 | All affected product was contained. | 16-Aug-2021 | 16-Aug-2021 | Procedure Not Followed | Procedure Not Followed | Human Error | The employee was retrained. See Human Error Form for details. The missing inspection was completed with passing results. |  |  | 0.0 | Closure | CLOSED | 15-Sep-2021 | 31-Aug-2021 | 14. | SS | Container
NC-IL081-33860 | 17-Aug-2021 | LAKEYSHA.SIMPSON | General | Internal | 131 | Cell 131-IL081 |  |  | 8506SA | CONTAINER BIOHAZARD 5 QT | CFF-F-Welding | Weld Strength failed during scheduled operator inspection. | 21E02463 | 14400.0 | Cell went down to perform End of Run Line Clearance. All affected product was placed on hold. | 16-Aug-2021 | 16-Aug-2021 | CFF-F-Welding | Weld Strength | Incomplete, improper weld | Incomplete, improper weld |  |  | 0.0 | Implementation | INWORKS | 30-Sep-2021 |  |  | SS | Container
NC-IL081-33869 | 17-Aug-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | UW-01 OUT OF TOLERANCE AS FOUND; ADJUSTED INTO TOLERANCE; PASSED CALIBRATION AS LEFT. |  |  | Unit adjusted 08/15/2021. | 17-Aug-2021 | 15-Aug-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 17-Sep-2021 | 30-Aug-2021 | 13.
NC-IL081-33874 | 17-Aug-2021 | TOLLIN.MCCANE | General | Internal | [6 empty] | Calibration Damage Gauge | MD-660, Weight Count Scale used in the Containers focused factory. Calibration as found damaged – repairable. |  |  | Scale red-tagged and removed from service. | 17-Aug-2021 | 23-Jul-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Seized mechanical hardware | A frozen bolt in the spider mechanism found during attempted calibration 07/23/2021 due to normal equipment wear. |  |  | 0.0 | Closure | CLOSED | 17-Sep-2021 | 20-Sep-2021 | 33.
NC-IL081-33897 | 18-Aug-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | GW-002; OUT OF TOLERANCE AS FOUND; ADJUSTED INTO TOLERANCE; CALIBRATION PASSED |  |  | Unit adjusted 08/17/2021. | 18-Aug-2021 | 17-Aug-2021 | [6 empty] | 0.0 | Investigation | INWORKS
NC-IL081-33943 | 19-Aug-2021 | GRACE.AUSTIN | General | Internal | AS1 | Assembly 1-IL081 |  |  |  |  | Documentation Incorrect/Unclear | Missing step for 100% inspection of molding part on the job 31143798. BPCS Hold number 21080022. |  |  |  | 19-Aug-2021 | 19-Aug-2021 | Documentation Incorrect/Unclear | Documentation Incorrect/Unclear | Missing SWI | When creating the document for a new process, including a step to visually check the containers was overlooked. |  |  | 0.0 | Closure | CLOSED | 18-Sep-2021 | 26-Aug-2021 | 6.
NC-IL081-33974 | 20-Aug-2021 | GRACE.AUSTIN | General | Internal | C43 | Press 43-IL081 |  |  |  |  | Line Clearance not completed | Line Clearance not completed for mold change prior Start Up checklist. Mold change from R7032D to R7032E. BPCS Hold number 21080024. |  |  | Verified documentation and performed End of Run Line Clearance for Tool R7032D and Start Up Checklist for mold R7032E. No product was affected. | 20-Aug-2021 | 20-Aug-2021 | Line Clearance not completed | Line Clearance not completed | Human Error | The Production Team Lead failed to perform a line clearance on the tool because they were not aware that it was required. |  |  | 0.0 | Closure | CLOSED | 19-Sep-2021 | 07-Sep-2021 | 18.
NC-IL081-33990 | 20-Aug-2021 | ANGELICA.AGUILAR | General | Internal | C61 | Press 61-IL081 |  |  |  |  | Documentation Incorrect | Incorrect PCR was used at the Start Up of the run. R7417C instead R7417B in the press. |  |  | Correct PCR is at the press and the parameters were verified. | 20-Aug-2021 | 20-Aug-2021 | documentation |  | Human Error | Human Error |  |  | 0.0 | Closure | CLOSED | 19-Sep-2021 | 20-Sep-2021 | 30.
NC-IL081-33992 | 20-Aug-2021 | MARIA.ROSALES | General | Internal | C64 | Press 64-IL081 |  |  | WP8982 | Foam in Place Gasketed Lid, Hinged, 2 Gallon, White | Documentation Incorrect/Unclear | Using Incorrect inspection form DC00042789 should be DC00045069. Lids were used into Finish Good 8982CRH. Incorrect inspection form only used for A shift on 08/20/2021. | 21E197348' '21C40863 | 48896.0 35860.0 4480.0 | Verify the correct Inspection form per Formula sheet and change the form. | 20-Aug-2021 | 20-Aug-2021 | Documentation Incomplete/Missing | Documentation Incomplete/Missing | Human Error | Human Error | Accept | 48896.0 35860.0 4480.0 | 0.0 | Closure | CLOSED | 19-Sep-2021 | 08-Sep-2021 | 18. | G2 | Container
NC-IL081-33993 | 21-Aug-2021 | ANGELICA.AGUILAR | General | Internal | 131 | Cell 131-IL081 |  |  | 8506SA | CONTAINER BIOHAZARD 5 QT | CFF-F-Welding | During scheduled inspection, operator failed the weld strength inspection. | 21E19063 | 3580.0 | Cell went down to perform End of Run Line Clearance. All affected product has been quarantined. | 20-Aug-2021 | 20-Aug-2021 | CFF-F-Welding | Weld Strength | Incomplete, improper weld | Incomplete, improper weld | Accept | 3580.0 | 0.0 | Closure | CLOSED | 30-Sep-2021 | 14-Sep-2021 | 24. | SS | Container
NC-IL081-34086 | 25-Aug-2021 | ANGELICA.AGUILAR | General | Internal | 131 | Cell 131-IL081 |  |  | 8506SA | CONTAINER BIOHAZARD 5 QT | CFF-F-Welding | Weld Strength failed during QA Final Inspection. | 21F11763 | 3340.0 | Cell went down to perform End of Run Line Clearance. All affected product has been placed on hold. | 24-Aug-2021 | 24-Aug-2021 | CFF-F-Welding | Weld Strength | Incomplete, improper weld | Incomplete, improper weld | Accept | 3340.0 | 0.0 | Closure | CLOSED | 30-Sep-2021 | 22-Sep-2021 | 28. | SS | Container
NC-IL081-34134 | 26-Aug-2021 | GRACE.AUSTIN | General | Internal | C54 | Press 54-IL081 |  |  | 8991 | 18 GA RENEWABLES HINGE LID | CFF-F-Molding 2 | Hinge Function Failed on WP7483 Lids. Lids have been used in Finish Good 8991. | 21F20463' '21F197966 | 25.0 9640.0 | Stop press and change over new Lot. All affected product was placed on hold. | 26-Aug-2021 | 26-Aug-2021 | CFF-F-Molding 2 | Hinged Lid | Human Error | The most probable Root Cause is the associate did not properly flex the hinge lid within one minute of ejection from the press. | Accept |  | 0.0 | Waiting for Closure | INWORKS | 25-Sep-2021 |  |  | G18 | Container
NC-IL081-34136 | 26-Aug-2021 | MARIA.ROSALES | General | Internal | C46 | Press 46-IL081 |  |  | 8537SA | 3 GA SHARPSTAR - TRANS RED | CFF-V-Label Verification and Barcode Verification 8 | QA found Incorrect product code and lot number on biohazard label during scheduled inspection. | 21F12263 | 240.0 | Press went down and End of run performed. All affected product has been contained. | 26-Aug-2021 | 26-Aug-2021 | [6 empty] | 0.0 | Investigation | INWORKS |  |  |  | SS | Container
NC-IL081-34206 | 30-Aug-2021 | GRACE.AUSTIN | General | Internal | C56 | Press 56-IL081 |  |  | 8980- | 8 GA RENEWABLES HINGE LID | CFF-F-Molding 7 | QA found during scheduled inspection lid with rotor no engaged. | 21F24963 | 4880.0 | Press went down and manufacturing proceeded with the End of Run Line Clearance. All affected product has been placed on hold. | 30-Aug-2021 | 30-Aug-2021 | CFF-F-Molding 7 | Disengaged Rotor | Human Error | The most probable root cause is that the Manufacturing Associate failed to follow Step 2 of DC00035768 rev. 0 where it states to place the rotor on the lid in the fully opened position. | Accept | 180.0 4700.0 | 0.0 | Closure | CLOSED | 29-Sep-2021 | 21-Sep-2021 | 22. | G8 | Container
NC-IL081-34216 | 30-Aug-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0499, MD-0500, and MD-0502 were out of tolerance as found; subsequently adjusted back into tolerance. See attachments. |  |  | MD-0499, MD-0500, and MD-0502 adjusted back into tolerance. | 30-Aug-2021 | 22-Aug-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 07-Oct-2021 | 22-Sep-2021 | 23.
NC-IL081-34218 | 30-Aug-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0500 OUT OF TOLERANCE AS FOUND (POST-MAINTENANCE REPAIR-W.O. CL-3333975); ADJUSTED INTO TOLERANCE. |  |  | MD-0500 ADJUSTED INTO TOLERANCE. | 30-Aug-2021 | 24-Aug-2021 | Calibration Out of Tolerance | Calibration Out of Tolerance | Equipment | Equipment |  |  | 0.0 | Closure | CLOSED | 07-Oct-2021 | 22-Sep-2021 | 23.
NC-IL081-34348 | 03-Sep-2021 | MONA.ABOKAUID | General | Internal | S7C | 7C-IL081 |  |  | 8881907102 | 10ML ORAL MED SYRINGE (CLEAR) | SFF-M-V-Shorts 2 | Shorts in the pathway Barrel tip cavity #20C/ Mold 58201/MP-1113 | 21G11863X' '21G11763X | 38500.0 22000.0 | Cleared line and performed checkback | 03-Sep-2021 | 02-Sep-2021 | SFF-M-V-Shorts 2 | Shorts not affecting sterility | Machine Failure | Machine Failure |  |  | 0.0 | Waiting for Closure | INWORKS | 03-Oct-2021 |  |  | 7 | Syringe
NC-IL081-34355 | 03-Sep-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | QA-0344 OUT OF TOLERANCE AS FOUND; ADJUSTED INTO TOLERANCE. |  |  | QA-0344 ADJUSTED INTO TOLERANCE. | 03-Sep-2021 | 26-Aug-2021 | [6 empty] | 0.0 | Investigation | VOID
NC-IL081-34424 | 07-Sep-2021 | MARIA.ROSALES | General | Internal | C81 | Press 81-IL081 |  |  | 89651 | 2GA HORIZONTAL DROP CONT - RED | CFF-V-Short | Short on the Lids found During QA final Good Inspection .component WP8965N. | 21F10563' '21G198168 | 55080.0 28760.0 | Manufacturing did End of Run the product. | 07-Sep-2021 | 07-Sep-2021 | CFF-V-Short | Part is not completely formed; Missing plastic | Air valve malfunction | Work order CL-3340633 was written to clean or verify valve gates on Press 24. | Accept' 'Reject | 360.0 26814.0 74.0 1872.0 54720.0 | 0.0 | Closure | CLOSED | 07-Oct-2021 | 22-Sep-2021 | 15. | G2 | Container
NC-IL081-34489 | 09-Sep-2021 | LAKEYSHA.SIMPSON | General | Internal | C43B | Press 43B-IL081 |  |  | 1522SA | 2.2 QT PHLEBOTOMY CONT | CCF-V-Quantity per case | Packaging-Missing Components. Finished Good case were missing 16ea. lids during QA Final Case Inspection. | 21F12863 | 480.0 | Manufacturing End of the run the product. | 09-Sep-2021 | 09-Sep-2021 | CFF-V-Quantity per Case | Missing product (Containers/Lids/IFU/Other Kit items/Other Labels) | Human Error | Human Error |  |  | 0.0 | Investigation Approval | INWORKS | 09-Oct-2021 |  |  | G2 | Container
NC-IL081-34500 | 10-Sep-2021 | LAKEYSHA.SIMPSON | General | Internal | S2R | RAM-IL081 |  |  | 8881112083 | 12ML SYRINGE LUER LOCK (NS) | SFF-M-V Grease/Lint/Particulate 1 | Extraneous matter in fluid pathway. | 21J01963X | 28000.0 | Line stopped, Cleared of product,. all impacted product placed on hold. | 10-Sep-2021 | 10-Sep-2021 | SFF-M-V-Grease/Lint/Particulate 1 | Grease/Lint/Particulate Type 1: Particulate in fluid pathway | Misaligned assembly transfer | The most probable root cause is that after the cleaning was performed on 09/07/2021 the transfer tubes were not aligned properly, resulting in increased contact between the barrels and air vents. |  |  | 0.0 | Investigation Approval | INWORKS | 10-Oct-2021 |  |  | 7 | Syringe
NC-IL081-34626 | 16-Sep-2021 | MARIA.ROSALES | General | Internal | 131 | Cell 131-IL081 |  |  | WP8506 | 8507SA LID | PCR out of parameters. | Press 18 MP-1070 was running with PCR Out of parameters. Fill time was documented 3.1sec. and should be 3.2-3.4sec. | 21J02963 | 14000.0 400.0 | Cell went down to perform the End of Run Line Clearance. All affected product has been placed on hold. | 16-Sep-2021 | 16-Sep-2021 | [6 empty] | 0.0 | Investigation | INWORKS |  |  |  | SS | Container
NC-IL081-34638 | 16-Sep-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Lost Gauge | QA-0341; LOST |  |  | QA-0341, 6" Digital Caliper moved to missing status in tracking system. | 16-Sep-2021 | 07-Sep-2021 | [6 empty] | 0.0 | Investigation | INWORKS
NC-IL081-34714 | 21-Sep-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | MD-0499 PRE-VALVE CHANGE CALIBRATION, MD-0499 POST-VALVE CHANGE CALIBRATION, UW-01 POST-VALVE CHANGE CALIBRATION, AND UW-02 POST-VALVE CHANGE CALIBRATION; ALL HAD OUT OF TOLERANCE AS FOUND ISSUES. THESE ARE ALL RELATED TO EARLY CALIBRATION AND VALVE CHANGE ENGINEERING REQUESTS AND THEREFORE HAVE BEEN GROUPED TOGETHER INTHIS NC. VALVE CHANGES PERFORMED AGAINST W.O. CL-3342018 AND CL-3340507. |  |  | See Description | 21-Sep-2021 | 11-Sep-2021 | [6 empty] | 0.0 | Investigation | INWORKS
NC-IL081-34730 | 22-Sep-2021 | FREDERICK.MCNATT | General | Internal | QA | QA-IL081 |  |  |  |  | Calibration Out of Tolerance | AP-0182; OUT OF TOLERANCE AS FOUND, HIGH. CANNOT BE ADJUSTED INTO TOLERANCE; GAGE DECOMISSIONED; REPLACEMENT GAGE SR-1204-1 CALIBRATED AND INSTALLED IN PLACE OF AP-0182. |  |  |  | 22-Sep-2021 | 20-Sep-2021 | [6 empty] | 0.0 | Initiation Verification | INWORKS
NC-IL081-34746 | 22-Sep-2021 | ANGELICA.AGUILAR | General | Internal | C67 | Press 67-IL081 |  |  | 31143665 | 10 GA SHARPS CONTAINER | CFF-V-Label Verification and Barcode Verification 2 | Missing lot number and incomplete product code was found during QA final inspection. | 21J04063 | 42.0 | Press went down to perform the End of Run Line Clearance. All affected product has been quarantined. | 22-Sep-2021 | 22-Sep-2021 | [6 empty] | 0.0 | Initiation Verification | INWORKS |  |  |  | G10 | Container

## TimeSeries Data
Date | FY TAT | Open NCs | Monthly TAT | Yearly TAT | Change in NCs
2019-01-31 |  | 10 | 1 | 1
2019-02-01 |  | 11 | 1 | 1 | 1
2019-02-02 |  | 11 | 1 | 1 | 0
2019-02-03 |  | 11 | 1 | 1 | 0
2019-02-04 |  | 13 | 1 | 1 | 2
2019-02-05 |  | 14 | 1 | 1 | 1
2019-02-06 |  | 15 | 1 | 1 | 1
2019-02-07 |  | 17 | 1 | 1 | 2
2019-02-08 |  | 17 | 1 | 1 | 0
2019-02-09 |  | 19 | 1 | 1 | 2
2019-02-10 |  | 21 | 0 | 1 | 2
2019-02-11 |  | 21 | 0 | 1 | 0
2019-02-12 |  | 22 | 0 | 1 | 1
2019-02-13 |  | 25 | 0 | 1 | 3
2019-02-14 |  | 28 | 0 | 1 | 3
2019-02-15 |  | 28 | 0 | 1 | 0
2019-02-16 |  | 29 | 0 | 1 | 1
2019-02-17 |  | 29 | 0 | 1 | 0
2019-02-18 |  | 29 | 0 | 1 | 0
2019-02-19 |  | 30 | 0 | 1 | 1
2019-02-20 |  | 32 | 0 | 1 | 2
2019-02-21 |  | 33 | 16 | 11 | 1
2019-02-22 |  | 31 | 16 | 11 | -2
2019-02-23 |  | 32 | 16 | 11 | 1
2019-02-24 |  | 32 | 16 | 11 | 0
2019-02-25 |  | 33 | 18.6667 | 14.25 | 1
2019-02-26 |  | 34 | 18.6667 | 14.25 | 1
2019-02-27 |  | 34 | 19.25 | 15.6 | 0
2019-02-28 |  | 33 | 19.25 | 15.6 | -1
2019-03-01 |  | 34 | 19.25 | 15.6 | 1
2019-03-02 |  | 34 | 19.25 | 15.6 | 0
2019-03-03 |  | 35 | 19.25 | 15.6 | 1
2019-03-04 |  | 38 | 19.25 | 15.6 | 3
2019-03-05 |  | 40 | 19.25 | 15.6 | 2
2019-03-06 |  | 41 | 19.25 | 15.6 | 1
2019-03-07 |  | 41 | 19.25 | 15.6 | 0
2019-03-08 |  | 42 | 19.25 | 15.6 | 1
2019-03-09 |  | 43 | 19.25 | 15.6 | 1
2019-03-10 |  | 43 | 19.25 | 15.6 | 0
2019-03-11 |  | 44 | 19.25 | 15.6 | 1
2019-03-12 |  | 45 | 19.25 | 15.6 | 1
2019-03-13 |  | 46 | 19.25 | 15.6 | 1
2019-03-14 |  | 46 | 19.25 | 15.6 | 0
2019-03-15 |  | 47 | 19.25 | 15.6 | 1
2019-03-16 |  | 48 | 19.25 | 15.6 | 1
2019-03-17 |  | 48 | 19.25 | 15.6 | 0
2019-03-18 |  | 52 | 19.25 | 15.6 | 4
2019-03-19 |  | 53 | 19.25 | 15.6 | 1
2019-03-20 |  | 54 | 19.25 | 15.6 | 1
2019-03-21 |  | 56 | 23.8 | 20 | 2
2019-03-22 |  | 55 | 25.8571 | 22.75 | -1
2019-03-23 |  | 56 | 25.8571 | 22.75 | 1
2019-03-24 |  | 56 | 29.8 | 22.75 | 0
2019-03-25 |  | 57 | 42.625 | 34 | 1
2019-03-26 |  | 56 | 42.625 | 34 | -1
2019-03-27 |  | 57 | 35.8 | 31.6667 | 1
2019-03-28 |  | 50 | 34.7333 | 30.4211 | -7
2019-03-29 |  | 50 | 36.0625 | 31.7 | 0
2019-03-30 |  | 49 | 37.0667 | 31.7 | -1
2019-03-31 |  | 49 | 37.0667 | 31.7 | 0
2019-04-01 |  | 49 | 35.3889 | 31.087 | 0
2019-04-02 |  | 47 | 36.55 | 32.36 | -2
2019-04-03 |  | 47 | 36.55 | 32.36 | 0
2019-04-04 |  | 48 | 36.55 | 32.36 | 1
2019-04-05 |  | 50 | 36.2381 | 32.2692 | 2
2019-04-06 |  | 49 | 36.2381 | 32.2692 | -1
2019-04-07 |  | 50 | 36.2381 | 32.2692 | 1
2019-04-08 |  | 51 | 36.2381 | 32.2692 | 1
2019-04-09 |  | 53 | 36.2381 | 32.2692 | 2
2019-04-10 |  | 53 | 36.4583 | 32.8621 | 0
2019-04-11 |  | 51 | 36.4583 | 32.8621 | -2
2019-04-12 |  | 52 | 36.4583 | 32.8621 | 1
2019-04-13 |  | 53 | 36.4583 | 32.8621 | 1
2019-04-14 |  | 53 | 36.4583 | 32.8621 | 0
2019-04-15 |  | 53 | 36.2759 | 33.2353 | 0
2019-04-16 |  | 51 | 36.2759 | 33.2353 | -2
2019-04-17 |  | 52 | 37.2667 | 34.1714 | 1
2019-04-18 |  | 51 | 37.2667 | 34.1714 | -1
2019-04-19 |  | 51 | 37.4062 | 34.4595 | 0
2019-04-20 |  | 49 | 37.4062 | 34.4595 | -2
2019-04-21 |  | 50 | 37.2581 | 34.4595 | 1
2019-04-22 |  | 50 | 37.6897 | 34.4595 | 0
2019-04-23 |  | 50 | 37.6897 | 34.4595 | 0
2019-04-24 |  | 51 | 39.4839 | 36.0513 | 1
2019-04-25 |  | 50 | 36.8571 | 36.0513 | -1
2019-04-26 |  | 52 | 36.8571 | 36.0513 | 2
2019-04-27 |  | 53 | 39.8095 | 36.0513 | 1
2019-04-28 |  | 53 | 41.4 | 36.0513 | 0
2019-04-29 |  | 53 | 39 | 35.8261 | 0
2019-04-30 |  | 46 | 38.5926 | 35.6596 | -7
2019-05-01 |  | 46 | 38.5926 | 35.6596 | 0
2019-05-02 |  | 47 | 40.1154 | 35.8776 | 1
2019-05-03 |  | 45 | 39.5417 | 35.8776 | -2
2019-05-04 |  | 46 | 39.5417 | 35.8776 | 1
2019-05-05 |  | 47 | 39.5417 | 35.8776 | 1
2019-05-06 |  | 47 | 39.9565 | 35.8776 | 0
2019-05-07 |  | 47 | 44.3077 | 38.2885 | 0
2019-05-08 |  | 44 | 45.7407 | 39.1321 | -3
2019-05-09 |  | 45 | 47.5714 | 40.2037 | 1
2019-05-10 |  | 45 | 47.5714 | 40.2037 | 0
2019-05-11 |  | 46 | 48.72 | 40.2037 | 1
2019-05-12 |  | 46 | 48.72 | 40.2037 | 0
2019-05-13 |  | 47 | 48.1538 | 40.0909 | 1
2019-05-14 |  | 47 | 48.1538 | 40.0909 | 0
2019-05-15 |  | 47 | 52.1724 | 42.5172 | 0
2019-05-16 |  | 44 | 55.6667 | 42.5172 | -3
2019-05-17 |  | 44 | 53.7308 | 42.1167 | 0
2019-05-18 |  | 45 | 53.24 | 42.1167 | 1
2019-05-19 |  | 45 | 53.24 | 42.1167 | 0
2019-05-20 |  | 46 | 56.125 | 42.9836 | 1
2019-05-21 |  | 46 | 56.125 | 42.9836 | 0
2019-05-22 |  | 46 | 56.125 | 42.9836 | 0
2019-05-23 |  | 48 | 56.125 | 42.9836 | 2
2019-05-24 |  | 48 | 56.125 | 42.9836 | 0
2019-05-25 |  | 49 | 55.2727 | 42.9836 | 1
2019-05-26 |  | 49 | 55.2727 | 42.9836 | 0
2019-05-27 |  | 49 | 55.2727 | 42.9836 | 0
2019-05-28 |  | 49 | 55.2727 | 42.9836 | 0
2019-05-29 |  | 51 | 55.2727 | 42.9836 | 2
2019-05-30 |  | 53 | 64.9333 | 42.9836 | 2
2019-05-31 |  | 53 | 64.5333 | 42.6452 | 0
2019-06-01 |  | 52 | 64.5333 | 42.6452 | -1
2019-06-02 |  | 55 | 68.1538 | 42.6452 | 3
2019-06-03 |  | 56 | 68.1538 | 42.6452 | 1
2019-06-04 |  | 56 | 68.1538 | 42.6452 | 0
2019-06-05 |  | 56 | 64.4 | 42.5625 | 0
2019-06-06 |  | 56 | 64.4 | 42.5625 | 0
2019-06-07 |  | 57 | 61.0833 | 42.5625 | 1
2019-06-08 |  | 57 | 59.0909 | 42.5625 | 0
2019-06-09 |  | 57 | 55.3 | 42.5625 | 0
2019-06-10 |  | 57 | 55.3 | 42.5625 | 0
2019-06-11 |  | 59 | 55.3 | 42.5625 | 2
2019-06-12 |  | 59 | 64.6667 | 44.6515 | 0
2019-06-13 |  | 58 | 70.1538 | 45.8382 | -1
2019-06-14 |  | 56 | 70.1538 | 45.8382 | -2
2019-06-15 |  | 56 | 65.1 | 45.8382 | 0
2019-06-16 |  | 56 | 65.1 | 45.8382 | 0
2019-06-17 |  | 56 | 78.4444 | 46.8551 | 0
2019-06-18 |  | 57 | 78.4444 | 46.8551 | 1
2019-06-19 |  | 57 | 84.3571 | 50.1081 | 0
2019-06-20 |  | 53 | 83.0714 | 50.4667 | -4
2019-06-21 |  | 53 | 75.875 | 49.8182 | 0
2019-06-22 |  | 51 | 75.875 | 49.8182 | -2
2019-06-23 |  | 52 | 75.875 | 49.8182 | 1
2019-06-24 |  | 52 | 71.8095 | 50.3659 | 0
2019-06-25 |  | 49 | 69.7273 | 50.0723 | -3
2019-06-26 |  | 48 | 66.0833 | 49.5059 | -1
2019-06-27 |  | 47 | 66.0833 | 49.5059 | -1
2019-06-28 |  | 47 | 66.0833 | 49.5059 | 0
2019-06-29 |  | 47 | 66.0833 | 49.5059 | 0
2019-06-30 |  | 47 | 66.0833 | 49.5059 | 0
2019-07-01 |  | 47 | 68 | 49.5059 | 0
2019-07-02 |  | 48 | 68.3333 | 49.814 | 1
2019-07-03 |  | 49 | 69.4 | 50.3333 | 1
2019-07-04 |  | 48 | 69.4 | 50.3333 | -1
2019-07-05 |  | 49 | 69.4 | 50.3333 | 1
2019-07-06 |  | 49 | 71.9565 | 50.3333 | 0
2019-07-07 |  | 49 | 71.9565 | 50.3333 | 0
2019-07-08 |  | 49 | 71.9565 | 50.3333 | 0
2019-07-09 |  | 50 | 71.9565 | 50.3333 | 1
2019-07-10 |  | 51 | 71.9565 | 50.3333 | 1
2019-07-11 |  | 53 | 71.9565 | 50.3333 | 2
2019-07-12 |  | 57 | 71.9565 | 50.3333 | 4
2019-07-13 |  | 57 | 68.1905 | 50.3333 | 0
2019-07-14 |  | 57 | 66.4211 | 50.3333 | 0
2019-07-15 |  | 57 | 66.4211 | 50.3333 | 0
2019-07-16 |  | 59 | 66.4211 | 50.3333 | 2
2019-07-17 |  | 61 | 64.6667 | 50.2809 | 2
2019-07-18 |  | 59 | 59.8571 | 49.8889 | -2
2019-07-19 |  | 58 | 62.4231 | 51.1158 | -1
2019-07-20 |  | 54 | 54.6667 | 51.1158 | -4
2019-07-21 |  | 54 | 53.55 | 51.1158 | 0
2019-07-22 |  | 54 | 60.5238 | 52.1122 | 0
2019-07-23 |  | 52 | 60.5238 | 52.1122 | -2
2019-07-24 |  | 52 | 58.6364 | 51.7778 | 0
2019-07-25 |  | 53 | 58.5882 | 51.7778 | 1
2019-07-26 |  | 54 | 60.8824 | 51.91 | 1
2019-07-27 |  | 53 | 65.5333 | 51.91 | -1
2019-07-28 |  | 53 | 65.5333 | 51.91 | 0
2019-07-29 |  | 55 | 65.5333 | 51.91 | 2
2019-07-30 |  | 56 | 65.2941 | 52.1373 | 1
2019-07-31 |  | 55 | 64.1429 | 52.4057 | -1
2019-08-01 |  | 52 | 64.1429 | 52.4057 | -3
2019-08-02 |  | 61 | 63.55 | 52.4057 | 9
2019-08-03 |  | 62 | 61.8947 | 52.4057 | 1
2019-08-04 |  | 62 | 61.8947 | 52.4057 | 0
2019-08-05 |  | 62 | 61.8947 | 52.4057 | 0
2019-08-06 |  | 63 | 61.8947 | 52.4057 | 1
2019-08-07 |  | 64 | 61.8947 | 52.4057 | 1
2019-08-08 |  | 64 | 61.8947 | 52.4057 | 0
2019-08-09 |  | 68 | 61.8947 | 52.4057 | 4
2019-08-10 |  | 69 | 61.8947 | 52.4057 | 1
2019-08-11 |  | 69 | 61.8947 | 52.4057 | 0
2019-08-12 |  | 71 | 65.7619 | 53.3333 | 2
2019-08-13 |  | 70 | 65.7619 | 53.3333 | -1
2019-08-14 |  | 71 | 65.6818 | 53.4312 | 1
2019-08-15 |  | 70 | 64.4348 | 53.2818 | -1
2019-08-16 |  | 69 | 64.4348 | 53.2818 | -1
2019-08-17 |  | 70 | 66 | 53.2818 | 1
2019-08-18 |  | 70 | 68.55 | 53.2818 | 0
2019-08-19 |  | 72 | 67 | 53.2818 | 2
2019-08-20 |  | 75 | 67 | 53.2818 | 3
2019-08-21 |  | 75 | 77.6316 | 55.5351 | 0
2019-08-22 |  | 71 | 76.5 | 55.5351 | -4
2019-08-23 |  | 72 | 72.8235 | 55.1739 | 1
2019-08-24 |  | 73 | 76.1875 | 55.1739 | 1
2019-08-25 |  | 73 | 76.1875 | 55.1739 | 0
2019-08-26 |  | 75 | 75.0526 | 55.605 | 2
2019-08-27 |  | 71 | 73.6 | 55.525 | -4
2019-08-28 |  | 71 | 76.6957 | 56.5447 | 0
2019-08-29 |  | 68 | 76.6957 | 56.5447 | -3
2019-08-30 |  | 69 | 77.9524 | 56.5447 | 1
2019-08-31 |  | 69 | 82.3529 | 56.5447 | 0
2019-09-01 |  | 70 | 82.3529 | 56.5447 | 1
2019-09-02 |  | 70 | 82.3529 | 56.5447 | 0
2019-09-03 |  | 71 | 81.4211 | 56.816 | 1
2019-09-04 |  | 71 | 77.3333 | 57.0077 | 0
2019-09-05 |  | 70 | 77.5 | 57.3485 | -1
2019-09-06 |  | 70 | 68.375 | 56.1087 | 0
2019-09-07 |  | 64 | 68.375 | 56.1087 | -6
2019-09-08 |  | 64 | 68.375 | 56.1087 | 0
2019-09-09 |  | 66 | 69.5 | 56.5571 | 2
2019-09-10 |  | 65 | 69.5 | 56.5571 | -1
2019-09-11 |  | 66 | 67.2368 | 56.3194 | 1
2019-09-12 |  | 62 | 63.8205 | 56.1156 | -4
2019-09-13 |  | 60 | 62.9512 | 55.9799 | -2
2019-09-14 |  | 58 | 62.925 | 55.9799 | -2
2019-09-15 |  | 58 | 63.5897 | 55.9799 | 0
2019-09-16 |  | 58 | 63.15 | 55.9133 | 0
2019-09-17 |  | 57 | 63.4634 | 56.0464 | -1
2019-09-18 |  | 58 | 62.2791 | 55.8105 | 1
2019-09-19 |  | 58 | 62.1364 | 55.8117 | 0
2019-09-20 |  | 58 | 62.1364 | 55.8117 | 0
2019-09-21 |  | 58 | 56.6 | 55.8117 | 0
2019-09-22 |  | 59 | 56.6 | 55.8117 | 1
2019-09-23 |  | 59 | 58.3256 | 56.0316 | 0
2019-09-24 |  | 57 | 58.3256 | 56.0316 | -2
2019-09-25 |  | 57 | 58.5909 | 56.1195 | 0
2019-09-26 |  | 57 | 61.8571 | 57.236 | 0
2019-09-27 |  | 58 | 62.2439 | 57.236 | 1
2019-09-28 |  | 58 | 59.4737 | 57.236 | 0
2019-09-29 |  | 58 | 59.4737 | 57.236 | 0
2019-09-30 |  | 59 | 59.2821 | 57.2037 | 1
2019-10-01 |  | 60 | 59.2821 | 57.2037 | 1
2019-10-02 |  | 61 | 59.5854 | 57.3049 | 1
2019-10-03 |  | 66 | 58.3542 | 57.0526 | 5
2019-10-04 |  | 59 | 57.6957 | 57.0526 | -7
2019-10-05 |  | 59 | 57.1951 | 57.0526 | 0
2019-10-06 |  | 61 | 56.0513 | 57.0526 | 2
2019-10-07 |  | 63 | 61 | 57.0526 | 2
2019-10-08 |  | 65 | 65.0556 | 57.9598 | 2
2019-10-09 |  | 62 | 64.0541 | 57.7886 | -3
2019-10-10 |  | 65 | 59.9474 | 57.2809 | 3
2019-10-11 |  | 63 | 58.3415 | 56.9613 | -2
2019-10-12 |  | 61 | 59.4595 | 56.9613 | -2
2019-10-13 |  | 61 | 60.6176 | 56.9613 | 0
2019-10-14 |  | 62 | 58.0286 | 56.3696 | 1
2019-10-15 |  | 60 | 58.4595 | 56.4731 | -2
2019-10-16 |  | 65 | 54.5714 | 55.6702 | 5
2019-10-17 |  | 61 | 52.3409 | 55.1031 | -4
2019-10-18 |  | 60 | 51.2444 | 54.9439 | -1
2019-10-19 |  | 58 | 51.8605 | 54.9439 | -2
2019-10-20 |  | 58 | 51.7619 | 54.9439 | 0
2019-10-21 |  | 58 | 52.0227 | 54.9697 | 0
2019-10-22 |  | 59 | 52.0227 | 54.9697 | 1
2019-10-23 |  | 59 | 52.0227 | 54.9697 | 0
2019-10-24 |  | 59 | 49.2143 | 54.6 | 0
2019-10-25 |  | 57 | 50.7955 | 54.8911 | -2
2019-10-26 |  | 57 | 50.3488 | 54.8911 | 0
2019-10-27 |  | 57 | 45.6829 | 54.8911 | 0
2019-10-28 |  | 60 | 44.9286 | 54.6897 | 3
2019-10-29 |  | 60 | 44.9286 | 54.6897 | 0
2019-10-30 |  | 60 | 45.3721 | 54.7353 | 0
2019-10-31 |  | 59 | 46.6667 | 54.913 | -1
2019-11-01 |  | 58 | 45.9783 | 54.7212 | -1
2019-11-02 |  | 57 | 45.0909 | 54.7212 | -1
2019-11-03 |  | 57 | 43.9459 | 54.7212 | 0
2019-11-04 |  | 58 | 44.359 | 54.6952 | 1
2019-11-05 |  | 57 | 43.875 | 54.5545 | -1
2019-11-06 |  | 57 | 43.4634 | 54.4245 | 0
2019-11-07 |  | 59 | 43.4634 | 54.4245 | 2
2019-11-08 |  | 59 | 38.2368 | 54.4245 | 0
2019-11-09 |  | 59 | 38.5135 | 54.4245 | 0
2019-11-10 |  | 59 | 39.4706 | 54.4245 | 0
2019-11-11 |  | 61 | 39.6129 | 54.4245 | 2
2019-11-12 |  | 62 | 37.7647 | 53.9256 | 1
2019-11-13 |  | 61 | 37.7647 | 53.9256 | -1
2019-11-14 |  | 63 | 39.4194 | 53.9256 | 2
2019-11-15 |  | 63 | 36.0625 | 53.4771 | 0
2019-11-16 |  | 60 | 37.963 | 53.4771 | -3
2019-11-17 |  | 60 | 40.3333 | 53.4771 | 0
2019-11-18 |  | 62 | 40.4091 | 53.4771 | 2
2019-11-19 |  | 64 | 40.4091 | 53.4771 | 2
2019-11-20 |  | 65 | 38.6957 | 53.2374 | 1
2019-11-21 |  | 64 | 36.9048 | 53.2374 | -1
2019-11-22 |  | 65 | 36.9048 | 53.2374 | 1
2019-11-23 |  | 65 | 36.9048 | 53.2374 | 0
2019-11-24 |  | 66 | 38.8947 | 53.2374 | 1
2019-11-25 |  | 67 | 33.5882 | 53.2374 | 1
2019-11-26 |  | 68 | 32.4444 | 53.0545 | 1
2019-11-27 |  | 67 | 32.4444 | 53.0545 | -1
2019-11-28 |  | 67 | 33.5294 | 53.0545 | 0
2019-11-29 |  | 67 | 33.5294 | 53.0545 | 0
2019-11-30 |  | 68 | 31.625 | 53.0545 | 1
2019-12-01 |  | 69 | 23.4615 | 53.0545 | 1
2019-12-02 |  | 71 | 24.1667 | 53.0545 | 2
2019-12-03 |  | 72 | 24.1667 | 53.0545 | 1
2019-12-04 |  | 74 | 21.0714 | 52.5991 | 2
2019-12-05 |  | 72 | 15.9167 | 52.5991 | -2
2019-12-06 |  | 73 | 15.0909 | 52.5991 | 1
2019-12-07 |  | 74 | 13.9 | 52.5991 | 1
2019-12-08 |  | 74 | 13.9 | 52.5991 | 0
2019-12-09 |  | 76 | 13.9 | 52.5991 | 2
2019-12-10 |  | 77 | 13.9 | 52.5991 | 1
2019-12-11 |  | 77 | 13.9 | 52.5991 | 0
2019-12-12 |  | 77 | 13.9 | 52.5991 | 0
2019-12-13 |  | 77 | 20.875 | 52.7399 | 0
2019-12-14 |  | 78 | 20.875 | 52.7399 | 1
2019-12-15 |  | 80 | 20.875 | 52.7399 | 2
2019-12-16 |  | 80 | 26.1429 | 52.6267 | 0
2019-12-17 |  | 78 | 26.1429 | 52.6267 | -2
2019-12-18 |  | 78 | 26.1429 | 52.6267 | 0
2019-12-19 |  | 79 | 26.1429 | 52.6267 | 1
2019-12-20 |  | 79 | 39.75 | 52.9912 | 0
2019-12-21 |  | 80 | 45.2857 | 52.9912 | 1
2019-12-22 |  | 82 | 45.2857 | 52.9912 | 2
2019-12-23 |  | 82 | 42.3333 | 52.807 | 0
2019-12-24 |  | 80 | 42.3333 | 52.807 | -2
2019-12-25 |  | 80 | 42.3333 | 52.807 | 0
2019-12-26 |  | 81 | 42.3333 | 52.807 | 1
2019-12-27 |  | 81 | 49.7778 | 52.9258 | 0
2019-12-28 |  | 80 | 49.7778 | 52.9258 | -1
2019-12-29 |  | 82 | 49.7778 | 52.9258 | 2
2019-12-30 |  | 83 | 49.7778 | 52.9258 | 1
2019-12-31 |  | 83 | 49.7778 | 52.9258 | 0
2020-01-01 |  | 83 | 49.7778 | 52.9258 | 0
2020-01-02 |  | 84 | 49.7778 | 52.9258 | 1
2020-01-03 |  | 86 | 49.7778 | 52.9258 | 2
2020-01-04 |  | 86 | 63.2857 | 52.9258 | 0
2020-01-05 |  | 86 | 63.2857 | 52.9258 | 0
2020-01-06 |  | 86 | 63.625 | 52.9826 | 0
2020-01-07 |  | 86 | 65.4615 | 53.3106 | 0
2020-01-08 |  | 81 | 63.3571 | 53.2373 | -5
2020-01-09 |  | 80 | 64.3529 | 53.4351 | -1
2020-01-10 |  | 79 | 63.7895 | 53.4813 | -1
2020-01-11 |  | 77 | 63.7895 | 53.7 | -2
2020-01-12 |  | 77 | 63.7895 | 53.7 | 0
2020-01-13 |  | 77 | 68.5238 | 54.3169 | 0
2020-01-14 |  | 75 | 65.8636 | 54.1352 | -2
2020-01-15 |  | 75 | 67.25 | 54.3659 | 0
2020-01-16 |  | 73 | 69.8261 | 54.4372 | -2
2020-01-17 |  | 73 | 69.8261 | 54.4372 | 0
2020-01-18 |  | 74 | 69.8261 | 54.4372 | 1
2020-01-19 |  | 74 | 69.8261 | 54.4372 | 0
2020-01-20 |  | 74 | 63.6667 | 54.2289 | 0
2020-01-21 |  | 72 | 63.6667 | 54.2289 | -2
2020-01-22 |  | 72 | 65 | 54.4422 | 0
2020-01-23 |  | 70 | 65.36 | 54.2579 | -2
2020-01-24 |  | 71 | 65.36 | 54.2579 | 1
2020-01-25 |  | 71 | 65.36 | 54.2579 | 0
2020-01-26 |  | 74 | 65.36 | 54.2579 | 3
2020-01-27 |  | 76 | 64.75 | 54.2579 | 2
2020-01-28 |  | 77 | 63.1538 | 54.1772 | 1
2020-01-29 |  | 77 | 62.0741 | 54.098 | 0
2020-01-30 |  | 77 | 62.0741 | 54.098 | 0
2020-01-31 |  | 77 | 62.0741 | 54.098 | 0
2020-02-01 |  | 78 | 62.0741 | 54.098 | 1
2020-02-02 |  | 80 | 62.0741 | 54.098 | 2
2020-02-03 |  | 80 | 59.9643 | 53.8984 | 0
2020-02-04 |  | 79 | 59.9643 | 53.8984 | -1
2020-02-05 |  | 80 | 61.2759 | 54.07 | 1
2020-02-06 |  | 80 | 61.1071 | 54.07 | 0
2020-02-07 |  | 80 | 59.5217 | 54.07 | 0
2020-02-08 |  | 80 | 60.5909 | 54.07 | 0
2020-02-09 |  | 82 | 59.2632 | 54.07 | 2
2020-02-10 |  | 82 | 61.1053 | 54.2432 | 0
2020-02-11 |  | 81 | 61.1053 | 54.2432 | -1
2020-02-12 |  | 82 | 64.45 | 54.5269 | 1
2020-02-13 |  | 84 | 57.5294 | 54.5269 | 2
2020-02-14 |  | 84 | 60.5 | 54.5269 | 0
2020-02-15 |  | 84 | 57.3571 | 54.5269 | 0
2020-02-16 |  | 84 | 56.2308 | 54.5269 | 0
2020-02-17 |  | 85 | 52.3333 | 54.3168 | 1
2020-02-18 |  | 84 | 57.4706 | 54.6326 | -1
2020-02-19 |  | 82 | 58.3889 | 54.7057 | -2
2020-02-20 |  | 81 | 62.0588 | 54.7293 | -1
2020-02-21 |  | 80 | 62.4444 | 54.7828 | -1
2020-02-22 |  | 79 | 60.125 | 55.0755 | -1
2020-02-23 |  | 81 | 63.6 | 55.0755 | 2
2020-02-24 |  | 82 | 60.3125 | 54.9098 | 1
2020-02-25 |  | 81 | 58.5882 | 54.8202 | -1
2020-02-26 |  | 80 | 58.5882 | 54.9361 | -1
2020-02-27 |  | 81 | 58.5882 | 54.9361 | 1
2020-02-28 |  | 82 | 59.2222 | 55.0373 | 1
2020-02-29 |  | 79 | 60.7059 | 55.0373 | -3
2020-03-01 |  | 80 | 60.7059 | 55.0373 | 1
2020-03-02 |  | 80 | 60.7059 | 55.0373 | 0
2020-03-03 |  | 80 | 60.7059 | 55.0373 | 0
2020-03-04 |  | 81 | 62.0526 | 55.1741 | 1
2020-03-05 |  | 81 | 63.8947 | 55.1107 | 0
2020-03-06 |  | 80 | 63.8947 | 55.1107 | -1
2020-03-07 |  | 80 | 62 | 55.1107 | 0
2020-03-08 |  | 80 | 62 | 55.1107 | 0
2020-03-09 |  | 80 | 60.0526 | 55 | 0
2020-03-10 |  | 80 | 60.0526 | 55 | 0
2020-03-11 |  | 81 | 60.0526 | 55 | 1
2020-03-12 |  | 82 | 58.1364 | 55.0578 | 1
2020-03-13 |  | 78 | 58.1364 | 55.0578 | -4
2020-03-14 |  | 79 | 54.8095 | 55.0578 | 1
2020-03-15 |  | 79 | 54.8095 | 55.0578 | 0
2020-03-16 |  | 82 | 54.8095 | 55.0578 | 3
2020-03-17 |  | 82 | 57 | 55.2302 | 0
2020-03-18 |  | 86 | 57 | 55.2302 | 4
2020-03-19 |  | 87 | 60 | 55.2302 | 1
2020-03-20 |  | 87 | 56 | 55.2302 | 0
2020-03-21 |  | 87 | 54.9412 | 55.278 | 0
2020-03-22 |  | 87 | 54.5625 | 55.4545 | 0
2020-03-23 |  | 88 | 63.4348 | 56.2014 | 1
2020-03-24 |  | 81 | 63.4348 | 56.2014 | -7
2020-03-25 |  | 84 | 63.4348 | 56.1179 | 3
2020-03-26 |  | 87 | 59 | 55.6303 | 3
2020-03-27 |  | 84 | 87.2581 | 59.3816 | -3
2020-03-28 |  | 78 | 87.2581 | 59.5638 | -6
2020-03-29 |  | 78 | 87.2581 | 59.5765 | 0
2020-03-30 |  | 79 | 90.9643 | 59.5765 | 1
2020-03-31 |  | 80 | 90.9643 | 59.5765 | 1
2020-04-01 |  | 82 | 88.5862 | 59.7921 | 2
2020-04-02 |  | 81 | 86.2581 | 59.8315 | -1
2020-04-03 |  | 79 | 86.4571 | 60.2297 | -2
2020-04-04 |  | 75 | 87.2424 | 60.2297 | -4
2020-04-05 |  | 76 | 88.7812 | 60.3369 | 1
2020-04-06 |  | 76 | 88.7812 | 60.3369 | 0
2020-04-07 |  | 77 | 88.7812 | 60.3369 | 1
2020-04-08 |  | 77 | 86.9444 | 60.5035 | 0
2020-04-09 |  | 74 | 89.6944 | 60.7247 | -3
2020-04-10 |  | 73 | 88.641 | 61.122 | -1
2020-04-11 |  | 70 | 88.641 | 61.122 | -3
2020-04-12 |  | 70 | 93.1176 | 61.122 | 0
2020-04-13 |  | 72 | 90.6944 | 61.0415 | 2
2020-04-14 |  | 71 | 88.3784 | 60.8483 | -1
2020-04-15 |  | 71 | 86.225 | 61.2778 | 0
2020-04-16 |  | 69 | 84.9756 | 61.1869 | -2
2020-04-17 |  | 68 | 85.9524 | 61.5379 | -1
2020-04-18 |  | 67 | 85.9524 | 61.5379 | -1
2020-04-19 |  | 68 | 85.9524 | 61.691 | 1
2020-04-20 |  | 70 | 85.9524 | 61.691 | 2
2020-04-21 |  | 72 | 85.8837 | 61.7647 | 2
2020-04-22 |  | 72 | 85.8837 | 61.7647 | 0
2020-04-23 |  | 73 | 86.8 | 61.7647 | 1
2020-04-24 |  | 73 | 87.1081 | 61.9516 | 0
2020-04-25 |  | 71 | 87.1081 | 61.9516 | -2
2020-04-26 |  | 71 | 95.0606 | 61.9516 | 0
2020-04-27 |  | 73 | 71.6667 | 61.9516 | 2
2020-04-28 |  | 75 | 71.9286 | 62.0103 | 2
2020-04-29 |  | 74 | 71.9286 | 62.689 | -1
2020-04-30 |  | 75 | 71.9286 | 62.8121 | 1
2020-05-01 |  | 76 | 67.5 | 62.4085 | 1
2020-05-02 |  | 74 | 69.069 | 62.5603 | -2
2020-05-03 |  | 74 | 70.2963 | 62.5603 | 0
2020-05-04 |  | 74 | 63.75 | 62.3031 | 0
2020-05-05 |  | 71 | 62.931 | 62.2257 | -3
2020-05-06 |  | 70 | 62.2 | 62.1522 | -1
2020-05-07 |  | 71 | 62.2 | 61.9895 | 1
2020-05-08 |  | 71 | 65.6667 | 62.316 | 0
2020-05-09 |  | 69 | 64.7586 | 62.1951 | -2
2020-05-10 |  | 69 | 62.6429 | 62.1951 | 0
2020-05-11 |  | 70 | 60.3704 | 62.1246 | 1
2020-05-12 |  | 69 | 60.3704 | 62.1246 | -1
2020-05-13 |  | 69 | 60.3704 | 62.2222 | 0
2020-05-14 |  | 69 | 60.5 | 62.1522 | 0
2020-05-15 |  | 68 | 68.4615 | 62.4146 | -1
2020-05-16 |  | 71 | 69.6087 | 62.4146 | 3
2020-05-17 |  | 71 | 71.1818 | 62.6386 | 0
2020-05-18 |  | 71 | 66.85 | 62.6386 | 0
2020-05-19 |  | 71 | 66.85 | 62.6386 | 0
2020-05-20 |  | 72 | 64.7143 | 62.3825 | 1
2020-05-21 |  | 71 | 62 | 62.1818 | -1
2020-05-22 |  | 71 | 61 | 62.1818 | 0
2020-05-23 |  | 71 | 61 | 62.1818 | 0
2020-05-24 |  | 72 | 61 | 62.1818 | 1
2020-05-25 |  | 72 | 57.6842 | 62.1818 | 0
2020-05-26 |  | 73 | 57.6842 | 62.1818 | 1
2020-05-27 |  | 74 | 57.65 | 62.1638 | 1
2020-05-28 |  | 73 | 57.65 | 62.1638 | -1
2020-05-29 |  | 74 | 68.6522 | 63.0447 | 1
2020-05-30 |  | 70 | 68.6522 | 63.0447 | -4
2020-05-31 |  | 70 | 68.6522 | 63.1862 | 0
2020-06-01 |  | 70 | 71.2917 | 63.0273 | 0
2020-06-02 |  | 68 | 71.2917 | 63.0273 | -2
2020-06-03 |  | 69 | 68.72 | 62.8367 | 1
2020-06-04 |  | 69 | 71.7619 | 62.7186 | 0
2020-06-05 |  | 70 | 72.381 | 62.8401 | 1
2020-06-06 |  | 69 | 73.95 | 62.8401 | -1
2020-06-07 |  | 69 | 73.95 | 62.8401 | 0
2020-06-08 |  | 69 | 65.5556 | 62.6339 | 0
2020-06-09 |  | 69 | 65.5556 | 62.6339 | 0
2020-06-10 |  | 69 | 64.3684 | 62.5676 | 0
2020-06-11 |  | 68 | 65.8235 | 62.5676 | -1
2020-06-12 |  | 72 | 65.8235 | 62.2347 | 4
2020-06-13 |  | 72 | 65.8235 | 62.0788 | 0
2020-06-14 |  | 72 | 67.3125 | 62.0788 | 0
2020-06-15 |  | 73 | 57.6667 | 62.0788 | 1
2020-06-16 |  | 73 | 57.6667 | 62.0788 | 0
2020-06-17 |  | 73 | 57.8824 | 61.8771 | 0
2020-06-18 |  | 71 | 57.8824 | 61.8771 | -2
2020-06-19 |  | 71 | 56.3158 | 61.1759 | 0
2020-06-20 |  | 70 | 58.2222 | 61.1211 | -1
2020-06-21 |  | 71 | 61.3529 | 61.3693 | 1
2020-06-22 |  | 71 | 58.0526 | 61.1522 | 0
2020-06-23 |  | 69 | 56.2381 | 61 | -2
2020-06-24 |  | 67 | 54.4348 | 60.8611 | -2
2020-06-25 |  | 66 | 54.4348 | 60.9826 | -1
2020-06-26 |  | 67 | 55.4 | 61.2648 | 1
2020-06-27 |  | 65 | 55.3333 | 61.2648 | -2
2020-06-28 |  | 65 | 55.3333 | 61.2648 | 0
2020-06-29 |  | 65 | 40.7619 | 61.1667 | 0
2020-06-30 |  | 66 | 40.7619 | 61.1667 | 1
2020-07-01 |  | 66 | 38.7391 | 60.8655 | 0
2020-07-02 |  | 64 | 37.4 | 60.8131 | -2
2020-07-03 |  | 64 | 37.4 | 60.6944 | 0
2020-07-04 |  | 64 | 39 | 60.6944 | 0
2020-07-05 |  | 64 | 39.6111 | 60.6944 | 0
2020-07-06 |  | 66 | 39.8333 | 60.6817 | 2
2020-07-07 |  | 65 | 39.8333 | 60.6817 | -1
2020-07-08 |  | 66 | 39.8333 | 60.6817 | 1
2020-07-09 |  | 66 | 44.3889 | 60.7621 | 0
2020-07-10 |  | 65 | 44.3889 | 60.7621 | -1
2020-07-11 |  | 65 | 44.4706 | 60.7621 | 0
2020-07-12 |  | 65 | 44.4706 | 60.7621 | 0
2020-07-13 |  | 65 | 55.7368 | 61.3836 | 0
2020-07-14 |  | 63 | 55.7368 | 61.3836 | -2
2020-07-15 |  | 63 | 67.5455 | 62.2068 | 0
2020-07-16 |  | 61 | 65.0833 | 62.0438 | -2
2020-07-17 |  | 59 | 61.8148 | 61.8725 | -2
2020-07-18 |  | 59 | 62 | 62.0303 | 0
2020-07-19 |  | 59 | 62 | 61.839 | 0
2020-07-20 |  | 59 | 63.6522 | 61.839 | 0
2020-07-21 |  | 59 | 63.6522 | 61.839 | 0
2020-07-22 |  | 59 | 62.3333 | 61.5103 | 0
2020-07-23 |  | 58 | 61.8333 | 61.2534 | -1
2020-07-24 |  | 57 | 61.3913 | 61.2089 | -1
2020-07-25 |  | 56 | 63.8571 | 61.2089 | -1
2020-07-26 |  | 56 | 63.8571 | 61.1959 | 0
2020-07-27 |  | 56 | 69.5714 | 61.6416 | 0
2020-07-28 |  | 55 | 67.3636 | 61.5034 | -1
2020-07-29 |  | 54 | 64.6087 | 61.3085 | -1
2020-07-30 |  | 54 | 62.25 | 61.0169 | 0
2020-07-31 |  | 52 | 62.25 | 61.0412 | -2
2020-08-01 |  | 53 | 66.3182 | 61.0412 | 1
2020-08-02 |  | 53 | 66.3182 | 61.0412 | 0
2020-08-03 |  | 54 | 68.4583 | 61.2526 | 1
2020-08-04 |  | 53 | 63.9259 | 60.9122 | -1
2020-08-05 |  | 52 | 62.5862 | 60.802 | -1
2020-08-06 |  | 50 | 62.7857 | 60.802 | -2
2020-08-07 |  | 52 | 62.069 | 60.7391 | 2
2020-08-08 |  | 51 | 62.069 | 60.7391 | -1
2020-08-09 |  | 51 | 61.2857 | 60.7391 | 0
2020-08-10 |  | 53 | 61.2857 | 60.7391 | 2
2020-08-11 |  | 53 | 61.2857 | 60.7391 | 0
2020-08-12 |  | 54 | 61.2857 | 60.4579 | 1
2020-08-13 |  | 55 | 51.0357 | 60.107 | 1
2020-08-14 |  | 53 | 52.2581 | 60.1296 | -2
2020-08-15 |  | 50 | 42.6071 | 60.2067 | -3
2020-08-16 |  | 50 | 42.9615 | 60.2067 | 0
2020-08-17 |  | 53 | 177.1897 | 81.5791 | 3
2020-08-18 |  | 19 | 175.8644 | 81.631 | -34
2020-08-19 |  | 18 | 170.4918 | 81.2189 | -1
2020-08-20 |  | 18 | 166.6825 | 81.0382 | 0
2020-08-21 |  | 16 | 166.6825 | 80.6042 | -2
2020-08-22 |  | 16 | 168.8548 | 80.6042 | 0
2020-08-23 |  | 16 | 173.6833 | 80.803 | 0
2020-08-24 |  | 16 | 176.5254 | 80.803 | 0
2020-08-25 |  | 16 | 176.5254 | 80.803 | 0
2020-08-26 |  | 17 | 174.8689 | 81.2282 | 1
2020-08-27 |  | 16 | 176.5085 | 81.3343 | -1
2020-08-28 |  | 16 | 176.2881 | 80.9667 | 0
2020-08-29 |  | 16 | 179.2586 | 80.9667 | 0
2020-08-30 |  | 16 | 184.9286 | 80.9667 | 0
2020-08-31 |  | 17 | 184.9286 | 80.9667 | 1
2020-09-01 |  | 17 | 184.9286 | 80.9667 | 0
2020-09-02 |  | 17 | 184.9286 | 80.9667 | 0
2020-09-03 |  | 17 | 188.3704 | 81.0122 | 0
2020-09-04 |  | 17 | 194.1538 | 81.0802 | 0
2020-09-05 |  | 16 | 200.14 | 81.0901 | -1
2020-09-06 |  | 16 | 200.14 | 82.0823 | 0
2020-09-07 |  | 16 | 203.3673 | 82.0823 | 0
2020-09-08 |  | 16 | 203.3673 | 82.0823 | 0
2020-09-09 |  | 17 | 193.4615 | 81.571 | 1
2020-09-10 |  | 14 | 193.4615 | 81.571 | -3
2020-09-11 |  | 16 | 193.4615 | 82 | 2
2020-09-12 |  | 16 | 193.4615 | 82.3452 | 0
2020-09-13 |  | 16 | 200.88 | 82.5812 | 0
2020-09-14 |  | 17 | 209.6383 | 82.5812 | 1
2020-09-15 |  | 17 | 209.6383 | 82.5812 | 0
2020-09-16 |  | 17 | 207.8367 | 83.2362 | 0
2020-09-17 |  | 15 | 62.9333 | 83.0777 | -2
2020-09-18 |  | 17 | 60.3571 | 83.3713 | 2
2020-09-19 |  | 17 | 68.4167 | 83.4608 | 0
2020-09-20 |  | 17 | 72 | 83.4608 | 0
2020-09-21 |  | 17 | 66 | 83.2085 | 0
2020-09-22 |  | 17 | 66 | 83.2085 | 0
2020-09-23 |  | 18 | 66 | 83.4554 | 1
2020-09-24 |  | 18 | 66 | 83.4554 | 0
2020-09-25 |  | 19 | 55.4286 | 82.8426 | 1
2020-09-26 |  | 16 | 43.6667 | 82.4257 | -3
2020-09-27 |  | 16 | 43.6667 | 82.4257 | 0
2020-09-28 |  | 16 | 46.9091 | 82.4257 | 0
2020-09-29 |  | 16 | 56.2308 | 82.5902 | 0
2020-09-30 |  | 15 | 52.6429 | 82.4393 | -1
2020-10-01 |  | 14 | 52.6429 | 82.4393 | -1
2020-10-02 |  | 15 | 49.8 | 82.3125 | 1
2020-10-03 |  | 14 | 49.8 | 83.0471 | -1
2020-10-04 |  | 14 | 49.8 | 83.0471 | 0
2020-10-05 |  | 14 | 52.8571 | 83.0471 | 0
2020-10-06 |  | 16 | 53.4667 | 82.9765 | 2
2020-10-07 |  | 16 | 48.1765 | 82.48 | 0
2020-10-08 |  | 16 | 44.5789 | 81.7492 | 0
2020-10-09 |  | 16 | 47.3333 | 81.8733 | 0
2020-10-10 |  | 14 | 49.9444 | 82.4209 | -2
2020-10-11 |  | 14 | 49.9444 | 82.8741 | 0
2020-10-12 |  | 14 | 49.9444 | 82.8741 | 0
2020-10-13 |  | 15 | 49.9444 | 82.8741 | 1
2020-10-14 |  | 17 | 49.9444 | 83.5155 | 2
2020-10-15 |  | 18 | 49.9444 | 83.6367 | 1
2020-10-16 |  | 18 | 49.9444 | 84.6549 | 0
2020-10-17 |  | 18 | 35.5 | 85.3559 | 0
2020-10-18 |  | 18 | 36.0667 | 85.6846 | 0
2020-10-19 |  | 18 | 34.625 | 85.425 | 0
2020-10-20 |  | 18 | 34.625 | 85.425 | 0
2020-10-21 |  | 19 | 33.3333 | 85.1786 | 1
2020-10-22 |  | 19 | 34.9412 | 85.1786 | 0
2020-10-23 |  | 19 | 33.7778 | 84.9253 | 0
2020-10-24 |  | 18 | 33.7778 | 85.405 | -1
2020-10-25 |  | 18 | 33.7778 | 85.4152 | 0
2020-10-26 |  | 18 | 36.0625 | 85.1763 | 0
2020-10-27 |  | 18 | 32.5789 | 84.4164 | 0
2020-10-28 |  | 16 | 31.3 | 84.3915 | -2
2020-10-29 |  | 16 | 32.2273 | 84.0883 | 0
2020-10-30 |  | 15 | 24.2857 | 83.9187 | -1
2020-10-31 |  | 14 | 25.2 | 84.1 | -1
2020-11-01 |  | 14 | 25.2 | 84.3477 | 0
2020-11-02 |  | 15 | 26 | 84.3477 | 1
2020-11-03 |  | 16 | 26 | 84.3477 | 1
2020-11-04 |  | 17 | 24.9 | 84.2914 | 1
2020-11-05 |  | 17 | 23.8095 | 84.2086 | 0
2020-11-06 |  | 17 | 23.9091 | 84.1254 | 0
2020-11-07 |  | 15 | 25.45 | 84.1254 | -2
2020-11-08 |  | 15 | 26.7222 | 84.1254 | 0
2020-11-09 |  | 16 | 20.875 | 84.1254 | 1
2020-11-10 |  | 17 | 20.875 | 84.1254 | 1
2020-11-11 |  | 17 | 20.875 | 84.1254 | 0
2020-11-12 |  | 18 | 20.5882 | 84.5884 | 1
2020-11-13 |  | 19 | 20.5882 | 84.5884 | 1
2020-11-14 |  | 19 | 20.5882 | 84.5884 | 0
2020-11-15 |  | 19 | 20.5882 | 85.281 | 0
2020-11-16 |  | 19 | 20.5882 | 85.281 | 0
2020-11-17 |  | 20 | 20.5882 | 85.281 | 1
2020-11-18 |  | 20 | 19.8333 | 84.9964 | 0
2020-11-19 |  | 19 | 20.4211 | 84.5415 | -1
2020-11-20 |  | 17 | 20.1 | 84.5884 | -2
2020-11-21 |  | 17 | 19.7778 | 84.5884 | 0
2020-11-22 |  | 17 | 19.7778 | 84.5884 | 0
2020-11-23 |  | 18 | 20.1176 | 84.5884 | 1
2020-11-24 |  | 19 | 20.1176 | 84.5884 | 1
2020-11-25 |  | 19 | 19.2105 | 84.0645 | 0
2020-11-26 |  | 17 | 19.2222 | 84.3201 | -2
2020-11-27 |  | 17 | 20.2667 | 84.3201 | 0
2020-11-28 |  | 17 | 21.2143 | 84.3201 | 0
2020-11-29 |  | 17 | 17.8333 | 84.3201 | 0
2020-11-30 |  | 17 | 18 | 84.3201 | 0
2020-12-01 |  | 17 | 21.8462 | 84.025 | 0
2020-12-02 |  | 16 | 20 | 83.4858 | -1
2020-12-03 |  | 14 | 20 | 83.4858 | -2
2020-12-04 |  | 15 | 20.5882 | 83.6454 | 1
2020-12-05 |  | 14 | 21.625 | 83.6454 | -1
2020-12-06 |  | 14 | 22.9333 | 83.6454 | 0
2020-12-07 |  | 15 | 26.2143 | 83.742 | 1
2020-12-08 |  | 15 | 26.2143 | 83.742 | 0
2020-12-09 |  | 15 | 26.2143 | 83.742 | 0
2020-12-10 |  | 16 | 24.8 | 83.4648 | 1
2020-12-11 |  | 16 | 24.8 | 83.4648 | 0
2020-12-12 |  | 16 | 24.8 | 83.4648 | 0
2020-12-13 |  | 16 | 27.7333 | 83.3803 | 0
2020-12-14 |  | 16 | 27.7333 | 83.3803 | 0
2020-12-15 |  | 17 | 28.25 | 83.214 | 1
2020-12-16 |  | 16 | 24.4737 | 82.6888 | -1
2020-12-17 |  | 13 | 23.7 | 82.4321 | -3
2020-12-18 |  | 17 | 23.7 | 82.4321 | 4
2020-12-19 |  | 17 | 24.5789 | 82.4321 | 0
2020-12-20 |  | 17 | 24.8824 | 82.2483 | 0
2020-12-21 |  | 19 | 28.5556 | 82.0417 | 2
2020-12-22 |  | 18 | 36 | 82.1862 | -1
2020-12-23 |  | 16 | 36 | 82.5347 | -2
2020-12-24 |  | 16 | 36 | 82.5347 | 0
2020-12-25 |  | 16 | 36 | 82.5347 | 0
2020-12-26 |  | 16 | 38.7222 | 82.5347 | 0
2020-12-27 |  | 16 | 38.7222 | 82.5436 | 0
2020-12-28 |  | 16 | 38.7222 | 82.5436 | 0
2020-12-29 |  | 16 | 38.7222 | 82.5436 | 0
2020-12-30 |  | 16 | 36.5 | 82.0865 | 0
2020-12-31 |  | 14 | 36.5 | 82.0865 | -2
2021-01-01 |  | 14 | 35.7778 | 82.0865 | 0
2021-01-02 |  | 14 | 39.25 | 82.0865 | 0
2021-01-03 |  | 14 | 39.25 | 82.0865 | 0
2021-01-04 |  | 15 | 41.2857 | 82.0865 | 1
2021-01-05 |  | 15 | 41.2857 | 82.0865 | 0
2021-01-06 |  | 17 | 41.2857 | 82.1424 | 2
2021-01-07 |  | 18 | 34.5 | 82.1514 | 1
2021-01-08 |  | 18 | 32.2667 | 82.0282 | 0
2021-01-09 |  | 17 | 32.2667 | 82.1673 | -1
2021-01-10 |  | 17 | 34.2143 | 82.3333 | 0
2021-01-11 |  | 17 | 32.1333 | 82.05 | 0
2021-01-12 |  | 16 | 32.3529 | 81.7092 | -1
2021-01-13 |  | 15 | 30.625 | 81.4731 | -1
2021-01-14 |  | 15 | 28.8571 | 80.6961 | 0
2021-01-15 |  | 10 | 26.4091 | 80.1519 | -5
2021-01-16 |  | 9 | 29.8947 | 80.1809 | -1
2021-01-17 |  | 9 | 31.0556 | 80.1809 | 0
2021-01-18 |  | 9 | 31.0556 | 80.1809 | 0
2021-01-19 |  | 10 | 31.0556 | 80.1809 | 1
2021-01-20 |  | 11 | 31.0556 | 80.55 | 1
2021-01-21 |  | 11 | 28.375 | 80.55 | 0
2021-01-22 |  | 11 | 23 | 80.1957 | 0
2021-01-23 |  | 8 | 23 | 80.4536 | -3
2021-01-24 |  | 8 | 23 | 80.4536 | 0
2021-01-25 |  | 8 | 23 | 80.4536 | 0
2021-01-26 |  | 9 | 23 | 80.4536 | 1
2021-01-27 |  | 9 | 23 | 80.4536 | 0
2021-01-28 |  | 10 | 23 | 80.7158 | 1
2021-01-29 |  | 12 | 23.8889 | 80.7338 | 2
2021-01-30 |  | 12 | 24.8125 | 80.7338 | 0
2021-01-31 |  | 12 | 24.8125 | 80.7338 | 0
2021-02-01 |  | 13 | 27.1111 | 80.4821 | 1
2021-02-02 |  | 12 | 27.1111 | 80.4821 | -1
2021-02-03 |  | 12 | 27.1111 | 80.7599 | 0
2021-02-04 |  | 12 | 27.1111 | 80.7599 | 0
2021-02-05 |  | 13 | 26.7368 | 80.4803 | 1
2021-02-06 |  | 13 | 26.7368 | 80.4803 | 0
2021-02-07 |  | 13 | 27.3333 | 80.4803 | 0
2021-02-08 |  | 13 | 28.8824 | 80.4803 | 0
2021-02-09 |  | 13 | 28.8824 | 80.4803 | 0
2021-02-10 |  | 14 | 27.8421 | 80.0681 | 1
2021-02-11 |  | 14 | 29.2222 | 80.0681 | 0
2021-02-12 |  | 15 | 28.625 | 79.8957 | 1
2021-02-13 |  | 15 | 28.625 | 79.8957 | 0
2021-02-14 |  | 15 | 31.0909 | 79.8957 | 0
2021-02-15 |  | 15 | 36.7778 | 79.8957 | 0
2021-02-16 |  | 17 | 36.7778 | 79.8957 | 2
2021-02-17 |  | 19 | 35.3 | 80.0686 | 2
2021-02-18 |  | 18 | 33.9091 | 79.7355 | -1
2021-02-19 |  | 17 | 33.9091 | 79.7564 | -1
2021-02-20 |  | 17 | 33.9091 | 79.8248 | 0
2021-02-21 |  | 17 | 33.9091 | 79.8645 | 0
2021-02-22 |  | 17 | 24 | 79.1196 | 0
2021-02-23 |  | 15 | 23 | 78.6763 | -2
2021-02-24 |  | 13 | 21.4286 | 78.6403 | -2
2021-02-25 |  | 12 | 22.3333 | 78.6547 | -1
2021-02-26 |  | 11 | 22.3333 | 78.6547 | -1
2021-02-27 |  | 12 | 22.3333 | 78.6547 | 1
2021-02-28 |  | 12 | 22.3333 | 78.9382 | 0
2021-03-01 |  | 12 | 21.1429 | 78.9382 | 0
2021-03-02 |  | 13 | 21.1429 | 78.9382 | 1
2021-03-03 |  | 13 | 21.1429 | 78.9382 | 0
2021-03-04 |  | 13 | 17.0833 | 78.9382 | 0
2021-03-05 |  | 13 | 17.0833 | 78.978 | 0
2021-03-06 |  | 14 | 17.0833 | 79.1287 | 1
2021-03-07 |  | 14 | 17.0833 | 79.1287 | 0
2021-03-08 |  | 16 | 16.8182 | 79.1287 | 2
2021-03-09 |  | 16 | 16.8182 | 79.1287 | 0
2021-03-10 |  | 18 | 16.8182 | 79.3284 | 2
2021-03-11 |  | 19 | 18.5833 | 79.1765 | 1
2021-03-12 |  | 18 | 16.375 | 78.1703 | -1
2021-03-13 |  | 14 | 16 | 78.5387 | -4
2021-03-14 |  | 14 | 16 | 78.5387 | 0
2021-03-15 |  | 15 | 16 | 78.5387 | 1
2021-03-16 |  | 16 | 16 | 78.5387 | 1
2021-03-17 |  | 17 | 21.3333 | 78.6029 | 1
2021-03-18 |  | 16 | 22.75 | 78.386 | -1
2021-03-19 |  | 17 | 22.75 | 78.386 | 1
2021-03-20 |  | 17 | 22.8 | 78.386 | 0
2021-03-21 |  | 17 | 23 | 78.386 | 0
2021-03-22 |  | 18 | 23 | 78.386 | 1
2021-03-23 |  | 23 | 23.9412 | 77.84 | 5
2021-03-24 |  | 22 | 23.9412 | 77.7191 | -1
2021-03-25 |  | 22 | 26.6429 | 77.7191 | 0
2021-03-26 |  | 23 | 25.4286 | 77.2082 | 1
2021-03-27 |  | 23 | 27.3077 | 78.0491 | 0
2021-03-28 |  | 23 | 26.6667 | 75.2162 | 0
2021-03-29 |  | 24 | 26.2308 | 75.0077 | 1
2021-03-30 |  | 25 | 26.2308 | 75.0077 | 1
2021-03-31 |  | 27 | 26.2308 | 75.0077 | 2
2021-04-01 |  | 28 | 26.2308 | 75.0077 | 1
2021-04-02 |  | 28 | 24.8571 | 74.95 | 0
2021-04-03 |  | 27 | 24.8571 | 75.124 | -1
2021-04-04 |  | 27 | 24.8571 | 74.9213 | 0
2021-04-05 |  | 27 | 24.9333 | 74.7294 | 0
2021-04-06 |  | 28 | 24.9333 | 74.7294 | 1
2021-04-07 |  | 28 | 24.9333 | 74.7294 | 0
2021-04-08 |  | 28 | 23.6471 | 74.2568 | 0
2021-04-09 |  | 32 | 23.6471 | 74.2885 | 4
2021-04-10 |  | 32 | 23.6471 | 74.0913 | 0
2021-04-11 |  | 32 | 22.75 | 74.0683 | 0
2021-04-12 |  | 33 | 24.6 | 73.3611 | 1
2021-04-13 |  | 30 | 24.875 | 73.1858 | -3
2021-04-14 |  | 29 | 24.875 | 73.3745 | -1
2021-04-15 |  | 29 | 24.875 | 73.648 | 0
2021-04-16 |  | 29 | 24.5294 | 73.5968 | 0
2021-04-17 |  | 29 | 20.0625 | 73.753 | 0
2021-04-18 |  | 29 | 18.4667 | 73.4204 | 0
2021-04-19 |  | 31 | 18.6111 | 72.7661 | 2
2021-04-20 |  | 28 | 18.7143 | 72.1275 | -3
2021-04-21 |  | 25 | 19.0645 | 70.1226 | -3
2021-04-22 |  | 16 | 19.0645 | 70.0731 | -9
2021-04-23 |  | 19 | 17.5517 | 69.8161 | 3
2021-04-24 |  | 18 | 17.5517 | 69.8161 | -1
2021-04-25 |  | 18 | 17.5517 | 69.6409 | 0
2021-04-26 |  | 18 | 18.8929 | 69.5192 | 0
2021-04-27 |  | 17 | 18.8929 | 69.5192 | -1
2021-04-28 |  | 17 | 18.7586 | 69.3103 | 0
2021-04-29 |  | 16 | 18.6786 | 69.2731 | -1
2021-04-30 |  | 16 | 18.2759 | 69.0345 | 0
2021-05-01 |  | 16 | 18.2759 | 69.0345 | 0
2021-05-02 |  | 16 | 18.2759 | 69.5251 | 0
2021-05-03 |  | 16 | 18.6786 | 69.5251 | 0
2021-05-04 |  | 19 | 18.6786 | 69.5251 | 3
2021-05-05 |  | 19 | 18.4483 | 69.7255 | 0
2021-05-06 |  | 20 | 18.1786 | 69.8425 | 1
2021-05-07 |  | 21 | 18.1786 | 69.9565 | 1
2021-05-08 |  | 21 | 18.1786 | 69.9565 | 0
2021-05-09 |  | 21 | 18.5 | 69.592 | 0
2021-05-10 |  | 21 | 18.5 | 69.592 | 0
2021-05-11 |  | 21 | 18.5 | 69.592 | 0
2021-05-12 |  | 21 | 18 | 69.4739 | 0
2021-05-13 |  | 21 | 28.8 | 70.308 | 0
2021-05-14 |  | 20 | 28.7917 | 70.308 | -1
2021-05-15 |  | 20 | 28.7917 | 70.4217 | 0
2021-05-16 |  | 20 | 28.7917 | 69.8508 | 0
2021-05-17 |  | 20 | 27.6154 | 69.1992 | 0
2021-05-18 |  | 17 | 27.6154 | 69.1992 | -3
2021-05-19 |  | 18 | 27.6154 | 69.1992 | 1
2021-05-20 |  | 19 | 28.6957 | 69.1992 | 1
2021-05-21 |  | 21 | 29.619 | 69.1912 | 2
2021-05-22 |  | 20 | 38.5455 | 69.448 | -1
2021-05-23 |  | 20 | 38.5455 | 69.448 | 0
2021-05-24 |  | 21 | 42.1 | 69.448 | 1
2021-05-25 |  | 23 | 42.1 | 69.448 | 2
2021-05-26 |  | 24 | 42.1 | 69.448 | 1
2021-05-27 |  | 25 | 42.5556 | 69.448 | 1
2021-05-28 |  | 26 | 42.5556 | 69.498 | 1
2021-05-29 |  | 26 | 46 | 69.498 | 0
2021-05-30 |  | 26 | 46 | 68.5714 | 0
2021-05-31 |  | 26 | 51.5714 | 68.5714 | 0
2021-06-01 |  | 26 | 46.125 | 68.3252 | 0
2021-06-02 |  | 25 | 46.125 | 68.5802 | -1
2021-06-03 |  | 29 | 46.125 | 68.5802 | 4
2021-06-04 |  | 29 | 45.6667 | 68.7243 | 0
2021-06-05 |  | 28 | 49.875 | 68.8926 | -1
2021-06-06 |  | 28 | 49.875 | 68.9585 | 0
2021-06-07 |  | 28 | 49 | 68.7654 | 0
2021-06-08 |  | 27 | 49 | 68.7654 | -1
2021-06-09 |  | 28 | 96.0909 | 71.0905 | 1
2021-06-10 |  | 28 | 88.25 | 70.8074 | 0
2021-06-11 |  | 27 | 88.25 | 70.9218 | -1
2021-06-12 |  | 27 | 95.8182 | 70.9218 | 0
2021-06-13 |  | 27 | 77.6 | 70.9218 | 0
2021-06-14 |  | 27 | 75.8182 | 70.8689 | 0
2021-06-15 |  | 27 | 72.8333 | 70.7429 | 0
2021-06-16 |  | 26 | 72.8333 | 70.7429 | -1
2021-06-17 |  | 26 | 81.6364 | 70.4534 | 0
2021-06-18 |  | 24 | 78.7692 | 70.4818 | -2
2021-06-19 |  | 22 | 78.7692 | 70.4818 | -2
2021-06-20 |  | 22 | 78.7692 | 70.7061 | 0
2021-06-21 |  | 22 | 86.7692 | 70.9228 | 0
2021-06-22 |  | 22 | 72 | 70.172 | 0
2021-06-23 |  | 18 | 69.8889 | 70.3494 | -4
2021-06-24 |  | 17 | 70.6316 | 70.6573 | -1
2021-06-25 |  | 16 | 70.6316 | 70.9431 | -1
2021-06-26 |  | 16 | 65.0952 | 70.4718 | 0
2021-06-27 |  | 14 | 65.0952 | 70.5041 | -2
2021-06-28 |  | 15 | 65.0952 | 70.5041 | 1
2021-06-29 |  | 15 | 65.0952 | 70.5041 | 0
2021-06-30 |  | 15 | 64.6818 | 70.5976 | 0
2021-07-01 | 0 | 14 | 64.6818 | 70.5976 | -1
2021-07-02 | 0 | 14 | 67.381 | 71.0328 | 0
2021-07-03 | 0 | 14 | 67.381 | 71.0328 | 0
2021-07-04 | 0 | 14 | 67.381 | 71.0328 | 0
2021-07-05 | 0 | 14 | 68.65 | 71.0328 | 0
2021-07-06 | 0 | 14 | 68.65 | 71.0328 | 0
2021-07-07 | 0 | 16 | 68.65 | 71.0905 | 2
2021-07-08 | 47 | 17 | 69.9474 | 70.9918 | 1
2021-07-09 | 100.5 | 16 | 74.15 | 71.3306 | -1
2021-07-10 | 100.5 | 16 | 48.2105 | 71.2787 | 0
2021-07-11 | 100.5 | 16 | 50.7778 | 71.2787 | 0
2021-07-12 | 81.6667 | 16 | 50.4211 | 71.1673 | 0
2021-07-13 | 81.6667 | 17 | 50.4211 | 71.1673 | 1
2021-07-14 | 79.8 | 17 | 52.9524 | 70.5592 | 0
2021-07-15 | 70.3333 | 15 | 51.2857 | 70.3659 | -2
2021-07-16 | 70.3333 | 14 | 51.85 | 69.4774 | -1
2021-07-17 | 70.3333 | 15 | 51.85 | 69.7386 | 1
2021-07-18 | 70.3333 | 15 | 53.7222 | 70.1681 | 0
2021-07-19 | 70.3333 | 17 | 52.5625 | 70.1681 | 2
2021-07-20 | 70.3333 | 17 | 52.5625 | 70.1681 | 0
2021-07-21 | 61.8571 | 17 | 50.1176 | 69.9205 | 0
2021-07-22 | 61.8571 | 16 | 45.5 | 69.9205 | -1
2021-07-23 | 61.8571 | 17 | 52.6667 | 70.0798 | 1
2021-07-24 | 61.8571 | 17 | 54.3636 | 70.4703 | 0
2021-07-25 | 61.8571 | 17 | 51.4 | 70.7447 | 0
2021-07-26 | 54.875 | 18 | 47.2727 | 70.4703 | 1
2021-07-27 | 54.875 | 17 | 55 | 70.4703 | -1
2021-07-28 | 54.875 | 18 | 55 | 69.9915 | 1
2021-07-29 | 61.1111 | 18 | 60.6 | 70.3761 | 0
2021-07-30 | 61.1111 | 18 | 60.6 | 70.6609 | 0
2021-07-31 | 61.1111 | 18 | 61.1111 | 71.0952 | 0
2021-08-01 | 61.1111 | 18 | 61.1111 | 71.0952 | 0
2021-08-02 | 61.1111 | 19 | 61.1111 | 71.0952 | 1
2021-08-03 | 54.7273 | 20 | 54.7273 | 70.7082 | 1
2021-08-04 | 52.3077 | 18 | 52.3077 | 70.2532 | -2
2021-08-05 | 51.2143 | 17 | 51.2143 | 70.6623 | -1
2021-08-06 | 45.8235 | 17 | 45.8235 | 70.2414 | 0
2021-08-07 | 45.8235 | 14 | 45.8235 | 70.2414 | -3
2021-08-08 | 45.8235 | 14 | 45.75 | 70.3636 | 0
2021-08-09 | 45.8235 | 14 | 38.5333 | 70.3636 | 0
2021-08-10 | 45.8235 | 15 | 38.5333 | 70.3636 | 1
2021-08-11 | 46.7222 | 16 | 40 | 70.3276 | 1
2021-08-12 | 46.7222 | 15 | 39.7333 | 70.3276 | -1
2021-08-13 | 46.7222 | 16 | 39.7333 | 70.3276 | 1
2021-08-14 | 46.7222 | 16 | 34 | 70.8696 | 0
2021-08-15 | 46.7222 | 16 | 34.9167 | 70.9648 | 0
2021-08-16 | 44.7895 | 17 | 33 | 70.6974 | 1
2021-08-17 | 42.1429 | 19 | 30.8667 | 70.2304 | 2
2021-08-18 | 45.875 | 18 | 37.7222 | 35.8687 | -1
2021-08-19 | 45.875 | 16 | 37.7222 | 35.5482 | -2
2021-08-20 | 44.56 | 19 | 36.4211 | 35.6735 | 3
2021-08-21 | 44.56 | 19 | 37.8333 | 35.5206 | 0
2021-08-22 | 44.56 | 19 | 37.8333 | 35.5206 | 0
2021-08-23 | 44.56 | 19 | 37.8333 | 35.5206 | 0
2021-08-24 | 43.8077 | 19 | 37.1579 | 35.4667 | 0
2021-08-25 | 45.5556 | 19 | 39.85 | 35.75 | 0
2021-08-26 | 44.1429 | 20 | 39.85 | 35.599 | 1
2021-08-27 | 43.2069 | 19 | 38.7619 | 34.5816 | -1
2021-08-28 | 43.2069 | 18 | 38.7619 | 34.5816 | -1
2021-08-29 | 43.2069 | 18 | 35.15 | 34.7179 | 0
2021-08-30 | 43.6774 | 21 | 36.5455 | 34.8782 | 3
2021-08-31 | 42.303 | 19 | 35.25 | 34.7387 | -2
2021-09-01 | 42.303 | 17 | 35.25 | 34.7387 | -2
2021-09-02 | 42.303 | 17 | 35.25 | 34.7387 | 0
2021-09-03 | 42.303 | 18 | 36.0909 | 34.7387 | 1
2021-09-04 | 42.303 | 18 | 35.8 | 34.7387 | 0
2021-09-05 | 42.303 | 18 | 35.7368 | 34.8788 | 0
2021-09-06 | 42.303 | 18 | 38.5625 | 34.8788 | 0
2021-09-07 | 46.1714 | 19 | 46.5 | 35.63 | 1
2021-09-08 | 45.3889 | 17 | 45 | 35.5423 | -2
2021-09-09 | 45.3889 | 17 | 45 | 35.5423 | 0
2021-09-10 | 45.3889 | 18 | 45 | 35.601 | 1
2021-09-11 | 45.3889 | 18 | 44.0556 | 35.601 | 0
2021-09-12 | 45.3889 | 18 | 44.0556 | 35.601 | 0
2021-09-13 | 45.3889 | 18 | 44.0556 | 35.601 | 0
2021-09-14 | 44.8108 | 18 | 43 | 35.5427 | 0
2021-09-15 | 44.8108 | 17 | 43 | 35.5427 | -1
2021-09-16 | 44.8108 | 19 | 44.8333 | 35.5427 | 2
2021-09-17 | 44.8108 | 19 | 48.3125 | 34.2234 | 0
2021-09-18 | 44.8108 | 19 | 42.8462 | 34.2602 | 0
2021-09-19 | 44.8108 | 19 | 42.8462 | 34.2602 | 0
2021-09-20 | 44.1282 | 19 | 43.3571 | 34.2323 | 0
2021-09-21 | 43.575 | 18 | 41.9333 | 34.1709 | -1
2021-09-22 | 41.6364 | 19 | 37.7895 | 34.0743 | 1
2021-09-23 | 41.6364 | 15 | 37.7895 | 34.0743 | -4

## precision
precision
0 | 0.079585
1 | 0.080321
2 | 0.092683
3 | 0.097938
4 | 0.098958
5 | 0.10053
6 | 0.099379
7 | 0.1
8 | 0.10063
9 | 0.10191
10 | 0.1039
11 | 0.10274
12 | 0.10345
13 | 0.10448
14 | 0.10606
15 | 0.10769
16 | 0.11382
17 | 0.11864
18 | 0.21622
19 | 0.21875
20 | 0.24138
21 | 0.26923
22 | 0.30435
23 | 0.33333
24 | 0.31579
25 | 0.75
26 | 0.83333
27 | 1

## thresholds
thresholds
0 | 0
1 | 0.033764
2 | 0.051887
3 | 0.065915
4 | 0.11111
5 | 0.14729
6 | 0.15385
7 | 0.17949
8 | 0.20833
9 | 0.25
10 | 0.26
11 | 0.26667
12 | 0.27477
13 | 0.27778
14 | 0.30134
15 | 0.32273
16 | 0.32353
17 | 0.32529
18 | 0.34783
19 | 0.35312
20 | 0.35879
21 | 0.40375
22 | 0.41176
23 | 0.41216
24 | 0.42424
25 | 0.42857
26 | 0.46317
27 | 0.62222

## recall
recall
0 | 0.86957
1 | 0.82609
2 | 0.82609
3 | 0.82609
4 | 0.82609
5 | 0.69565
6 | 0.69565
7 | 0.69565
8 | 0.69565
9 | 0.69565
10 | 0.65217
11 | 0.65217
12 | 0.6087
13 | 0.6087
14 | 0.6087
15 | 0.6087
16 | 0.6087
17 | 0.34783
18 | 0.30435
19 | 0.30435
20 | 0.30435
21 | 0.30435
22 | 0.30435
23 | 0.26087
24 | 0.26087
25 | 0.21739
26 | 0.17391
27 | 0